## Microbalance
Microbalance AXR 2MA
Next-generation weighing technology with a solid, durable construction for long-term reliability. High-precision performance with 2.1 g capacity and 1 µg readability for accurate micro-weighing. Intuitive color TFT touchscreen display with internal calibration and easy leveling alerts. Supports 21 CFR Part 11 compliant user management for secure and regulated laboratory operations. Multiple connectivity options including USB, RS232, and Ethernet for seamless data integration. Compact design with overload protection and power-saving mode, ideal for modern laboratories.
| Weighing Capacity | 2.1 g
| Readability | 1 µg
| Minimum Weight (USP, 0.1%, typical) | 1 mg
Advanced EMFC Technology for Trusted Results | Excellent Weighing Performance | Exceptional Accuracy for small samples | New Touchscreen as Smartphone | Touch Free Draftshield Operational | 21 CFR Part 11 Compliance | Minimum Sample Weight | Diversity | Digital Level | Audit Safe Result | 10 Easy-to-use Built-in-Applications | Data Free Transfer
Advanced EMFC sensor technology delivers international-standard performance. Engineered in India, AXR Series balances provide long-term accuracy, fast stabilization, excellent repeatability, overload protection, self-test reliability, and consistent results. | The Microbalance, with its two-terminal concept, offers excellent weighing performance by separating the electronics from the precise mechanical measuring cell. This design helps to eliminate the effects of temperature and its influence on the balance's performance. | With one microgram readability, low minimum weight, and an improved repeatability by 20 percent, you can accurately weigh small samples with the utmost accuracy. This is one of the biggest merits of this machine. | It’s 5.7" large colour TFT display and the capacitive touchscreen allows you to effortless swipe through operations, configurations, Tasks and Users. It has an Intuitive interface which is operable through cotton, silicon or rubber gloves. | The draft shield has wide opening, clear visibility and Infrared sensors to open and close the draft-shield doors with the wave of a hand. It simplifies dosing, minimizes spills, and reduces cross-contamination. | Our machine fully meets the requirements of 21 CFR Part 11/ EU GMP Annex 11. It also provides features like Full data security, Electronic signatures, Validations, Electronic records and Audit trails. | The balance displays whether the value of the weight is below/above the minimum weight threshold. 3 MSW method can be configure e.g. USP, GMP and User defined with and Without Tare Weight. | The tasks can be custom-made for each small unique configuration of parameter and make it quick, errorless operation. This saves you time and provides higher productivity. It also has customised print format. | Instead of manually keep track of your level your scale. It automatically alerts before performance can be compromised. Digital level saves your time and gives you maximum precision in your measurements. | Innovative audit safe with alibi and audit trial. As you carry out your operation, your results and changed parameters are automatically recorded on Alibi and Audi trial makes it operation complete audit safe. | The ten easy to use build in applications include density determination, check weighing, totaling, and formulation, simplify your routine weighing tasks. These applications make your analytical operations much easier and accurate. | Simply transfer task parameters and results to a PC via many interface ports like RS232, USB A, USB B, Ethernet and wifi. Transcription errors are completely eliminated and reliability of your data is preserved.
| Specifications - Ultra Microbalance AXR 2UA | Specifications - Microbalance AXR 2MA | Specifications - Microbalance AXR 6MA | Specifications - Microbalance AXR 11MA
Maximum Capacity | 2.1 g | 2.1 g | 6.1 g | 11.1 g
Minimum load (OIML) | 0.01 mg | 0.1 mg | 0.1 mg | 0.1 mg
Readability | 0.1 µg | 1 µg | 1 µg | 1 µg
Verification unit [e] | 1 mg | 1 mg | 1 mg | 1 mg
Repeatability | 0.1 µg | 1 µg | 1 µg | 1.2 µg
Linearity | 1.5 µg | 3 µg | 6 µg | 6 µg
Tare Range (Subtractive) | -2.1 g | -2.1 g | -6.1 g | -11.1 g
Standard repeatability [5% Max] | 0.25 µg | 0.5 µg | 1.2 µg | 1.2 µg
Minimum Weight USP | 0.5 mg | 1 mg | 2.4 mg | 2.4 mg
Minimum weight ((U=1%, k=2) | 0.05 mg | 0.1 mg | 0.24 mg | 0.24 mg
Weighing Pan Diameter | Ø 16 mm | Ø 26 mm | Ø 26 mm | Ø 26 mm
Setting Time | 10-20 Sec | 8-10 Sec | 8-10 Sec | 8-10 Sec
Display | 5.7'' large colour TFT display | 5.7'' large colour TFT display | 5.7'' large colour TFT display | 5.7'' large colour TFT display
Calibration | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC)
User Management | Log History [21 CFR Part 11 Compliant ] 7 User management Password protection AuditTrai Minimum Sample Weight (MSW) | Log History [21 CFR Part 11 Compliant ] 7 User management Password protection AuditTrai Minimum Sample Weight (MSW) | Log History [21 CFR Part 11 Compliant ] 7 User management Password protection AuditTrai Minimum Sample Weight (MSW) | Log History [21 CFR Part 11 Compliant ] 7 User management Password protection AuditTrai Minimum Sample Weight (MSW)
Features | Touch-free operation Infrared Proximity Sensor [Automatic Door] 4 built-in languages English, German, French, Spanish Levelling guidance | Touch-free operation Infrared Proximity Sensor [Automatic Door] 4 built-in languages English, German, French, Spanish Levelling guidance | Touch-free operation Infrared Proximity Sensor [Automatic Door] 4 built-in languages English, German, French, Spanish Levelling guidance | Touch-free operation Infrared Proximity Sensor [Automatic Door] 4 built-in languages English, German, French, Spanish Levelling guidance
Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
Operating temperature | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC
Relative humidity | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80%
Sensitivity Drift | ± 1ppm (1 x 10-6 / ºC) | ± 1ppm (1 x 10-6 / ºC) | ± 1ppm (1 x 10-6 / ºC) | ± 1ppm (1 x 10-6 / ºC)
Protection class | IP 43 | IP 43 | IP 43 | IP 43
Leveling system | Manual | Manual | Manual | Manual
Interfaces | RS232/USB | RS232/USB | RS232/USB | RS232/USB
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold
Power Supply | DC Adapter, input 100 ~ 240V 0.5A output 13V / 0.9A | DC Adapter, input 100 ~ 240V 0.5A output 13V / 0.9A | DC Adapter, input 100 ~ 240V 0.5A output 13V / 0.9A | DC Adapter, input 100 ~ 240V 0.5A output 13V / 0.9A
Compliance | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP
Dimensions for Electronic Box (W x D x H)
Dimensions for Mechanical Box (W x D x H)
Packaging dimensions (W x D X H) | 520 x 780 x 680 mm | 520 x 780 x 680 mm | 520 x 780 x 680 mm | 520 x 780 x 680 mm
Net weight
Gross weight | 18 Kg | 18 Kg | 18 Kg | 18 Kg
Filter Microbalance AXR 2MF
Advanced weighing technology with a solid, robust construction designed for filter microbalance applications. High-accuracy performance with 2.1 g capacity and 1 µg readability for ultra-precise measurements. Ideal for filter weighing in environmental monitoring, air quality analysis, and particulate matter (PM) testing. User-friendly color TFT touchscreen with internal calibration and easy leveling alerts for consistent results. Secure 21 CFR Part 11 compliant user management with USB, RS232, and Ethernet connectivity. Compact footprint with overload protection and power-saving mode, optimized for daily laboratory workflows.
| Weighing Capacity | 2.1 g
| Readability | 1 µg
| Minimum Weight (USP, 0.1%, typical) | 1 mg
Advanced EMFC Technology for Trusted Results | Excellent Weighing Performance | Exceptional Accuracy for small samples | New Touchscreen as Smartphone | Touch Free Draftshield Operational | 21 CFR Part 11 Compliance | Minimum Sample Weight | Diversity | Digital Level | Audit Safe Result | 10 Easy-to-use Built-in-Applications | Data Free Transfer
Advanced EMFC sensor technology delivers international-standard performance. Engineered in India, AXR Series balances provide long-term accuracy, fast stabilization, excellent repeatability, overload protection, self-test reliability, and consistent results. | The Microbalance, with its two-terminal concept, offers excellent weighing performance by separating the electronics from the precise mechanical measuring cell. This design helps to eliminate the effects of temperature and its influence on the balance's performance. | With one microgram readability, low minimum weight, and an improved repeatability by 20 percent, you can accurately weigh small samples with the utmost accuracy. This is one of the biggest merits of this machine. | It’s 5.7" large colour TFT display and the capacitive touchscreen allows you to effortless swipe through operations, configurations, Tasks and Users. It has an Intuitive interface which is operable through cotton, silicon or rubber gloves. | The draft shield has wide opening, clear visibility and Infrared sensors to open and close the draft-shield doors with the wave of a hand. It simplifies dosing, minimizes spills, and reduces cross-contamination. | Our machine fully meets the requirements of 21 CFR Part 11/ EU GMP Annex 11. It also provides features like Full data security, Electronic signatures, Validations, Electronic records and Audit trails. | The balance displays whether the value of the weight is below/above the minimum weight threshold. 3 MSW method can be configure e.g. USP, GMP and User defined with and Without Tare Weight. | The tasks can be custom-made for each small unique configuration of parameter and make it quick, errorless operation. This saves you time and provides higher productivity. It also has customised print format. | Instead of manually keep track of your level your scale. It automatically alerts before performance can be compromised. Digital level saves your time and gives you maximum precision in your measurements. | Innovative audit safe with alibi and audit trial. As you carry out your operation, your results and changed parameters are automatically recorded on Alibi and Audi trial makes it operation complete audit safe. | The ten easy to use build in applications include density determination, check weighing, totaling, and formulation, simplify your routine weighing tasks. These applications make your analytical operations much easier and accurate. | Simply transfer task parameters and results to a PC via many interface ports like RS232, USB A, USB B, Ethernet and wifi. Transcription errors are completely eliminated and reliability of your data is preserved.
| Specifications - Microbalance Filter AXR 2MF | Specifications - Microbalance Filter AXR 5MF | Specifications - Microbalance Filter AXR 6MF
Maximum Capacity | 2.1 g | 5.1 g | 6.1 g
Minimum load (OIML) | 0.1 mg | 0.1 mg | 0.1 mg
Readability | 1 µg | 1 µg | 1 µg
Verification unit [e] | 1 mg | 1 mg | 1 mg
Repeatability | 1 µg | 1 µg | 1 µg
Linearity | 3 µg | 6 µg | 6 µg
Tare Range (Subtractive) | -2.1 g | -6.1 g | -6.1 g
Standard repeatability [5% Max] | 0.5 µg | 1.2 µg | 1.2 µg
Minimum Weight USP | 1 mg | 2.4 mg | 2.4 mg
Minimum weight ((U=1%, k=2) | 0.1 mg | 0.24 mg | 0.24 mg
Weighing Pan Diameter | Ø 50 mm & Optional Ø 70 mm | Ø 50 mm & Optional Ø 70 mm | Ø 50 mm & Optional Ø 70 mm
Setting Time | 8-10 Sec | 8-10 Sec | 8-10 Sec
Display | 5.7'' large colour TFT display | 5.7'' large colour TFT display | 5.7'' large colour TFT display
Calibration | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC)
User Management | Log History [21 CFR Part 11 Compliant ] 7 User management Password protection AuditTrai Minimum Sample Weight (MSW) | Log History [21 CFR Part 11 Compliant ] 7 User management Password protection AuditTrai Minimum Sample Weight (MSW) | Log History [21 CFR Part 11 Compliant ] 7 User management Password protection AuditTrai Minimum Sample Weight (MSW)
Features | Manual Door 4 built-in languages English, German, French, Spanish Levelling guidance | Manual Door 4 built-in languages English, German, French, Spanish Levelling guidance | Manual Door 4 built-in languages English, German, French, Spanish Levelling guidance
Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
Operating temperature | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC
Relative humidity | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80%
Sensitivity Drift | ± 1ppm (1 x 10-6 / ºC) | ± 1ppm (1 x 10-6 / ºC) | ± 1ppm (1 x 10-6 / ºC)
Protection class | IP 43 | IP 43 | IP 43
Leveling system | Manual | Manual | Manual
Interfaces | RS232/USB | RS232/USB | RS232/USB
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold
Power Supply | DC Adapter, input 100 ~ 240V 0.5A output 13V / 0.9A | DC Adapter, input 100 ~ 240V 0.5A output 13V / 0.9A | DC Adapter, input 100 ~ 240V 0.5A output 13V / 0.9A
Compliance | ISO, GLP / GMP & USP | ISO, GLP / GMP & USP | ISO, GLP / GMP & USP
Dimensions for Electronic Box (W x D x H)
Dimensions for Mechanical Box (W x D x H)
Packaging dimensions (W x D X H) | 520 x 780 x 680 mm | 520 x 780 x 680 mm | 520 x 780 x 680 mm
Net weight
Gross weight | 18 Kg | 18 Kg | 18 Kg
Pipette Microbalance AXR 11MP
Advanced weighing technology with a solid, robust construction designed for precision microbalance applications. High-accuracy performance with 11.1 g capacity and 1 µg readability for ultra-precise gravimetric measurements. Ideal for pipette calibration and gravimetric testing in accordance with ISO 8655 standards. User-friendly color TFT touchscreen display with internal calibration and easy leveling alerts for reliable results. Supports 21 CFR Part 11 compliant user management with USB, RS232, and Ethernet connectivity. Compact footprint with overload protection and power-saving mode, optimized for regulated laboratory environments.
| Weighing Capacity | 11.1 g
| Readability | 1 µg
| Minimum Weight (USP, 0.1%, typical) | 2.4 mg
Advanced EMFC Technology for Trusted Results | Excellent Weighing Performance | Exceptional Accuracy for small samples | New Touchscreen as Smartphone | Touch Free Draftshield Operational | 21 CFR Part 11 Compliance | Minimum Sample Weight | Diversity | Digital Level | Audit Safe Result | 10 Easy-to-use Built-in-Applications | Data Free Transfer
Advanced EMFC sensor technology delivers international-standard performance. Engineered in India, AXR Series balances provide long-term accuracy, fast stabilization, excellent repeatability, overload protection, self-test reliability, and consistent results. | The Microbalance, with its two-terminal concept, offers excellent weighing performance by separating the electronics from the precise mechanical measuring cell. This design helps to eliminate the effects of temperature and its influence on the balance's performance. | With one microgram readability, low minimum weight, and an improved repeatability by 20 percent, you can accurately weigh small samples with the utmost accuracy. This is one of the biggest merits of this machine. | It’s 5.7" large colour TFT display and the capacitive touchscreen allows you to effortless swipe through operations, configurations, Tasks and Users. It has an Intuitive interface which is operable through cotton, silicon or rubber gloves. | The draft shield has wide opening, clear visibility and Infrared sensors to open and close the draft-shield doors with the wave of a hand. It simplifies dosing, minimizes spills, and reduces cross-contamination. | Our machine fully meets the requirements of 21 CFR Part 11/ EU GMP Annex 11. It also provides features like Full data security, Electronic signatures, Validations, Electronic records and Audit trails. | The balance displays whether the value of the weight is below/above the minimum weight threshold. 3 MSW method can be configure e.g. USP, GMP and User defined with and Without Tare Weight. | The tasks can be custom-made for each small unique configuration of parameter and make it quick, errorless operation. This saves you time and provides higher productivity. It also has customised print format. | Instead of manually keep track of your level your scale. It automatically alerts before performance can be compromised. Digital level saves your time and gives you maximum precision in your measurements. | Innovative audit safe with alibi and audit trial. As you carry out your operation, your results and changed parameters are automatically recorded on Alibi and Audi trial makes it operation complete audit safe. | The ten easy to use build in applications include density determination, check weighing, totaling, and formulation, simplify your routine weighing tasks. These applications make your analytical operations much easier and accurate. | Simply transfer task parameters and results to a PC via many interface ports like RS232, USB A, USB B, Ethernet and wifi. Transcription errors are completely eliminated and reliability of your data is preserved.
Specifications - Microbalance Filter AXR 11MP
Maximum Capacity | 11.1 g
Minimum load (OIML) | 0.1 mg
Readability | 1 µg
Verification unit [e] | 1 mg
Repeatability | 1.2 µg
Linearity | 6 µg
Tare Range (Subtractive) | -11.1 g
Standard repeatability [5% Max] | 1.2 µg
Minimum Weight USP | 2.4 mg
Minimum weight ((U=1%, k=2) | 0.24 mg
Weighing Pan Diameter | Ø 20 mm
Setting Time | 8-10 Sec
Display | 5.7'' large colour TFT display
Calibration | Motorised Internal Calibration (automatic/PSAC)
User Management | Log History [21 CFR Part 11 Compliant ] 7 User management Password protection AuditTrai Minimum Sample Weight (MSW)
Features | Touch-free operation Infrared Proximity Sensor [Automatic Door] 4 built-in languages English, German, French, Spanish Levelling guidance
Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
Operating temperature | 15º to 35ºC
Relative humidity | 40% ÷ 80%
Sensitivity Drift | ± 1ppm (1 x 10-6 / ºC)
Protection class | IP 43
Leveling system | Manual
Interfaces | RS232/USB
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold
Power Supply | DC Adapter, input 100 ~ 240V 0.5A output 13V / 0.9A
Compliance | ISO, GLP, GMP, USP
Dimensions for Electronic Box (W x D x H)
Dimensions for Mechanical Box (W x D x H)
Packaging dimensions (W x D X H) | 520 x 780 x 680 mm
Net weight
Gross weight | 18 Kg
STENT Microbalance AXR 6MS
Advanced weighing technology with a solid, robust construction designed for precision microbalance applications. High-accuracy performance with 6.1 g capacity and 1 µg readability, ideal for stent weighing in medical device manufacturing requiring ultra-precise mass determination. User-friendly color TFT touchscreen display with internal calibration and easy leveling alerts for reliable results. Supports 21 CFR Part 11–compliant user management with USB, RS232, and Ethernet connectivity. Compact footprint with overload protection and power-saving mode, optimized for regulated laboratory environments.
| Weighing Capacity | 6.1 g
| Readability | 1 µg
| Minimum Weight (USP, 0.1%, typical) | 2 mg
Advanced EMFC Technology for Trusted Results | Excellent Weighing Performance | Environment and Internal Calibration | New Touchscreen as Smartphone | Touch Free Draftshield Operational | 21 CFR Part 11 Compliance | Minimum Sample Weight | Diversity | Digital Level | Audit Safe Result | 10 Easy-to-use Built-in-Applications | Data Free Transfer
Advanced EMFC sensor technology delivers international-standard performance. Engineered in India, AXR Series balances provide long-term accuracy, fast stabilization, excellent repeatability, overload protection, self-test reliability, and consistent results. | The Mono-Tech technology sensor truly stands out with its unparalleled design based on electromagnetic force compensation. Its exceptional performance meets the highest international standards. | Motorized built-in weight internal calibration ensures precise weighing performance in any condition. The PSAC Cal. function ensures accurate measurements in any condition. It performs automatic internal calibration based on time or temperature changes to maintain precision. | It’s 5.7" large colour TFT display and the capacitive touchscreen allows you to effortless swipe through operations, configurations, Tasks and Users. It has an Intuitive interface which is operable through cotton, silicon or rubber gloves. | The draft shield has wide opening, clear visibility and Infrared sensors to open and close the draft-shield doors with the wave of a hand. It simplifies dosing, minimizes spills, and reduces cross-contamination. | Our machine fully meets the requirements of 21 CFR Part 11/ EU GMP Annex 11. It also provides features like Full data security, Electronic signatures, Validations, Electronic records and Audit trails. | The balance displays whether the value of the weight is below/above the minimum weight threshold. 3 MSW method can be configure e.g. USP, GMP and User defined with and Without Tare Weight. | The tasks can be custom-made for each small unique configuration of parameter and make it quick, errorless operation. This saves you time and provides higher productivity. It also has customised print format. | Instead of manually keep track of your level your scale. It automatically alerts before performance can be compromised. Digital level saves your time and gives you maximum precision in your measurements. | Innovative audit safe with alibi and audit trial. As you carry out your operation, your results and changed parameters are automatically recorded on Alibi and Audi trial makes it operation complete audit safe. | The ten easy to use build in applications include density determination, check weighing, totaling, and formulation, simplify your routine weighing tasks. These applications make your analytical operations much easier and accurate. | Simply transfer task parameters and results to a PC via many interface ports like RS232, USB A, USB B, Ethernet and wifi. Transcription errors are completely eliminated and reliability of your data is preserved.
Specifications - Microbalance Filter AXR 6MS
Maximum Capacity | 6.1 g
Minimum load (OIML) | 0.1 mg
Readability | 1 µg
Verification unit [e] | 1 mg
Repeatability | 2 µg
Linearity | 6 µg
Tare Range (Subtractive) | -6.1 g
Standard repeatability [5% Max] | 1.2 µg
Minimum Weight USP | 2.4 mg
Minimum weight ((U=1%, k=2) | 0.24 mg
Weighing Pan Diameter | Ø 50 mm & Optional Ø 70 mm
Setting Time | 8-10 Sec
Display | 5.7'' large colour TFT display
Calibration | Motorised Internal Calibration (automatic/PSAC)
User Management | Log History [21 CFR Part 11 Compliant ] 7 User management Password protection AuditTrai Minimum Sample Weight (MSW)
Features | Manual Door 4 built-in languages English, German, French, Spanish Levelling guidance
Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
Operating temperature | 15º to 35ºC
Relative humidity | 40% ÷ 80%
Sensitivity Drift | ± 1ppm (1 x 10-6 / ºC)
Protection class | IP 43
Leveling system | Manual
Interfaces | RS232/USB
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold
Power Supply | DC Adapter, input 100 ~ 240V 0.5A output 13V / 0.9A
Compliance | ISO, GLP / GMP & USP
Dimensions for Electronic Box (W x D x H)
Dimensions for Mechanical Box (W x D x H)
Packaging dimensions (W x D X H) | 520 x 780 x 680 mm
Net weight
Gross weight | 18 Kg
Microbalance CM 2
Intuitive to operate with solid built construction; 2.1 g capacity; 1 µg readability; large, bright LCD backlit display; internal calibration; easy leveling; USB and RS232 connection; overload protection; application: accurate weighing of microgram samples for pharmaceutical, analytical, and research laboratory use.
| Weighing Capacity | 2.1 g
| Readability | 1 µg
| Minimum Weight (USP, 0.1%, typical) | 1 mg
Technology with Precise Weighing | Excellent Weighing Performance | Brilliant Proven Technology | Automatic Internal Calibration (PSAC) | Revolution in Mass Measurement Quantity | Touch Free Draftshield Operational | Bright, High-Contrast Display | 10 Easy-to-Use Built-In Applications | ISO/GLP Documentation | Seamless Connectivity | Quick & Simple Leveling | Easy Maintenance & Cleaning
The technology has unparalleled design based on electromagnetic force compensation (EMFC). Its exceptional performance meets the highest international standards. The robust stainless-steel housing guarantees short stabilization time and eliminates heat transfer. | The Microbalance, with its two-terminal concept, offers excellent weighing performance by separating the electronics from the precise mechanical measuring cell. This design helps to eliminate the effects of temperature and its influence on the balance's performance. | The sensor truly stands out with its unparalleled design based on the electromagnetic force compensation (EMFC). It provides an exceptional performance that meets the highest of the set international standards. | Motorised, built-in weight internal calibration ensures precise weighing under all conditions. The Perfect Self Automatic Calibration (PSAC) function automatically maintains accuracy by adjusting for time and temperature changes. | Our weighing system sets a new benchmark for mass measurement quality with the best weighing accuracy and sd ≤ 1d repeatability. We adhere to strict USP requirements (chapters 41 and 1251) to ensure the highest level of quality in our measurements. | The draft shield has wide opening, clear visibility and Infrared sensors to open and close the draft-shield doors with the wave of a hand. It simplifies dosing, minimizes spills, and reduces cross-contamination. | Excellent readable display with bright LED-backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | The built-in real-time clock ensures accurate, time-stamped data, providing traceable documentation that reflects true test results and supports compliance with SOPs and GLP regulations. | The USB and RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data. | An easy-to-view circular bubble level and large adjustable feet make balance leveling quick, simple, and precise for stable, accurate weighing. | Washable parts dismantle in seconds without tools. Dishwasher-safe components and a smooth housing design make routine cleaning quick, simple, and hygienic for everyday laboratory use.
| Specifications - Microbalance CM 2 | Specifications - Microbalance CM 5 | Specifications - Microbalance CM 6 | Specifications - Microbalance CM 11
Maximum Capacity | 2.1 g | 5.1g | 6.1 g | 11.1 g
Minimum load (OIML) | 0.1 mg | 0.1 mg | 0.1 mg | 0.1 mg
Readability | 1 µg | 1 µg | 1 µg | 1 µg
Verification unit [e] | 1 mg | 1 mg | 1 mg | 1 mg
Repeatability | 1 µg | 1 µg | 1 µg | 1.2 µg
Linearity | 3 µg | 6 µg | 6 µg | 6 µg
Tare Range (Subtractive) | -2.1 g | -5. 1g | -6.1 g | -11.1 g
Standard repeatability [5% Max] | 0.5 µg | 1.2 µg | 1.2 µg | 1.2 µg
Minimum Weight USP | 1 mg | 2.4 mg | 2.4 mg | 2.4 mg
Minimum weight ((U=1%, k=2) | 0.1 mg | 0.24 mg | 0.24 mg | 0.24 mg
Weighing Pan Diameter | Ø 30 mm | Ø 30 mm | Ø 30 mm | Ø 30 mm
Setting Time | 8-10 Sec | 8-10 Sec | 8-10 Sec | 8-10 Sec
Display | LCD Backlit | LCD Backlit | LCD Backlit | LCD Backlit
Calibration | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC)
Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
Operating temperature | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC
Relative humidity | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80%
Sensitivity Drift | ± 1ppm (1 x 10-6 / ºC) | ± 1ppm (1 x 10-6 / ºC) | ± 1ppm (1 x 10-6 / ºC) | ± 1ppm (1 x 10-6 / ºC)
Protection class | IP 43 | IP 43 | IP 43 | IP 43
Leveling system | manual | manual | manual | manual
Interfaces | RS232/USB | RS232/USB | RS232/USB | RS232/USB
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold
Power Supply | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A
Compliance | ISO, GLP / GMP & USP | ISO, GLP / GMP & USP | ISO, GLP / GMP & USP | ISO, GLP / GMP & USP
Weighing chamber height | Ø75 x 86.5 mm | Ø75 x 86.5 mm | Ø75 x 86.5 mm | Ø75 x 86.5 mm
Dimensions for Electronic Box (W x D x H) | 240 x 210 x 65 mm | 240 x 210 x 65 mm | 240 x 210 x 65 mm | 240 x 210 x 65 mm
Dimensions for Mechanical Box (W x D x H) | 360 x 140 x 146 mm | 360 x 140 x 146 mm | 360 x 140 x 146 mm | 360 x 140 x 146 mm
Packaging dimensions (W x D X H) | 510 x 350 x 540 mm | 510 x 350 x 540 mm | 510 x 350 x 540 mm | 510 x 350 x 540 mm
Net weight
Gross weight | 14.180 Kg | 14.180 Kg | 14.180 Kg | 14.180 Kg
Filter Microbalance CM 2F
High-Precision Microbalance – Solid construction, 2.1 g capacity, 1 µg readability, large backlit LCD, internal calibration, easy leveling, USB/RS232 connectivity, overload protection. Designed for environmental monitoring, filter microbalance measurements, and particulate matter (PM) analysis for pollution control.
| Weighing Capacity | 2.1 g
| Readability | 1 µg
| Minimum Weight (USP, 0.1%, typical) | 1 mg
Technology with Precise Weighing | Excellent Weighing Performance | Automatic Internal Calibration (PSAC) | Revolution in Mass Measurement Quantity | Simplified Operations and Protection | Bright, High-Contrast Display | 10 Easy-to-Use Built-In Applications | ISO/GLP Documentation | Seamless Connectivity | Quick & Simple Leveling
The technology has unparalleled design based on electromagnetic force compensation (EMFC). Its exceptional performance meets the highest international standards. The robust stainless-steel housing guarantees short stabilization time and eliminates heat transfer. | Filter microbalances that are perfect for measuring particulate emissions. CM-F Series Micro Balances have excellent performance in challenging conditions. They have a wide range of capacities from 2g up to 11g and a readability of 1µg, making them an excellent choice for accurate and precise weighing to measure the smallest sample weights. | Motorised, built-in weight internal calibration ensures precise weighing under all conditions. The Perfect Self Automatic Calibration (PSAC) function automatically maintains accuracy by adjusting for time and temperature changes. | Our weighing system sets a new benchmark for mass measurement quality with the best weighing accuracy and sd ≤ 1d repeatability. We adhere to strict USP requirements (chapters 41 and 1251) to ensure the highest level of quality in our measurements. | Equiped with weighing chamber of cylindrical shaped aluminium draft shield for filter weighing minimizing influence of electrostatic charges on weighing result. A built-in capacity tracker to prevent overloading the weighing cell. | Excellent readable display with bright LED-backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | The built-in real-time clock ensures accurate, time-stamped data, providing traceable documentation that reflects true test results and supports compliance with SOPs and GLP regulations. | The USB and RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data. | An easy-to-view circular bubble level and large adjustable feet make balance leveling quick, simple, and precise for stable, accurate weighing.
| Specifications - Microbalance CM 2F | Specifications - Microbalance CM 5F | Specifications - Microbalance CM 6F | Specifications - Microbalance CM 11F
Maximum Capacity | 2.1 g | 5 g | 6.1 g | 11.1 g
Minimum load (OIML) | 0.1 mg | 0.1 mg | 0.1 mg | 0.1 mg
Readability | 1 µg | 1 µg | 1 µg | 1 µg
Verification unit [e] | 1 mg | 1 mg | 1 mg | 1 mg
Repeatability | 1 µg | 1 µg | 1 µg | 1.2 µg
Linearity | 3 µg | 6 µg | 6 µg | 6 µg
Tare Range (Subtractive) | -2.1 g | -5 g | -6.1 g | -11.1 g
Standard repeatability [5% Max] | 0.5 µg | 1.2 µg | 1.2 µg | 1.2 µg
Minimum Weight USP | 1 mg | 2.4 mg | 2.4 mg | 2.4 mg
Minimum weight ((U=1%, k=2) | 0.1 mg | 0.24 mg | 0.24 mg | 0.24 mg
Weighing Pan Diameter | 50 Ø | 50 Ø | 50 Ø | 50 Ø
Setting Time | 8-10 Sec | 8-10 Sec | 8-10 Sec | 8-10 Sec
Display | Alpha Numeric Backlit LCD Display | Alpha Numeric Backlit LCD Display | Alpha Numeric Backlit LCD Display | Alpha Numeric Backlit LCD Display
Calibration | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC)
Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
Operating temperature | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC
Relative humidity | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80%
Sensitivity Temperature Drift | ± 1ppm (1 x 10-6 / ºC) | ± 1ppm (1 x 10-6 / ºC) | ± 1ppm (1 x 10-6 / ºC) | ± 1ppm (1 x 10-6 / ºC)
Protection class | IP 43 | IP 43 | IP 43 | IP 43
Leveling system | manual | manual | manual | manual
Interfaces | RS232/USB | RS232/USB | RS232/USB | RS232/USB
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold
Power Supply | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A
Compliance | ISO, GLP / GMP & USP | ISO, GLP / GMP & USP | ISO, GLP / GMP & USP | ISO, GLP / GMP & USP
Weighing chamber height
Dimensions for Electronic Box (W x D x H) | 240 x 210 x 65 mm | 240 x 210 x 65 mm | 240 x 210 x 65 mm | 240 x 210 x 65 mm
Dimensions for Mechanical Box (W x D x H)
Packaging dimensions (W x D X H) | 510 x 350 x 540 mm | 510 x 350 x 540 mm | 510 x 350 x 540 mm | 510 x 350 x 540 mm
Net weight
Gross weight | 14.180 Kg | 14.180 Kg | 14.180 Kg | 14.180 Kg
Pipettes Microbalance CM 11P
High-Precision Microbalance – Solid construction, 11 g capacity, 1 µg readability, large backlit LCD, internal calibration, easy leveling, USB/RS232 connectivity, overload protection. Ideal for precise pipette calibration in laboratories.
| Weighing Capacity | 11.1 g
| Readability | 1 µg
| Minimum Weight (USP, 0.1%, typical) | 2.4 mg
The Brilliant Proven Technology | Environmental and Internal Calibration | Revolution in Mass Measurement Quantity | Overload Protection and Touch-Free Shield | Well Built and Wider Viewing Display | Optional Evaporation Trap Kit | 10 Easy-to-use Built-in-Applications | Communication and GLP Documentation | Easy Levelling and Cleaning
The technology sensor truly stands out with its unparalleled design based on electromagnetic force compensation (EMFC). Its exceptional performance meets the highest international standards. | The Perfect Self Automatic Calibration function maintains accuracy by automatically calibrating the system based on time or temperature changes. Motorized built-in weight internal calibration ensures precise weighing performance in any condition. | Weighing system sets a new benchmark for mass measurement quality with the best accuracy and repeatability. Strict USP requirements (chapters 41 and 1251) ensure the highest level of quality in measurements. | Capacity tracker to prevent overloading the weighing cell. A wide opening, clear visibility and Infrared sensors enable you to open\close draft shield with just wave of a hand which simplifies your operations. | The solid construction along with high-quality materials provide protective cover to safeguard the keypad. A bright LED backlit LCD display with large numbers and wider viewing angle with an Intuitive interface. | The CET 01 evaporation traps stabilize the humidity condition inside the weighing chamber to minimize the evaporation of the liquid that causes errors. A weighing chamber made of glass minimizes influence of electrostatic charges on the result. | The ten easy to use build in applications including density determination, check weighing, totalling, and formulation, simplify your routine weighing tasks. | Standard USB and RS232 interface provides windows direct communication to all Windows applications. Built-in real-time clock, the Data is an accurate reflection of results in compliance with SOPS and GLP regulations. | Adjust the feet of the balance until indicator shows that air bubble is centred. The draft shields can be easily dismantled and assembled making it convenient to keep the balance clean.
Specifications - Microbalance Pipette CM 11P
Maximum Capacity | 11.1 g
Minimum load (OIML) | 0.1 mg
Readability | 1 µg
Verification unit [e] | 1 mg
Repeatability | 1.2 µg
Linearity | 6 µg
Tare Range (Subtractive) | -11.1 g
Standard repeatability [5% Max] | 1.2 µg
Minimum Weight USP | 2.4 mg
Minimum weight ((U=1%, k=2) | 0.24 mg
Weighing Pan Diameter | 20 Ø
Setting Time | 8-10 Sec
Display | Alpha Numeric Backlit LCD Display
Calibration | Motorised Internal Calibration (automatic/PSAC)
Units of measure | gm, mg, ct, GN, mo, oz, dwt, t1T, t1H, t1S, mom, Bat, MS
Operating temperature | 18º to 30ºC
Relative humidity | 40% ÷ 80%
Sensitivity Drift | ± 1ppm (1 x 10-6/ ºC)
Protection class | IP 43
Leveling system | manual
Interfaces | RS232/USB
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold
Power Supply | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A
Compliance | ISO, GLP / GMP & USP
Weighing chamber height | Ø75 x 86.5 mm
Dimensions for Electronic Box (W x D x H) | 240 x 210 x 65 mm
Dimensions for Mechanical Box (W x D x H) | 360 x 140 x 146 mm
Packaging dimensions (W x D X H) | 510 x 350 x 540 mm
Net weight
Gross weight | 14.180 Kg

## Analytical - Semi Micro
Analytical Balance CY 285C
Intuitive to operate with Solid built construction. 82/220 g capacity; 0.01/0.1 mg readability; large, bright LCD backlit display; internal Calibration; easy leveling; USB and RS232 connection; overload protection.
| Weighing Capacity | 82/220 g
| Readability | 0.01/0.1 mg
| Minimum Weight (USP, 0.1%, typical) | 30 mg
Brilliant Proven Technology | Automatic Internal Calibration (PSAC) | Revolution in Mass Measurement Quantity | New Smart Grid Pan | Bright, High-Contrast Display | Automatic Internal Calibration (PSAC) | 10 Easy-to-Use Built-In Applications | ISO/GLP Documentation | Seamless Connectivity | Easy Maintenance & Cleaning | Quick & Simple Leveling
The sensor really stands out because of its unique design based on electromagnetic force compensation (EMFC). It works very well and meets the highest international standards. | Motorised, built-in weight internal calibration ensures precise weighing under all conditions. The Perfect Self Automatic Calibration (PSAC) function automatically maintains accuracy by adjusting for time and temperature changes. | Our weighing system sets a new standard for the quality of mass measurement by meeting strict USP requirements (chapters 41 and 1251) for the best weighing accuracy and sd ≤ 1d repeatability. | The Smart Grid design lets air flow through it and keeps air currents from affecting the weighing process as much, which cuts down on problems caused by moving air. It also has a large-grade 304 stainless steel pan that doesn't attract magnets. | Excellent readable display with bright LED-backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | Motorised, built-in weight internal calibration ensures precise weighing under all conditions. The Perfect Self Automatic Calibration (PSAC) function automatically maintains accuracy by adjusting for time and temperature changes. | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | The built-in real-time clock ensures accurate, time-stamped data, providing traceable documentation that reflects true test results and supports compliance with SOPs and GLP regulations. | The USB and RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data. | Washable parts dismantle in seconds without tools. Dishwasher-safe components and a smooth housing design make routine cleaning quick, simple, and hygienic for everyday laboratory use. | An easy-to-view circular bubble level and large adjustable feet make balance leveling quick, simple, and precise for stable, accurate weighing.
| Specifications - Analytical Balance CY 285C | Specifications - Analytical Balance CY 205C
Maximum Capacity | 82 / 220 g | 210 g
Minimum load (OIML) | 1 mg | 1 mg
Readability | 0.01 / 0.1 mg | 0.01 mg
Verification unit [e] | 1 mg | 1 mg
Repeatability | 0.02 / 0.1 mg | 0.03 mg
Linearity | 0.05 / 0.2 mg | 0.07 mg
Tare Range (Subtractive) | -220 g | -210 g
Standard repeatability [5% Max] | 0.015 mg | 0.015 mg
Minimum Weight USP | 30 mg | 30 mg
Minimum weight ((U=1%, k=2) | 3 mg | 3 mg
Weighing Pan Diameter | Ø 80 mm Grid pan | Ø 80 mm Grid pan
Setting Time | 5- 6 Sec | 5- 6 Sec
Display | LCD Backlit | LCD Backlit
Calibration | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC)
Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
Operating temperature | 15º to 35ºC | 15º to 35ºC
Relative humidity | 40% ÷ 80% | 40% ÷ 80%
Sensitivity Drift | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC)
Protection class | IP 43 | IP 43
Leveling system | manual | manual
Interfaces | RS232/USB | RS232/USB
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density
Power Supply | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A
Compliance | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP
Balance Dimension (W x D X H) | 520 x 430 x 540mm | 520 x 430 x 540mm
Packaging dimensions (W x D X H)
Net weight | 8 kg | 8 kg
Gross weight | 12 kg | 12 kg
Note on B22: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 90 mm
Note on C22: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 90 mm grid pan

## Analytical Balances
Analytical Balance CY 224C
Intuitive to operate with solid built construction. 220 g capacity; 0.1 mg readability; large, bright LCD backlit display; internal Calibration; easy leveling; USB and RS232 connection. overload protection; small footprint.
| Weighing Capacity | 220 g
| Readability | 0.1 mg
| Minimum Weight (USP, 0.1%, typical) | 160 mg
Brilliant Proven Technology | Solid & Reliable Design | Bright, High-Contrast Display | Automatic Internal Calibration (PSAC) | Seamless Connectivity | 10 Easy-to-Use Built-In Applications | ISO/GLP Documentation | Quick & Simple Leveling | Easy Maintenance & Cleaning
The sensor truly stands out with its unparalleled design based on the electromagnetic force compensation (EMFC). It provides an exceptional performance that meets the highest of the set international standards. | Built with a solid design and high-quality materials for long-term durability, ensuring reliable and consistent operation with accurate results. Supplied with a standard protective cover for added safety. | Excellent readable display with bright LED-backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | Motorised, built-in weight internal calibration ensures precise weighing under all conditions. The Perfect Self Automatic Calibration (PSAC) function automatically maintains accuracy by adjusting for time and temperature changes. | The USB and RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data. | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | The built-in real-time clock ensures accurate, time-stamped data, providing traceable documentation that reflects true test results and supports compliance with SOPs and GLP regulations. | An easy-to-view circular bubble level and large adjustable feet make balance leveling quick, simple, and precise for stable, accurate weighing. | Washable parts dismantle in seconds without tools. Dishwasher-safe components and a smooth housing design make routine cleaning quick, simple, and hygienic for everyday laboratory use.
| Specifications - Analytical Balance CY 64C | Specifications - Analytical Balance CY 124C | Specifications - Analytical Balance CY 224C | Specifications - Analytical Balance CY 314C
Maximum Capacity | 60 g | 120 g | 220 g | 310 g
Minimum load (OIML) | 10 mg | 10 mg | 10 mg | 10 mg
Readability | 0.1 mg | 0.1 mg | 0.1 mg | 0.1 mg
Verification unit [e] | 1 mg | 1 mg | 1 mg | 1 mg
Repeatability | 0.1 mg | 0.1 mg | 0.1 mg | 0.2 mg
Linearity | 0.2 mg | 0.2 mg | 0.2 mg | 0.3 mg
Tare Range (Subtractive) | -60 g | -120 g | -220 g | -310 g
Standard repeatability [5% Max] | 0.08 mg | 0.08 mg | 0.08 mg | 0.08 mg
Minimum Weight USP | 160 mg | 160 mg | 160 mg | 160 mg
Minimum weight ((U=1%, k=2) | 16 mg | 16 mg | 16 mg | 16 mg
Weighing Pan Diameter | Ø 90 mm | Ø 90 mm | Ø 90 mm | Ø 90 mm
Setting Time | 2- 3 Sec | 2- 3 Sec | 2- 3 Sec | 2- 3 Sec
Display | LCD Backlit | LCD Backlit | LCD Backlit | LCD Backlit
Calibration | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC)
Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
Operating temperature | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC
Relative humidity | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80%
Sensitivity Drift | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC)
Protection class | IP 43 | IP 43 | IP 43 | IP 43
Leveling system | manual | manual | manual | manual
Interfaces | RS232/USB | RS232/USB | RS232/USB | RS232/USB
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold
Power Supply | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A
Compliance | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP
Balance Dimension (W x D X H) | 342.5 x 212 x 350 mm | 342.5 x 212 x 350 mm | 342.5 x 212 x 350 mm | 342.5 x 212 x 350 mm
Packaging dimensions (W x D X H) | 500 x 350 x 540 mm | 500 x 350 x 540 mm | 500 x 350 x 540 mm | 500 x 350 x 540 mm
Net weight | 8 kg | 8 kg | 8 kg | 8 kg
Gross weight | 12 kg | 12 kg | 12 kg | 12 kg
Analytical Balance CY 224
Intuitive to operate with solid-built construction. 220 g capacity; 0.1 mg readability; large, bright LCD backlit display; external calibration; easy leveling; USB and RS232 connection; overload protection; small footprint.
| Weighing Capacity | 220 g
| Readability | 0.1 mg
| Minimum Weight (USP, 0.1%, typical) | 160 mg
Brilliant Proven Technology | Solid & Reliable Design | Bright, High-Contrast Display | Seamless Connectivity | 10 Easy-to-Use Built-In Applications | ISO/GLP Documentation | Quick & Simple Leveling | Easy Maintenance & Cleaning
The sensor truly stands out with its unparalleled design based on the electromagnetic force compensation (EMFC). It provides an exceptional performance that meets the highest of the set international standards. | Built with a solid design and high-quality materials for long-term durability, ensuring reliable and consistent operation with accurate results. Supplied with a standard protective cover for added safety. | Excellent readable display with bright LED-backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | The USB and RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data. | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | The built-in real-time clock ensures accurate, time-stamped data, providing traceable documentation that reflects true test results and supports compliance with SOPs and GLP regulations. | An easy-to-view circular bubble level and large adjustable feet make balance leveling quick, simple, and precise for stable, accurate weighing. | Washable parts dismantle in seconds without tools. Dishwasher-safe components and a smooth housing design make routine cleaning quick, simple, and hygienic for everyday laboratory use.
| Specifications - Analytical Balance CY 64 | Specifications - Analytical Balance CY 124 | Specifications - Analytical Balance CY 224
Maximum Capacity | 60 g | 120 g | 220 g
Minimum load (OIML) | 10 mg | 10 mg | 10 mg
Readability | 0.1 mg | 0.1 mg | 0.1 mg
Verification unit [e] | 1 mg | 1 mg | 1 mg
Repeatability | 0.1 mg | 0.1 mg | 0.1 mg
Linearity | 0.2 mg | 0.2 mg | 0.2 mg
Tare Range (Subtractive) | -60 g | -120 g | -220 g
Standard repeatability [5% Max] | 0.08 mg | 0.08 mg | 0.08 mg
Minimum Weight USP | 160 mg | 160 mg | 160 mg
Minimum weight ((U=1%, k=2) | 16 mg | 16 mg | 16 mg
Weighing Pan Diameter | Ø 90 mm | Ø 90 mm | Ø 90 mm
Setting Time | 2- 3 Sec | 2- 3 Sec | 2- 3 Sec
Display | LCD Backlit | LCD Backlit | LCD Backlit
Calibration | External Calibration | External Calibration | External Calibration
Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
Operating temperature | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC
Relative humidity | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80%
Sensitivity Drift | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC)
Protection class | IP 43 | IP 43 | IP 43
Leveling system | manual | manual | manual
Interfaces | RS232/USB | RS232/USB | RS232/USB
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold
Power Supply | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A
Compliance | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP
Balance Dimension (W x D X H) | 342.5 x 212 x 350 mm | 342.5 x 212 x 350 mm | 342.5 x 212 x 350 mm
Packaging dimensions (W x D X H) | 500 x 350 x 540 mm | 500 x 350 x 540 mm | 500 x 350 x 540 mm
Net weight | 8 kg | 8 kg | 8 kg
Gross weight | 12 kg | 12 kg | 12 kg
Analytical Balance ALT 224C
New generation technology with Solid built Construction. 220 g capacity; 0.1 mg readability; Colour TFT touchscreen display; internal Calibration; optimal feature leveling alert, easy leveling;and user management simplify and optimize your daily weighing processes. USB and RS232 connection; overload protection; small footprint; Power saving mode.
| Weighing Capacity | 220 g
| Readability | 0.1 mg
| Minimum Weight (USP, 0.1%, typical) | 160 mg
Advanced EMFC Technology for Trusted Results | Intuitive Touchscreen | Automatic Internal Adjustment with PSAC | Designed for Easy Handling and Access | Advanced User Management & Access Control | Infrared Proximity Sensor | ALIBI Memory secure storage of measurements | 10 Easy-to-Use Built-In Applications | Easy Leveling for Accurate Results | Optional On-Screen Leveling Assistance | Audit Trial Report | Seamless Connectivity
Advanced EMFC sensor technology delivers international-standard performance. Engineered in India, ALT Series balances provide long-term accuracy, fast stabilization, excellent repeatability, overload protection, self-test reliability, and consistent results. | The large 4.3” color TFT touchscreen ensures clear visibility and smooth operation, even while wearing gloves. Its intuitive interface and wipe-clean surfaces allow quick, hygienic cleaning after every weighing task. | Automatic motorized, built-in weight internal adjustment ensures continuous accuracy using internal test weights. The balance performs PSAC at scheduled times or automatically whenever ambient temperature changes are detected. | The spacious weighing chamber provides extra height for large containers, convenient access to the weighing pan, smooth, removable top and side sliding doors, and a fully detachable draft shield for easy handling and cleaning. | User Management ensures secure operation by allowing administrators to create up to 7 users and assign specific access rights, ensuring controlled usage, data integrity, and protection against unauthorized actions. | Equipped with two programmable hands-free sensors, the balance allows convenient contactless operation. Each sensor can be assigned functions such as Zero, Tare, Print, Home Screen, or user-defined tasks, improving workflow efficiency and maintaining hygiene. | ALIBI memory securely stores up to 1000 reports. ensuring reliable data storage, full traceability, and long-term safety. It guarantees continuity and protection of vital measurement data over extended periods. USB-A (Optional for mass data storage in a pendrive ) | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | ALT Series balances ensure maximum stability with adjustable feet and a front bubble level, delivering perfectly leveled, accurate, and reliable weighing results. | Optional real-time digital level guidance uses an integrated sensor to continuously monitor the balance position and provides on-screen instructions, enabling quick and accurate leveling within seconds. | AuditTrail adds value by providing complete, gapless documentation with automatic recording of calibration and leveling processes, including secure date and time stamps for full traceability. | The RS232, USB-C, and Ethernet interfaces allow easy connection to a printer, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data.
| Specifications - Analytical Balance ALT 64C | Specifications - Analytical Balance ALT 124C | Specifications - Analytical Balance ALT 224C | Specifications - Analytical Balance ALT 324C | Specifications - Analytical Balance ALT 285C
Maximum Capacity | 60 g | 120 g | 220 g | 320 g | 82 / 220 g
Minimum load (OIML) | 10 mg | 10 mg | 10 mg | 10 mg | 1 mg
Readability | 0.1 mg | 0.1 mg | 0.1 mg | 0.1 mg | 0.01 / 0.1 mg
Verification unit [e] | 1 mg | 1 mg | 1 mg | 1 mg | 1 mg
Repeatability | 0.1 mg | 0.1 mg | 0.1 mg | 0.2 mg | 0.02 / 0.1 mg
Linearity | 0.2 mg | 0.2 mg | 0.2 mg | 0.3 mg | 0.04 / 0.2 mg
Tare Range (Subtractive) | -60 g | -120 g | -220 g | -320 g | -220 g
Standard repeatability [5% Max] | 0.08 mg | 0.08 mg | 0.08 mg | 0.08 mg | 0.015 mg
Minimum Weight USP | 160 mg | 160 mg | 160 mg | 160 mg | 30 mg
Minimum weight ((U=1%, k=2) | 16 mg | 16 mg | 16 mg | 16 mg | 3 mg
Weighing Pan Diameter | Ø 90 mm | Ø 90 mm | Ø 90 mm | Ø 90 mm | Ø 90 mm Grid pan
Setting Time | 2- 3 Sec | 2- 3 Sec | 2- 3 Sec | 2- 3 Sec | 5- 6 Sec
Display | 4.3" color TFT touchscreen | 4.3" color TFT touchscreen | 4.3" color TFT touchscreen | 4.3" color TFT touchscreen | 4.3" color TFT touchscreen
Calibration | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC)
Features | 7 User management Password protection AuditTrai ALIBI Memory ( 1000 reports are stored. ) Minimum Sample Weight (MSW) Touch-free operation Infrared Proximity Sensor 4 English, German, French, Spanish built-in languages A Levelling guidance ( Optional ) USB-A (Optional for mass data storage in a pendrive ) | 7 User management Password protection AuditTrai ALIBI Memory ( 1000 reports are stored. ) Minimum Sample Weight (MSW) Touch-free operation Infrared Proximity Sensor 4 English, German, French, Spanish built-in languages A Levelling guidance ( Optional ) USB-A (Optional for mass data storage in a pendrive ) | 7 User management Password protection AuditTrai ALIBI Memory ( 1000 reports are stored. ) Minimum Sample Weight (MSW) Touch-free operation Infrared Proximity Sensor 4 English, German, French, Spanish built-in languages A Levelling guidance ( Optional ) USB-A (Optional for mass data storage in a pendrive ) | 7 User management Password protection AuditTrai ALIBI Memory ( 1000 reports are stored. ) Minimum Sample Weight (MSW) Touch-free operation Infrared Proximity Sensor 4 English, German, French, Spanish built-in languages A Levelling guidance ( Optional ) USB-A (Optional for mass data storage in a pendrive ) | 7 User management Password protection AuditTrai ALIBI Memory ( 1000 reports are stored. ) Minimum Sample Weight (MSW) Touch-free operation Infrared Proximity Sensor 4 English, German, French, Spanish built-in languages A Levelling guidance ( Optional ) USB-A (Optional for mass data storage in a pendrive )
Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
Operating temperature | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC
Relative humidity | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80%
Sensitivity Drift | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC)
Protection class | IP 43 | IP 43 | IP 43 | IP 43 | IP 43
Leveling system | manual | manual | manual | manual | manual
Interface | RS232/USB | RS232/USB | RS232/USB | RS232/USB | RS232/USB
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic weighting Check weighting Formulation Totalization Density Statistics Peak Hold
Power Supply | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz. | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz. | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz. | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz. | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz.
Compliance | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP
Balance Dimension (W x D X H) | 211 x 310 x 315 mm | 211 x 310 x 315 mm | 211 x 310 x 315 mm | 211 x 310 x 315 mm | 211 x 310 x 315 mm
Packaging dimensions (W x D X H) | 350 x 510 x 540 mm | 350 x 510 x 540 mm | 350 x 510 x 540 mm | 350 x 510 x 540 mm
Net weight | 5 kg | 5 kg | 5 kg | 5 kg | 8 kg
Gross weight | 8.4 kg | 8.4 kg | 8.4 kg | 8.4 kg | 12 kg
Analytical Balance ALT 224
New generation technology with Solid built Construction. 220 g capacity; 0.1 mg readability; Colour TFT touchscreen display; external Calibration; optonal feature leveling alert, easy leveling;and user management simplify and optimize your daily weighing processes. USB and RS232 connection; overload protection; small footprint; Power saving mode.
| Weighing Capacity | 220 g
| Readability | 0.1 mg
| Minimum Weight (USP, 0.1%, typical) | 160 mg
Advanced EMFC Technology for Trusted Results | Intuitive Touchscreen | Designed for Easy Handling and Access | Advanced User Management & Access Control | Infrared Proximity Sensor | ALIBI Memory secure storage of measurements | 10 Easy-to-Use Built-In Applications | Easy Leveling for Accurate Results | Optional On-Screen Leveling Assistance | Audit Trial Report | Seamless Connectivity
Advanced EMFC sensor technology delivers international-standard performance. Engineered in India, ALT Series balances provide long-term accuracy, fast stabilization, excellent repeatability, overload protection, self-test reliability, and consistent results. | The large 4.3” color TFT touchscreen ensures clear visibility and smooth operation, even while wearing gloves. Its intuitive interface and wipe-clean surfaces allow quick, hygienic cleaning after every weighing task. | The spacious weighing chamber provides extra height for large containers, convenient access to the weighing pan, smooth, removable top and side sliding doors, and a fully detachable draft shield for easy handling and cleaning. | User Management ensures secure operation by allowing administrators to create up to 7 users and assign specific access rights, ensuring controlled usage, data integrity, and protection against unauthorized actions. | Equipped with two programmable hands-free sensors, the balance allows convenient contactless operation. Each sensor can be assigned functions such as Zero, Tare, Print, Home Screen, or user-defined tasks, improving workflow efficiency and maintaining hygiene. | ALIBI memory securely stores up to 1000 reports. ensuring reliable data storage, full traceability, and long-term safety. It guarantees continuity and protection of vital measurement data over extended periods. USB-A (Optional for mass data storage in a pendrive ) | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | ALT Series balances ensure maximum stability with adjustable feet and a front bubble level, delivering perfectly leveled, accurate, and reliable weighing results. | Optional real-time digital level guidance uses an integrated sensor to continuously monitor the balance position and provides on-screen instructions, enabling quick and accurate leveling within seconds. | AuditTrail adds value by providing complete, gapless documentation with automatic recording of calibration and leveling processes, including secure date and time stamps for full traceability. | The RS232, USB-C, and Ethernet interfaces allow easy connection to a printer, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data.
| Specifications - Analytical Balance ALT 64 | Specifications - Analytical Balance ALT 124 | Specifications - Analytical Balance ALT 224
Maximum Capacity | 60 g | 120 g | 220 g
Minimum load (OIML) | 10 mg | 10 mg | 10 mg
Readability | 0.1 mg | 0.1 mg | 0.1 mg
Verification unit [e] | 1 mg | 1 mg | 1 mg
Repeatability | 0.1 mg | 0.1 mg | 0.1 mg
Linearity | 0.2 mg | 0.2 mg | 0.2 mg
Tare Range (Subtractive) | -60 g | -120 g | -220 g
Standard repeatability [5% Max] | 0.08 mg | 0.08 mg | 0.08 mg
Minimum Weight USP | 160 mg | 160 mg | 160 mg
Minimum weight ((U=1%, k=2) | 16 mg | 16 mg | 16 mg
Weighing Pan Diameter | Ø 90 mm | Ø 90 mm | Ø 90 mm
Setting Time | 2- 3 Sec | 2- 3 Sec | 2- 3 Sec
Display | 4.3" color TFT touchscreen | 4.3" color TFT touchscreen | 4.3" color TFT touchscreen
Calibration | Eternal Calibration | Eternal Calibration | Eternal Calibration
Features | 7 User management Password protection AuditTrai ALIBI Memory ( 1000 reports are stored. ) Minimum Sample Weight (MSW) Touch-free operation Infrared Proximity Sensor 4 built-in languages English, German, French, Spanish Levelling guidance ( Optional ) USB-A (Optional for mass data storage in a pendrive ) | 7 User management Password protection AuditTrai ALIBI Memory ( 1000 reports are stored. ) Minimum Sample Weight (MSW) Touch-free operation Infrared Proximity Sensor 4 English, German, French, Spanish built-in languages A Levelling guidance ( Optional ) USB-A (Optional for mass data storage in a pendrive ) | 7 User management Password protection AuditTrai ALIBI Memory ( 1000 reports are stored. ) Minimum Sample Weight (MSW) Touch-free operation Infrared Proximity Sensor 4 English, German, French, Spanish built-in languages A Levelling guidance ( Optional ) USB-A (Optional for mass data storage in a pendrive )
Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
Operating temperature | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC
Relative humidity | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80%
Sensitivity Drift | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC)
Protection class | IP 43 | IP 43 | IP 43
Leveling system | manual | manual | manual
Interfaces | RS232/USB | RS232/USB | RS232/USB
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold
Power Supply | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz. | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz. | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz.
Compliance | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP
Balance Dimension (W x D X H) | 211 x 310 x 315 mm | 211 x 310 x 315 mm | 211 x 310 x 315 mm
Packaging dimensions (W x D X H) | 350 x 510 x 540 mm | 350 x 510 x 540 mm | 350 x 510 x 540 mm
Net weight | 5 kg | 5 kg | 5 kg
Gross weight | 8.4 kg | 8.4 kg | 8.4 kg
Accessories
Data Printer CPR 02
Dot matrix printer, RS232 interface, print speed 2.3 lines per second,
Anti-Vibration Pad VRD 03
Anti-Vibration Pad for microbalance, analytical & precision balance.
Remote Display SRD 01
LCD backlit display powered by the balance; RJ45 interface sutable for CY series balances
Density Kit (Analytical) CDK O1
It allows for fast and precise density determination of solid, floating substances.
Density Kit (Analytical) CDK 01-L
It allows for fast and precise density determination liquid substances.
Weigh-Below Scales For Specific Gravity Stand LBA-X 13
A rigid frame (often mild steel) that supports the balance above a water tank. sutable for 0.01/0.1 / 1 / 10 mg balances

## Precision Balances
Precision Balance CY 223C
Intuitive to operate with solid built construction. 220 g capacity; 0.001 g readability; large, bright LCD backlit display; internal Calibration; easy leveling; USB and RS232 connection.
| Weighing Capacity | 220 g
| Readability | 0.001 g
| Minimum Weight (USP, 0.1%, typical) | 1 g
Brilliant Proven Technology | Solid & Reliable Design | Bright, High-Contrast Display | Automatic Internal Calibration (PSAC) | Seamless Connectivity | 10 Easy-to-Use Built-In Applications | ISO/GLP Documentation | Quick & Simple Leveling | Easy Maintenance & Cleaning
The sensor truly stands out with its unparalleled design based on the electromagnetic force compensation (EMFC). It provides an exceptional performance that meets the highest of the set international standards. | Built with a solid design and high-quality materials for long-term durability, ensuring reliable and consistent operation with accurate results. Supplied with a standard protective cover for added safety. | Excellent readable display with bright LED-backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | Motorised, built-in weight internal calibration ensures precise weighing under all conditions. The Perfect Self Automatic Calibration (PSAC) function automatically maintains accuracy by adjusting for time and temperature changes. | The USB and RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data. | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | The built-in real-time clock ensures accurate, time-stamped data, providing traceable documentation that reflects true test results and supports compliance with SOPs and GLP regulations. | An easy-to-view circular bubble level and large adjustable feet make balance leveling quick, simple, and precise for stable, accurate weighing. | Washable parts dismantle in seconds without tools. Dishwasher-safe components and a smooth housing design make routine cleaning quick, simple, and hygienic for everyday laboratory use.
| Specifications - Precision Balance CY 223C | Specifications - Precision Balance CY 323C | Specifications - Precision Balance CY 513C | Specifications - Precision Balance CY 723C | Specifications - Precision Balance CY 1003C | Specifications - Precision Balance CY 2003C | Specifications - Precision Balance CY 3003C
Maximum Capacity | 220 g | 320 g | 510 g | 720 g | 1000 g | 2000 g | 3000 g
Minimum load (OIML) | 20 mg | 20 mg | 20 mg | 20 mg | 20 mg | -- | --
Readability | 0.001 g | 0.001 g | 0.001 g | 0.001 g | 0.001 g | 0.001 g | 0.001 g
Verification unit [e] | 10 mg | 10 mg | 10 mg | 10 mg | 10 mg | -- | --
Repeatability | 0.001 g | 0.001 g | 0.001 g | 0.002 g | 0.002 g | 0.004 g | 0.005 g
Linearity | 0.002 g | 0.002 g | 0.002 g | 0.003 g | 0.003 g | 0.005 g | 0.006 g
Tare Range (Subtractive) | -220 g | -320 g | -510 g | -720 g | -1000 g | -2000 g | -3000 g
Standard repeatability [5% Max] | 0.5 mg | 0.5 mg | 0.5 mg | 0.5 mg | 0.5 mg | 0.6 mg | 0.6 mg
Minimum Weight USP | 1 g | 1 g | 1 g | 1 g | 1 g | 1.2 g | 1.2 g
Minimum weight ((U=1%, k=2) | 0.1 g | 0.1 g | 0.1 g | 0.1 g | 0.1 g | 0.12 g | 0.12 g
Weighing Pan | 128 X 128 mm | 128 X 128 mm | 128 X 128 mm | 128 X 128 mm | 128 X 128 mm | 128 X 128 mm | 128 X 128 mm
Setting Time | 2- 3 Sec | 2- 3 Sec | 2- 3 Sec | 2- 3 Sec | 2- 3 Sec | 2- 3 Sec | 2- 3 Sec
Display | LCD Backlit | LCD Backlit | LCD Backlit | LCD Backlit | LCD Backlit | LCD Backlit | LCD Backlit
Calibration | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC)
Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
Operating temperature | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC
Relative humidity | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80%
Sensitivity Drift | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC)
Protection class | IP 43 | IP 43 | IP 43 | IP 43 | IP 43 | IP 43 | IP 43
Leveling system | manual | manual | manual | manual | manual | manual | manual
Interfaces | RS-232/USB | RS-232/USB | RS-232/USB | RS-232/USB | RS-232/USB | RS-232/USB | RS-232/USB
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold
Power Supply | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A
Compliance | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP
Balance Dimension (W x D X H) | 342.5 x 212 x 165 mm | 342.5 x 212 x 165 mm | 342.5 x 212 x 165 mm | 342.5 x 212 x 165 mm | 342.5 x 212 x 165 mm | 342.5 x 212 x 165 mm | 342.5 x 212 x 165 mm
Packaging dimensions (W x D X H) | 510 x 410 x 330 mm | 510 x 410 x 330 mm | 510 x 410 x 330 mm | 510 x 410 x 330 mm | 510 x 410 x 330 mm | 510 x 410 x 330 mm | 510 x 410 x 330 mm
Net weight | 7.8 kg | 7.8 kg | 7.8 kg | 7.8 kg | 7.8 kg | 7.8 kg | 7.8 kg
Gross weight | 12 kg | 12 kg | 12 kg | 12 kg | 12 kg | 12 kg | 12 kg
Precision Balance CY 223
Intuitive to operate with solid built construction. 220 g capacity; 0.001 g readability; large, bright LCD backlit display; external Calibration; easy leveling; USB and RS232 connection.
| Weighing Capacity | 220 g
| Readability | 0.001 g
| Minimum Weight (USP, 0.1%, typical) | 1 g
Brilliant Proven Technology | Solid & Reliable Design | Bright, High-Contrast Display | Seamless Connectivity | 10 Easy-to-Use Built-In Applications | ISO/GLP Documentation | Quick & Simple Leveling | Easy Maintenance & Cleaning
The sensor truly stands out with its unparalleled design based on the electromagnetic force compensation (EMFC). It provides an exceptional performance that meets the highest of the set international standards. | Built with a solid design and high-quality materials for long-term durability, ensuring reliable and consistent operation with accurate results. Supplied with a standard protective cover for added safety. | Excellent readable display with bright LED-backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | The USB and RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data. | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | The built-in real-time clock ensures accurate, time-stamped data, providing traceable documentation that reflects true test results and supports compliance with SOPs and GLP regulations. | An easy-to-view circular bubble level and large adjustable feet make balance leveling quick, simple, and precise for stable, accurate weighing. | Washable parts dismantle in seconds without tools. Dishwasher-safe components and a smooth housing design make routine cleaning quick, simple, and hygienic for everyday laboratory use.
| Specifications - Precision Balance CY 223 | Specifications - Precision Balance CY 323 | Specifications - Precision Balance CY 513 | Specifications - Precision Balance CY 723 | Specifications - Precision Balance CY 1003
Maximum Capacity | 220 g | 320 g | 510 g | 720 g | 1000 g
Minimum load (OIML) | 20 mg | 20 mg | 20 mg | 20 mg | 20 mg
Readability | 0.001 g | 0.001 g | 0.001 g | 0.001 g | 0.001 g
Verification unit [e] | 10 mg | 10 mg | 10 mg | 10 mg | 10 mg
Repeatability | 0.001 g | 0.001 g | 0.001 g |  | 0.002 g
Linearity | 0.002 g | 0.002 g | 0.002 g | 0.003 g | 0.003 g
Tare Range (Subtractive) | -220 g | -320 g | -510 g | -720 g | -1000 g
Standard repeatability [5% Max] | 0.5 mg | 0.5 mg | 0.5 mg | 0.5 mg | 0.5 mg
Minimum Weight USP | 1 g | 1 g | 1 g | 1 g | 1 g
Minimum weight ((U=1%, k=2) | 0.1 g | 0.1 g | 0.1 g | 0.1 g | 0.1 g
Weighing Pan | 128 X 128 mm | 128 X 128 mm | 128 X 128 mm | 128 X 128 mm | 128 X 128 mm
Setting Time | 2- 3 Sec | 2- 3 Sec | 2- 3 Sec | 2- 3 Sec | 2- 3 Sec
Display | LCD Backlit | LCD Backlit | LCD Backlit | LCD Backlit | LCD Backlit
Calibration | External Calibration | External Calibration | External Calibration | External Calibration | External Calibration
Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
Operating temperature | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC
Relative humidity | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80%
Sensitivity Drift | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC)
Protection class | IP 43 | IP 43 | IP 43 | IP 43 | IP 43
Leveling system | manual | manual | manual | manual | manual
Interfaces | RS-232/USB | RS-232/USB | RS-232/USB | RS-232/USB | RS-232/USB
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold
Power Supply | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A
Compliance | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP
Balance Dimension (W x D X H) | 342.5 x 212 x 165 mm | 342.5 x 212 x 165 mm | 342.5 x 212 x 165 mm | 342.5 x 212 x 165 mm | 342.5 x 212 x 165 mm
Packaging dimensions (W x D X H) | 510 x 410 x 330 mm | 510 x 410 x 330 mm | 510 x 410 x 330 mm | 510 x 410 x 330 mm | 510 x 410 x 330 mm
Net weight | 7.8 kg | 7.8 kg | 7.8 kg | 7.8 kg | 7.8 kg
Gross weight | 12 kg | 12 kg | 12 kg | 12 kg | 12 kg
Precision Balance ALT 223C
New generation technology with Solid built Construction. 220 g capacity; 0.001 g readability; Colour TFT touchscreen display; internal Calibration; optimal feature leveling alert, easy leveling;and user management simplify and optimize your daily weighing processes. USB and RS232 connection; overload protection; small footprint; Power saving mode. .
| Weighing Capacity | 220 g
| Readability | 0.001 g
| Minimum Weight (USP, 0.1%, typical) | 1 g
Advanced EMFC Technology for Trusted Results | Intuitive Touchscreen | Automatic Internal Adjustment with PSAC | Designed for Easy Handling and Access | Advanced User Management & Access Control | Infrared Proximity Sensor | ALIBI Memory secure storage of measurements | 10 Easy-to-Use Built-In Applications | Easy Leveling for Accurate Results | Optional On-Screen Leveling Assistance | Audit Trial Report | Seamless Connectivity
Advanced EMFC sensor technology delivers international-standard performance. Engineered in India, ALT Series balances provide long-term accuracy, fast stabilization, excellent repeatability, overload protection, self-test reliability, and consistent results. | The large 4.3” color TFT touchscreen ensures clear visibility and smooth operation, even while wearing gloves. Its intuitive interface and wipe-clean surfaces allow quick, hygienic cleaning after every weighing task. | Automatic motorized, built-in weight internal adjustment ensures continuous accuracy using internal test weights. The balance performs PSAC at scheduled times or automatically whenever ambient temperature changes are detected. | The spacious weighing chamber provides extra height for large containers, convenient access to the weighing pan, smooth, removable top and side sliding doors, and a fully detachable draft shield for easy handling and cleaning. | User Management ensures secure operation by allowing administrators to create up to 7 users and assign specific access rights, ensuring controlled usage, data integrity, and protection against unauthorized actions. | Equipped with two programmable hands-free sensors, the balance allows convenient contactless operation. Each sensor can be assigned functions such as Zero, Tare, Print, Home Screen, or user-defined tasks, improving workflow efficiency and maintaining hygiene. | ALIBI memory securely stores up to 1000 reports. ensuring reliable data storage, full traceability, and long-term safety. It guarantees continuity and protection of vital measurement data over extended periods. USB-A (Optional for mass data storage in a pendrive ) | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | ALT Series balances ensure maximum stability with adjustable feet and a front bubble level, delivering perfectly leveled, accurate, and reliable weighing results. | Optional real-time digital level guidance uses an integrated sensor to continuously monitor the balance position and provides on-screen instructions, enabling quick and accurate leveling within seconds. | AuditTrail adds value by providing complete, gapless documentation with automatic recording of calibration and leveling processes, including secure date and time stamps for full traceability. | The RS232, USB-C, and Ethernet interfaces allow easy connection to a printer, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data.
| Specifications - Precision Balance ALT 223C | Specifications - Precision Balance ALT 513C | Specifications - Precision Balance ALT 1003C
Maximum Capacity | 320 g | 510 g | 1000 g
Minimum load (OIML) | 20 mg | 20 mg | 20 mg
Readability | 0.001 g | 0.001 g | 0.001 g
Verification unit [e] | 10 mg | 10 mg | 10 mg
Repeatability | 0.001 g | 0.001 g | 0.002 g
Linearity | 0.002 g | 0.002 g | 0.003 g
Tare Range (Subtractive) | -320 g | -510 g | -1000 g
Standard repeatability [5% Max] | 0.5 mg | 0.5 mg | 0.5 mg
Minimum Weight USP | 1 g | 1 g | 1 g
Minimum weight ((U=1%, k=2) | 0.1 g | 0.1 g | 0.1 g
Weighing Pan | 128 X 128 mm | 128 X 128 mm | 128 X 128 mm
Setting Time | 2- 3 Sec | 2- 3 Sec | 2- 3 Sec
Display | 4.3" color TFT touchscreen | 4.3" color TFT touchscreen | 4.3" color TFT touchscreen
Calibration | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC)
Features | 7 User management Password protection AuditTrai ALIBI Memory ( 1000 reports are stored. ) Minimum Sample Weight (MSW) Touch-free operation Infrared Proximity Sensor 4 English, German, French, Sp built-in languagesanish Levelling guidance ( Optional ) USB-A (Optional for mass data storage in a pendrive ) | 7 User management Password protection AuditTrai ALIBI Memory ( 1000 reports are stored. ) Minimum Sample Weight (MSW) Touch-free operation Infrared Proximity Sensor 4 built-in languages English, German, French, Spanish Levelling guidance ( Optional ) USB-A (Optional for mass data storage in a pendrive ) | 7 User management Password protection AuditTrai ALIBI Memory ( 1000 reports are stored. ) Minimum Sample Weight (MSW) Touch-free operation Infrared Proximity Sensor 4 built-in languages English, German, French, Spanish Levelling guidance ( Optional ) USB-A (Optional for mass data storage in a pendrive )
Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
Operating temperature | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC
Relative humidity | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80%
Sensitivity Drift | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC)
Protection class | IP 43 | IP 43 | IP 43
Leveling system | manual | manual | manual
Interfaces | RS232/USB | RS232/USB | RS232/USB
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold
Power Supply | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz. | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz. | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz.
Compliance | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP
Balance Dimension (W x D X H) | 211 x 310 x 315 mm | 211 x 310 x 315 mm | 211 x 310 x 315 mm
Packaging dimensions (W x D X H) | 410 x 520 x 340 mm | 410 x 520 x 340 mm | 410 x 520 x 340 mm
Net weight | 5 kg | 5 kg | 5 kg
Gross weight | 6.3 kg | 6.3 kg | 6.3 kg
Precision Balance ALT 223
New generation technology with Solid built Construction. 220 g capacity; 0.001 g readability; Colour TFT touchscreen display; external calibration; optimal feature leveling alert, easy leveling;and user management simplify and optimize your daily weighing processes. USB and RS232 connection; overload protection; small footprint; Power saving mode. .
| Weighing Capacity | 220 g
| Readability | 0.001 g
| Minimum Weight (USP, 0.1%, typical) | 1 g
Advanced EMFC Technology for Trusted Results | Intuitive Touchscreen | Designed for Easy Handling and Access | Advanced User Management & Access Control | Infrared Proximity Sensor | ALIBI Memory secure storage of measurements | 10 Easy-to-Use Built-In Applications | Easy Leveling for Accurate Results | Optional On-Screen Leveling Assistance | Audit Trial Report | Seamless Connectivity
Advanced EMFC sensor technology delivers international-standard performance. Engineered in India, ALT Series balances provide long-term accuracy, fast stabilization, excellent repeatability, overload protection, self-test reliability, and consistent results. | The large 4.3” color TFT touchscreen ensures clear visibility and smooth operation, even while wearing gloves. Its intuitive interface and wipe-clean surfaces allow quick, hygienic cleaning after every weighing task. | The spacious weighing chamber provides extra height for large containers, convenient access to the weighing pan, smooth, removable top and side sliding doors, and a fully detachable draft shield for easy handling and cleaning. | User Management ensures secure operation by allowing administrators to create up to 7 users and assign specific access rights, ensuring controlled usage, data integrity, and protection against unauthorized actions. | Equipped with two programmable hands-free sensors, the balance allows convenient contactless operation. Each sensor can be assigned functions such as Zero, Tare, Print, Home Screen, or user-defined tasks, improving workflow efficiency and maintaining hygiene. | ALIBI memory securely stores up to 1000 reports. ensuring reliable data storage, full traceability, and long-term safety. It guarantees continuity and protection of vital measurement data over extended periods. USB-A (Optional for mass data storage in a pendrive ) | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | ALT Series balances ensure maximum stability with adjustable feet and a front bubble level, delivering perfectly leveled, accurate, and reliable weighing results. | Optional real-time digital level guidance uses an integrated sensor to continuously monitor the balance position and provides on-screen instructions, enabling quick and accurate leveling within seconds. | AuditTrail adds value by providing complete, gapless documentation with automatic recording of calibration and leveling processes, including secure date and time stamps for full traceability. | The RS232, USB-C, and Ethernet interfaces allow easy connection to a printer, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data.
| Specifications - Precision Balance ALT 223 | Specifications - Precision Balance ALT 513 | Specifications - Precision Balance ALT 1003
Maximum Capacity | 320 g | 510 g | 1000 g
Minimum load (OIML) | 20 mg | 20 mg | 20 mg
Readability | 0.001 g | 0.001 g | 0.001 g
Verification unit [e] | 10 mg | 10 mg | 10 mg
Repeatability | 0.001 g | 0.001 g | 0.002 g
Linearity | 0.002 g | 0.002 g | 0.003 g
Tare Range (Subtractive) | -320 g | -510 g | -1000 g
Standard repeatability [5% Max] | 0.5 mg | 0.5 mg | 0.5 mg
Minimum Weight USP | 1 g | 1 g | 1 g
Minimum weight ((U=1%, k=2) | 0.1 g | 0.1 g | 0.1 g
Weighing Pan | 128 X 128 mm | 128 X 128 mm | 128 X 128 mm
Setting Time | 2- 3 Sec | 2- 3 Sec | 2- 3 Sec
Display | 4.3" color TFT touchscreen | 4.3" color TFT touchscreen | 4.3" color TFT touchscreen
Calibration | External Calibration | External Calibration | External Calibration
Features | 7 User management Password protection AuditTrai ALIBI Memory ( 1000 reports are stored. ) Minimum Sample Weight (MSW) Touch-free operation Infrared Proximity Sensor 4 built-in languages English, German, French, Spanish Levelling guidance ( Optional ) USB-A (Optional for mass data storage in a pendrive ) | 7 User management Password protection AuditTrai ALIBI Memory ( 1000 reports are stored. ) Minimum Sample Weight (MSW) Touch-free operation Infrared Proximity Sensor 4 built-in languages English, German, French, Spanish Levelling guidance ( Optional ) USB-A (Optional for mass data storage in a pendrive ) | 7 User management Password protection AuditTrai ALIBI Memory ( 1000 reports are stored. ) Minimum Sample Weight (MSW) Touch-free operation Infrared Proximity Sensor 4 built-in languages English, German, French, Spanish Levelling guidance ( Optional ) USB-A (Optional for mass data storage in a pendrive )
Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
Operating temperature | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC
Relative humidity | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80%
Sensitivity Drift | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC)
Protection class | IP 43 | IP 43 | IP 43
Leveling system | manual | manual | manual
Interfaces | RS232/USB | RS232/USB | RS232/USB
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold
Power Supply | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz. | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz. | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz.
Compliance | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP
Balance Dimension (W x D X H) | 211 x 310 x 315 mm | 211 x 310 x 315 mm | 211 x 310 x 315 mm
Packaging dimensions (W x D X H) | 410 x 520 x 340 mm | 410 x 520 x 340 mm | 410 x 520 x 340 mm
Net weight | 5 kg | 5 kg | 5 kg
Gross weight | 6.3 kg | 6.3 kg | 6.3 kg
Precision Balance CY 3102C
Intuitive to operate with solid built construction. 3100 g capacity; 0.01 g readability; large, bright LCD backlit display; internal Calibration; easy leveling; USB and RS232 connection.
| Weighing Capacity | 3100 g
| Readability | 0.01 g
| Minimum Weight (USP, 0.1%, typical) | 10 g
Brilliant Proven Technology | Solid & Reliable Design | Bright, High-Contrast Display | Automatic Internal Calibration (PSAC) | Seamless Connectivity | 10 Easy-to-Use Built-In Applications | ISO/GLP Documentation | Quick & Simple Leveling | Easy Maintenance & Cleaning
The Precision Balances use a MonoTech single-block weighing sensor ensuring uniform load transmission, thermal stability, rapid stabilization, accurate corner-load performance, high repeatability, minimal drift, and long-term reliability in continuous operation. | Built with a solid design and high-quality materials for long-term durability, ensuring reliable and consistent operation with accurate results. Supplied with a standard protective cover for added safety. | Excellent readable display with bright LED-backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | Motorised, built-in weight internal calibration ensures precise weighing under all conditions. The Perfect Self Automatic Calibration (PSAC) function automatically maintains accuracy by adjusting for time and temperature changes. | The USB and RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data. | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | The built-in real-time clock ensures accurate, time-stamped data, providing traceable documentation that reflects true test results and supports compliance with SOPs and GLP regulations. | An easy-to-view circular bubble level and large adjustable feet make balance leveling quick, simple, and precise for stable, accurate weighing. | Washable parts dismantle in seconds without tools. Dishwasher-safe components and a smooth housing design make routine cleaning quick, simple, and hygienic for everyday laboratory use.
| Specifications - Precision Balance CY 1202C | Specifications - Precision Balance CY 2202C | Specifications - Precision Balance CY 3102C | Specifications - Precision Balance CY 4102C | Specifications - Precision Balance CY 6102C
Maximum Capacity | 1200 g | 2200 g | 3100 g | 4100 g | 6100 g
Minimum load (OIML) | 0.5 g | 0.5 g | 0.5 g | 0.5 g | 0.5 g
Readability | 0.01 g | 0.01 g | 0.01 g | 0.01 g | 0.01 g
Verification unit [e] | 0.1 g | 0.1 g | 0.1 g | 0.1 g | 0.1 g
Repeatability | 0.01 g | 0.01 g | 0.01 g | 0.01 g | 0.01 g
Linearity | 0.02 g | 0.02 g | 0.02 g | 0.02 g | 0.03 g
Tare Range (Subtractive) | -1200 g | -2200 g | -3100 g | -4100 g | -6100 g
Standard repeatability [5% Max] | 0.005 g | 0.005 g | 0.005 g | 0.005 g | 0.005 g
Minimum Weight USP | 10 g | 10 g | 10 g | 10 g | 10 g
Minimum weight ((U=1%, k=2) | 1 g | 1 g | 1 g | 1 g | 1 g
Weighing Pan | 180 x 195 mm | 180 x 195 mm | 180 x 195 mm | 180 x 195 mm | 180 x 195 mm
Setting Time | 2- 3 Sec | 2- 3 Sec | 2- 3 Sec | 2- 3 Sec | 2- 3 Sec
Display | LCD Backlit | LCD Backlit | LCD Backlit | LCD Backlit | LCD Backlit
Calibration | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC)
Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
Operating temperature | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC
Relative humidity | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80%
Sensitivity Drift | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC)
Protection class | IP 43 | IP 43 | IP 43 | IP 43 | IP 43
Leveling system | manual | manual | manual | manual | manual
Interfaces | RS-232/USB | RS-232/USB | RS-232/USB | RS-232/USB | RS-232/USB
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold
Power Supply | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A
Compliance | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP
Balance Dimension (W x D X H) | 342.5 x 212 x 89.5 mm | 342.5 x 212 x 89.5 mm | 342.5 x 212 x 89.5 mm | 342.5 x 212 x 89.5 mm | 342.5 x 212 x 89.5 mm
Packaging dimensions (W x D X H) | 490 × 400 × 520 mm | 490 × 400 × 520 mm | 490 × 400 × 520 mm | 490 × 400 × 520 mm | 490 × 400 × 520 mm
Net weight | 7.5 kg | 7.5 kg | 7.5 kg | 7.5 kg | 7.5 kg
Gross weight | 12 kg | 12 kg | 12 kg | 12 kg | 12 kg
Precision Balance CY 3102
Intuitive to operate with Solid built construction. 3100 g capacity; 0.01 g readability; large, bright LCD backlit display; external calibration; easy leveling; USB and RS232 connection; overload protection; small footprint.
| Weighing Capacity | 3100 g
| Readability | 0.01 g
| Minimum Weight (USP, 0.1%, typical) | 10 g
Brilliant Proven Technology | Solid & Reliable Design | Bright, High-Contrast Display | Seamless Connectivity | 10 Easy-to-Use Built-In Applications | ISO/GLP Documentation | Quick & Simple Leveling | Easy Maintenance & Cleaning
The Precision Balances use a MonoTech single-block weighing sensor ensuring uniform load transmission, thermal stability, rapid stabilization, accurate corner-load performance, high repeatability, minimal drift, and long-term reliability in continuous operation. | Built with a solid design and high-quality materials for long-term durability, ensuring reliable and consistent operation with accurate results. Supplied with a standard protective cover for added safety. | Excellent readable display with bright LED-backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | The USB and RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data. | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | The built-in real-time clock ensures accurate, time-stamped data, providing traceable documentation that reflects true test results and supports compliance with SOPs and GLP regulations. | An easy-to-view circular bubble level and large adjustable feet make balance leveling quick, simple, and precise for stable, accurate weighing. | Washable parts dismantle in seconds without tools. Dishwasher-safe components and a smooth housing design make routine cleaning quick, simple, and hygienic for everyday laboratory use.
| Specifications - Precision Balance CY 1202 | Specifications - Precision Balance CY 2202 | Specifications - Precision Balance CY 3102 | Specifications - Precision Balance CY 4102 | Specifications - Precision Balance CY 6102
Maximum Capacity | 1200 g | 2200 g | 3100 g | 4100 g | 6100 g
Minimum load (OIML) | 0.5 g | 0.5 g | 0.5 g | 0.5 g | 0.5 g
Readability | 0.01 g | 0.01 g | 0.01 g | 0.01 g | 0.01 g
Verification unit [e] | 0.1 g | 0.1 g | 0.1 g | 0.1 g | 0.1 g
Repeatability | 0.01 g | 0.01 g | 0.01 g | 0.01 g | 0.01 g
Linearity | 0.02 g | 0.02 g | 0.02 g | 0.02 g | 0.03 g
Tare Range (Subtractive) | -1200 g | -2200 g | -3100 g | -4100 g | -6100 g
Standard repeatability [5% Max] | 0.005 g | 0.005 g | 0.005 g | 0.005 g | 0.005 g
Minimum Weight USP | 10 g | 10 g | 10 g | 10 g | 10 g
Minimum weight ((U=1%, k=2) | 1 g | 1 g | 1 g | 1 g | 1 g
Weighing Pan | 180 x 195 mm | 180 x 195 mm | 180 x 195 mm | 180 x 195 mm | 180 x 195 mm
Setting Time | 2- 3 Sec | 2- 3 Sec | 2- 3 Sec | 2- 3 Sec | 2- 3 Sec
Display | LCD Backlit | LCD Backlit | LCD Backlit | LCD Backlit | LCD Backlit
Calibration | External Calibration | External Calibration | External Calibration | External Calibration | External Calibration
Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
Operating temperature | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC
Relative humidity | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80%
Sensitivity Drift | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC)
Protection class | IP 43 | IP 43 | IP 43 | IP 43 | IP 43
Leveling system | manual | manual | manual | manual | manual
Interfaces | RS-232/USB | RS-232/USB | RS-232/USB | RS-232/USB | RS-232/USB
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold
Power Supply | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A
Compliance | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP
Balance Dimension (W x D X H) | 342.5 x 212 x 89.5 mm | 342.5 x 212 x 89.5 mm | 342.5 x 212 x 89.5 mm | 342.5 x 212 x 89.5 mm | 342.5 x 212 x 89.5 mm
Packaging dimensions (W x D X H) | 490 × 400 × 520 mm | 490 × 400 × 520 mm | 490 × 400 × 520 mm | 490 × 400 × 520 mm | 490 × 400 × 520 mm
Net weight | 7.5 kg | 7.5 kg | 7.5 kg | 7.5 kg | 7.5 kg
Gross weight | 12 kg | 12 kg | 12 kg | 12 kg | 12 kg
Precision Balance ALT 3102C
New generation technology with Solid built Construction. 3100 g capacity; 0.01 g readability; Colour TFT touchscreen display; internal Calibration; optimal feature leveling alert, easy leveling;and user management simplify and optimize your daily weighing processes. USB and RS232 connection; overload protection; small footprint; Power saving mode.
| Weighing Capacity | 3100 g
| Readability | 0.01 g
| Minimum Weight (USP, 0.1%, typical) | 10 g
Advanced Technology for Trusted Results | Intuitive Touchscreen | Automatic Internal Adjustment with PSAC | Advanced User Management & Access Control | Infrared Proximity Sensor | ALIBI Memory secure storage of measurements | 10 Easy-to-Use Built-In Applications | Easy Leveling for Accurate Results | Optional On-Screen Leveling Assistance | Audit Trial Report | Seamless Connectivity
The Precision Balances use a MonoTech single-block weighing sensor ensuring uniform load transmission, thermal stability, rapid stabilization, accurate corner-load performance, high repeatability, minimal drift, and long-term reliability in continuous operation. | The large 4.3” color TFT touchscreen ensures clear visibility and smooth operation, even while wearing gloves. Its intuitive interface and wipe-clean surfaces allow quick, hygienic cleaning after every weighing task. | Automatic motorized, built-in weight internal adjustment ensures continuous accuracy using internal test weights. The balance performs PSAC at scheduled times or automatically whenever ambient temperature changes are detected. | User Management ensures secure operation by allowing administrators to create up to 7 users and assign specific access rights, ensuring controlled usage, data integrity, and protection against unauthorized actions. | Equipped with two programmable hands-free sensors, the balance allows convenient contactless operation. Each sensor can be assigned functions such as Zero, Tare, Print, Home Screen, or user-defined tasks, improving workflow efficiency and maintaining hygiene. | ALIBI memory securely stores up to 1000 reports. ensuring reliable data storage, full traceability, and long-term safety. It guarantees continuity and protection of vital measurement data over extended periods. USB-A (Optional for mass data storage in a pendrive ) | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | ALT Series balances ensure maximum stability with adjustable feet and a front bubble level, delivering perfectly leveled, accurate, and reliable weighing results. | Optional real-time digital level guidance uses an integrated sensor to continuously monitor the balance position and provides on-screen instructions, enabling quick and accurate leveling within seconds. | AuditTrail adds value by providing complete, gapless documentation with automatic recording of calibration and leveling processes, including secure date and time stamps for full traceability. | The RS232, USB-C, and Ethernet interfaces allow easy connection to a printer, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data.
| Specifications - Precision Balance ALT 3102C | Specifications - Precision Balance ALT 6102C
Maximum Capacity | 3100 g | 6100 g
Minimum load (OIML) | 0.5 g | 0.5 g
Readability | 0.01 g | 0.01 g
Verification unit [e] | 0.1 g | 0.1 g
Repeatability | 0.01 g | 0.01 g
Linearity | 0.02 g | 0.03 g
Tare Range (Subtractive) | -3100 g | -6100 g
Standard repeatability [5% Max] | 0.005 g | 0.005 g
Minimum Weight USP | 10 g | 10 g
Minimum weight ((U=1%, k=2) | 1 g | 1 g
Weighing Pan | 190 x 205 mm | 190 x 205 mm
Setting Time | 2- 3 Sec | 2- 3 Sec
Display | 4.3" color TFT touchscreen | 4.3" color TFT touchscreen
Calibration | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC)
Features | 7 User management Password protection AuditTrai ALIBI Memory ( 1000 reports are stored. ) Minimum Sample Weight (MSW) Touch-free operation Infrared Proximity Sensor 4 English, German, French, Sp built-in languagesanish Levelling guidance ( Optional ) USB-A (Optional for mass data storage in a pendrive ) | 7 User management Password protection AuditTrai ALIBI Memory ( 1000 reports are stored. ) Minimum Sample Weight (MSW) Touch-free operation Infrared Proximity Sensor 4 built-in languages English, German, French, Spanish Levelling guidance ( Optional ) USB-A (Optional for mass data storage in a pendrive )
Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
Operating temperature | 15º to 35ºC | 15º to 35ºC
Relative humidity | 40% ÷ 80% | 40% ÷ 80%
Sensitivity Drift | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC)
Protection class | IP 43 | IP 43
Leveling system | manual | manual
Interfaces | RS232/USB | RS232/USB
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold
Power Supply | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz. | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz.
Compliance | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP
Balance Dimension (W x D X H)
Packaging dimensions (W x D X H) | 400 x 510 x 270 mm | 400 x 510 x 270 mm
Net weight
Gross weight | 5.9 kg | 5.9 kg
Precision Balance ALT 3102
New generation technology with Solid built Construction. 3100 g capacity; 0.01 g readability; Colour TFT touchscreen display; external Calibration; optimal feature leveling alert, easy leveling;and user management simplify and optimize your daily weighing processes. USB and RS232 connection; overload protection; small footprint; Power saving mode.
| Weighing Capacity | 3100 g
| Readability | 0.01 g
| Minimum Weight (USP, 0.1%, typical) | 10 g
Advanced Technology for Trusted Results | Intuitive Touchscreen | Advanced User Management & Access Control | Infrared Proximity Sensor | ALIBI Memory secure storage of measurements | 10 Easy-to-Use Built-In Applications | Easy Leveling for Accurate Results | Optional On-Screen Leveling Assistance | Audit Trial Report | Seamless Connectivity
The Precision Balances use a MonoTech single-block weighing sensor ensuring uniform load transmission, thermal stability, rapid stabilization, accurate corner-load performance, high repeatability, minimal drift, and long-term reliability in continuous operation. | The large 4.3” color TFT touchscreen ensures clear visibility and smooth operation, even while wearing gloves. Its intuitive interface and wipe-clean surfaces allow quick, hygienic cleaning after every weighing task. | User Management ensures secure operation by allowing administrators to create up to 7 users and assign specific access rights, ensuring controlled usage, data integrity, and protection against unauthorized actions. | Equipped with two programmable hands-free sensors, the balance allows convenient contactless operation. Each sensor can be assigned functions such as Zero, Tare, Print, Home Screen, or user-defined tasks, improving workflow efficiency and maintaining hygiene. | ALIBI memory securely stores up to 1000 reports. ensuring reliable data storage, full traceability, and long-term safety. It guarantees continuity and protection of vital measurement data over extended periods. USB-A (Optional for mass data storage in a pendrive ) | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | ALT Series balances ensure maximum stability with adjustable feet and a front bubble level, delivering perfectly leveled, accurate, and reliable weighing results. | Optional real-time digital level guidance uses an integrated sensor to continuously monitor the balance position and provides on-screen instructions, enabling quick and accurate leveling within seconds. | AuditTrail adds value by providing complete, gapless documentation with automatic recording of calibration and leveling processes, including secure date and time stamps for full traceability. | The RS232, USB-C, and Ethernet interfaces allow easy connection to a printer, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data.
| Specifications - Precision Balance ALT 3102 | Specifications - Precision Balance ALT 6102
Maximum Capacity | 3100 g | 6100 g
Minimum load (OIML) | 0.5 g | 0.5 g
Readability | 0.01 g | 0.01 g
Verification unit [e] | 0.1 g | 0.1 g
Repeatability | 0.01 g | 0.01 g
Linearity | 0.02 g | 0.03 g
Tare Range (Subtractive) | -3100 g | -6100 g
Standard repeatability [5% Max] | 0.005 g | 0.005 g
Minimum Weight USP | 10 g | 10 g
Minimum weight ((U=1%, k=2) | 1 g | 1 g
Weighing Pan | 190 x 205 mm | 190 x 205 mm
Setting Time | 2- 3 Sec | 2- 3 Sec
Display | 4.3" color TFT touchscreen | 4.3" color TFT touchscreen
Calibration | External Calibration | External Calibration
Features | 7 User management Password protection AuditTrai ALIBI Memory ( 1000 reports are stored. ) Minimum Sample Weight (MSW) Touch-free operation Infrared Proximity Sensor 4 English, German, French, Spanish built-in languages A Levelling guidance ( Optional ) USB-A (Optional for mass data storage in a pendrive ) | 7 User management Password protection AuditTrai ALIBI Memory ( 1000 reports are stored. ) Minimum Sample Weight (MSW) Touch-free operation Infrared Proximity Sensor 4 English, German, French, Spanish built-in languages A Levelling guidance ( Optional ) USB-A (Optional for mass data storage in a pendrive )
Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
Operating temperature | 15º to 35ºC | 15º to 35ºC
Relative humidity | 40% ÷ 80% | 40% ÷ 80%
Sensitivity Drift | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC)
Protection class | IP 43 | IP 43
Leveling system | manual | manual
Interfaces | RS232/USB | RS232/USB
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold
Power Supply | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz. | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz.
Compliance | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP
Balance Dimension (W x D X H)
Packaging dimensions (W x D X H) | 400 x 510 x 270 mm | 400 x 510 x 270 mm
Net weight
Gross weight | 5.9 kg | 5.9 kg
Precision Balance CY 3101C
Intuitive to operate with solid built construction. 3100 g capacity; 0.1 g readability; large, bright LCD backlit display; internal Calibration; easy leveling; USB and RS232 connection.
| Weighing Capacity | 3100 g
| Readability | 0.1 g
| Minimum Weight (USP, 0.1%, typical) | 82 g
Brilliant Proven Technology | Solid & Reliable Design | Bright, High-Contrast Display | Automatic Internal Calibration (PSAC) | Seamless Connectivity | 10 Easy-to-Use Built-In Applications | ISO/GLP Documentation | Quick & Simple Leveling | Easy Maintenance & Cleaning
The Precision Balances use a MonoTech single-block weighing sensor ensuring uniform load transmission, thermal stability, rapid stabilization, accurate corner-load performance, high repeatability, minimal drift, and long-term reliability in continuous operation. | Built with a solid design and high-quality materials for long-term durability, ensuring reliable and consistent operation with accurate results. Supplied with a standard protective cover for added safety. | Excellent readable display with bright LED-backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | Motorised, built-in weight internal calibration ensures precise weighing under all conditions. The Perfect Self Automatic Calibration (PSAC) function automatically maintains accuracy by adjusting for time and temperature changes. | The USB and RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data. | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | The built-in real-time clock ensures accurate, time-stamped data, providing traceable documentation that reflects true test results and supports compliance with SOPs and GLP regulations. | An easy-to-view circular bubble level and large adjustable feet make balance leveling quick, simple, and precise for stable, accurate weighing. | Washable parts dismantle in seconds without tools. Dishwasher-safe components and a smooth housing design make routine cleaning quick, simple, and hygienic for everyday laboratory use.
| Specifications - Precision Balance CY 3101C | Specifications - Precision Balance CY 6101C | Specifications - Precision Balance CY 10001C
Maximum Capacity | 3100 g | 6100 g | 10000 g
Minimum load (OIML) | 5 g | 5 g | 5 g
Readability | 0.1 g | 0.1 g | 0.1 g
Verification unit [e] | 5 g | 5 g | 5 g
Repeatability | 1 g | 1 g | 1 g
Linearity | 0.2 g | 0.2 g | 0.2 g
Tare Range (Subtractive) | -3100 g | -6100 g | -10000 g
Standard repeatability [5% Max] | 0.041 g | 0.041 g | 0.041 g
Minimum Weight USP | 82 g | 82 g | 82 g
Minimum weight ((U=1%, k=2) | 8.2 g | 8.2 g | 8.2 g
Weighing Pan | 180 x 195 mm | 180 x 195 mm | 180 x 195 mm
Setting Time | 2- 3 Sec | 2- 3 Sec | 2- 3 Sec
Display | LCD Backlit | LCD Backlit | LCD Backlit
Calibration | Internal Calibration | Internal Calibration | Internal Calibration
Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
Operating temperature | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC
Relative humidity | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80%
Sensitivity Drift | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC)
Protection class | IP 43 | IP 43 | IP 43
Leveling system | manual | manual | manual
Interfaces | RS-232/USB | RS-232/USB | RS-232/USB
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold
Power Supply | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A
Compliance | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP
Balance Dimension (W x D X H) | 342.5 x 212 x 89.5 mm | 342.5 x 212 x 89.5 mm | 342.5 x 212 x 89.5 mm
Packaging dimensions (W x D X H) | 490 × 400 × 520 mm | 490 × 400 × 520 mm | 490 × 400 × 520 mm
Net weight | 7.5 kg | 7.5 kg | 7.5 kg
Gross weight | 12 kg | 12 kg | 12 kg
Precision Balance CY 3101
Intuitive to operate with Solid built construction. 3100 g capacity; 0.1 g readability; large, bright LCD backlit display; external calibration; easy leveling; USB and RS232 connection; overload protection; small footprint.
| Weighing Capacity | 3100 g
| Readability | 0.1 g
| Minimum Weight (USP, 0.1%, typical) | 82 g
Brilliant Proven Technology | Solid & Reliable Design | Bright, High-Contrast Display | Seamless Connectivity | 10 Easy-to-Use Built-In Applications | ISO/GLP Documentation | Quick & Simple Leveling | Easy Maintenance & Cleaning
The Precision Balances use a MonoTech single-block weighing sensor ensuring uniform load transmission, thermal stability, rapid stabilization, accurate corner-load performance, high repeatability, minimal drift, and long-term reliability in continuous operation. | Built with a solid design and high-quality materials for long-term durability, ensuring reliable and consistent operation with accurate results. Supplied with a standard protective cover for added safety. | Excellent readable display with bright LED-backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | The USB and RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data. | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | The built-in real-time clock ensures accurate, time-stamped data, providing traceable documentation that reflects true test results and supports compliance with SOPs and GLP regulations. | An easy-to-view circular bubble level and large adjustable feet make balance leveling quick, simple, and precise for stable, accurate weighing. | Washable parts dismantle in seconds without tools. Dishwasher-safe components and a smooth housing design make routine cleaning quick, simple, and hygienic for everyday laboratory use.
| Specifications - Precision Balance CY 3101 | Specifications - Precision Balance CY 6101 | Specifications - Precision Balance CY 10001
Maximum Capacity | 3100 g | 6100 g | 10000 g
Minimum load (OIML) | 5 g | 5 g | 5 g
Readability | 0.1 g | 0.1 g | 0.1 g
Verification unit [e] | 1 g | 1 g | 1 g
Repeatability | 0.1 g | 0.1 g | 0.1 g
Linearity | 0.2 g | 0.2 g | 0.2 g
Tare Range (Subtractive) | -3100 g | -6100 g | -10000 g
Standard repeatability [5% Max] | 0.041 g | 0.041 g | 0.041 g
Minimum Weight USP | 82 g | 82 g | 82 g
Minimum weight ((U=1%, k=2) | 8.2 g | 8.2 g | 8.2 g
Weighing Pan | 180 x 195 mm | 180 x 195 mm | 180 x 195 mm
Setting Time | 2- 3 Sec | 2- 3 Sec | 2- 3 Sec
Display | LCD Backlit | LCD Backlit | LCD Backlit
Calibration | External Calibration | External Calibration | External Calibration
Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
Operating temperature | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC
Relative humidity | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80%
Sensitivity Drift | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC)
Protection class | IP 43 | IP 43 | IP 43
Leveling system | manual | manual | manual
Interfaces | RS-232/USB | RS-232/USB | RS-232/USB
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold
Power Supply | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A
Compliance | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP
Balance Dimension (W x D X H) | 342.5 x 212 x 89.5 mm | 342.5 x 212 x 89.5 mm | 342.5 x 212 x 89.5 mm
Packaging dimensions (W x D X H) | 490 × 400 × 520 mm | 490 × 400 × 520 mm | 490 × 400 × 520 mm
Net weight | 7.5 kg | 7.5 kg | 7.5 kg
Gross weight | 12 kg | 12 kg | 12 kg
Precision Balance ALT 3101C
New generation technology with Solid built Construction. 3100 g capacity; 0.1 g readability; Colour TFT touchscreen display; internal Calibration; optimal feature leveling alert, easy leveling;and user management simplify and optimize your daily weighing processes. USB and RS232 connection; overload protection; small footprint; Power saving mode.
| Weighing Capacity | 3100 g
| Readability | 0.1 g
| Minimum Weight (USP, 0.1%, typical) | 82 g
Advanced Technology for Trusted Results | Intuitive Touchscreen | Automatic Internal Adjustment with PSAC | Advanced User Management & Access Control | Infrared Proximity Sensor | ALIBI Memory secure storage of measurements | 10 Easy-to-Use Built-In Applications | Easy Leveling for Accurate Results | Optional On-Screen Leveling Assistance | Audit Trial Report | Seamless Connectivity | Seamless Connectivity
The Precision Balances use a MonoTech single-block weighing sensor ensuring uniform load transmission, thermal stability, rapid stabilization, accurate corner-load performance, high repeatability, minimal drift, and long-term reliability in continuous operation. | The large 4.3” color TFT touchscreen ensures clear visibility and smooth operation, even while wearing gloves. Its intuitive interface and wipe-clean surfaces allow quick, hygienic cleaning after every weighing task. | Automatic motorized, built-in weight internal adjustment ensures continuous accuracy using internal test weights. The balance performs PSAC at scheduled times or automatically whenever ambient temperature changes are detected. | User Management ensures secure operation by allowing administrators to create up to 7 users and assign specific access rights, ensuring controlled usage, data integrity, and protection against unauthorized actions. | Equipped with two programmable hands-free sensors, the balance allows convenient contactless operation. Each sensor can be assigned functions such as Zero, Tare, Print, Home Screen, or user-defined tasks, improving workflow efficiency and maintaining hygiene. | ALIBI memory securely stores up to 1000 reports. ensuring reliable data storage, full traceability, and long-term safety. It guarantees continuity and protection of vital measurement data over extended periods. USB-A (Optional for mass data storage in a pendrive ) | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | ALT Series balances ensure maximum stability with adjustable feet and a front bubble level, delivering perfectly leveled, accurate, and reliable weighing results. | Optional real-time digital level guidance uses an integrated sensor to continuously monitor the balance position and provides on-screen instructions, enabling quick and accurate leveling within seconds. | AuditTrail adds value by providing complete, gapless documentation with automatic recording of calibration and leveling processes, including secure date and time stamps for full traceability. | The RS232, USB-C, and Ethernet interfaces allow easy connection to a printer, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data. | The RS232, USB-C, and Ethernet interfaces allow easy connection to a printer, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data.
| Specifications - Precision Balance ALT 3101C | Specifications - Precision Balance ALT 6101C | Specifications - Precision Balance ALT 8101C
Maximum Capacity | 3100 g | 6100 g | 8100 g
Minimum load (OIML) | 5 g | 5 g | 5 g
Readability | 0.1 g | 0.1 g | 0.1 g
Verification unit [e] | 1 g | 1 g | 1 g
Repeatability | 0.1 g | 0.1 g | 0.1 g
Linearity | 0.2 g | 0.3 g | 0.3 g
Tare Range (Subtractive) | -3100 g | -6100 g | -8100 g
Standard repeatability [5% Max] | 0.041 g | 0.041 g | 0.041 g
Minimum Weight USP | 82 g | 82 g | 82 g
Minimum weight ((U=1%, k=2) | 8.2 g | 8.2 g | 8.2 g
Weighing Pan | 190 x 205 mm | 190 x 205 mm | 190 x 205 mm
Setting Time | 2- 3 Sec | 2- 3 Sec | 2- 3 Sec
Display | 4.3" color TFT touchscreen | 4.3" color TFT touchscreen | 4.3" color TFT touchscreen
Calibration | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC)
Features | 7 User management Password protection AuditTrai ALIBI Memory ( 1000 reports are stored. ) Minimum Sample Weight (MSW) Touch-free operation Infrared Proximity Sensor 4 English, German, French, Spanish built-in languages A Levelling guidance ( Optional ) USB-A (Optional for mass data storage in a pendrive ) | 7 User management Password protection AuditTrai ALIBI Memory ( 1000 reports are stored. ) Minimum Sample Weight (MSW) Touch-free operation Infrared Proximity Sensor 4 English, German, French, Spanish built-in languages A Levelling guidance ( Optional ) USB-A (Optional for mass data storage in a pendrive ) | 7 User management Password protection AuditTrai ALIBI Memory ( 1000 reports are stored. ) Minimum Sample Weight (MSW) Touch-free operation Infrared Proximity Sensor 4 English, German, French, Spanish built-in languages A Levelling guidance ( Optional ) USB-A (Optional for mass data storage in a pendrive )
Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
Operating temperature | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC
Relative humidity | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80%
Sensitivity Drift | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC)
Protection class | IP 43 | IP 43 | IP 43
Leveling system | manual | manual | manual
Interfaces | RS232/USB | RS232/USB | RS232/USB
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold
Power Supply | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz. | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz. | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz.
Compliance | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP
Balance Dimension (W x D X H)
Packaging dimensions (W x D X H) | 400 x 510 x 270 mm | 400 x 510 x 270 mm | 400 x 510 x 270 mm
Net weight
Gross weight | 5.9 kg | 5.9 kg | 5.9 kg
Precision Balance ALT 3101
New generation technology with Solid built Construction. 3100 g capacity; 0.1 g readability; Colour TFT touchscreen display; external Calibration; optimal feature leveling alert, easy leveling;and user management simplify and optimize your daily weighing processes. USB and RS232 connection; overload protection; small footprint; Power saving mode.
| Weighing Capacity | 3100 g
| Readability | 0.1 g
| Minimum Weight (USP, 0.1%, typical) | 82 g
Advanced Technology for Trusted Results | Intuitive Touchscreen | Advanced User Management & Access Control | Infrared Proximity Sensor | ALIBI Memory secure storage of measurements | 10 Easy-to-Use Built-In Applications | Easy Leveling for Accurate Results | Optional On-Screen Leveling Assistance | Audit Trial Report | Seamless Connectivity
The Precision Balances use a MonoTech single-block weighing sensor ensuring uniform load transmission, thermal stability, rapid stabilization, accurate corner-load performance, high repeatability, minimal drift, and long-term reliability in continuous operation. | The large 4.3” color TFT touchscreen ensures clear visibility and smooth operation, even while wearing gloves. Its intuitive interface and wipe-clean surfaces allow quick, hygienic cleaning after every weighing task. | User Management ensures secure operation by allowing administrators to create up to 7 users and assign specific access rights, ensuring controlled usage, data integrity, and protection against unauthorized actions. | Equipped with two programmable hands-free sensors, the balance allows convenient contactless operation. Each sensor can be assigned functions such as Zero, Tare, Print, Home Screen, or user-defined tasks, improving workflow efficiency and maintaining hygiene. | ALIBI memory securely stores up to 1000 reports. ensuring reliable data storage, full traceability, and long-term safety. It guarantees continuity and protection of vital measurement data over extended periods. USB-A (Optional for mass data storage in a pendrive ) | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | ALT Series balances ensure maximum stability with adjustable feet and a front bubble level, delivering perfectly leveled, accurate, and reliable weighing results. | Optional real-time digital level guidance uses an integrated sensor to continuously monitor the balance position and provides on-screen instructions, enabling quick and accurate leveling within seconds. | AuditTrail adds value by providing complete, gapless documentation with automatic recording of calibration and leveling processes, including secure date and time stamps for full traceability. | The RS232, USB-C, and Ethernet interfaces allow easy connection to a printer, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data.
| Specifications - Precision Balance ALT 3101 | Specifications - Precision Balance ALT 6101 | Specifications - Precision Balance ALT 8101
Maximum Capacity | 3100 g | 6100 g | 8100 g
Minimum load (OIML) | 5 g | 5 g | 5 g
Readability | 0.1 g | 0.1 g | 0.1 g
Verification unit [e] | 1 g | 1 g | 1 g
Repeatability | 0.1 g | 0.1 g | 0.1 g
Linearity | 0.2 g | 0.3 g | 0.3 g
Tare Range (Subtractive) | -3100 g | -6100 g | -8100 g
Standard repeatability [5% Max] | 0.041 g | 0.041 g | 0.041 g
Minimum Weight USP | 82 g | 82 g | 82 g
Minimum weight ((U=1%, k=2) | 8.2 g | 8.2 g | 8.2 g
Weighing Pan | 190 x 205 mm | 190 x 205 mm | 190 x 205 mm
Setting Time | 2- 3 Sec | 2- 3 Sec | 2- 3 Sec
Display | 4.3" color TFT touchscreen | 4.3" color TFT touchscreen | 4.3" color TFT touchscreen
Calibration | External Calibration | External Calibration | External Calibration
Features | 7 User management Password protection AuditTrai ALIBI Memory ( 1000 reports are stored. ) Minimum Sample Weight (MSW) Touch-free operation Infrared Proximity Sensor 4 English, German, French, Spanish built-in languages A Levelling guidance ( Optional ) USB-A (Optional for mass data storage in a pendrive ) | 7 User management Password protection AuditTrai ALIBI Memory ( 1000 reports are stored. ) Minimum Sample Weight (MSW) Touch-free operation Infrared Proximity Sensor 4 English, German, French, Spanish built-in languages A Levelling guidance ( Optional ) USB-A (Optional for mass data storage in a pendrive ) | 7 User management Password protection AuditTrai ALIBI Memory ( 1000 reports are stored. ) Minimum Sample Weight (MSW) Touch-free operation Infrared Proximity Sensor 4 English, German, French, Spanish built-in languages A Levelling guidance ( Optional ) USB-A (Optional for mass data storage in a pendrive )
Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
Operating temperature | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC
Relative humidity | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80%
Sensitivity Drift | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC) | ± 2ppm (2 x 10-6 / ºC)
Protection class | IP 43 | IP 43 | IP 43
Leveling system | manual | manual | manual
Interfaces | RS232/USB | RS232/USB | RS232/USB
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold
Power Supply | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz. | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz. | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz.
Compliance | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP
Balance Dimension (W x D X H)
Packaging dimensions (W x D X H) | 400 x 510 x 270 mm | 400 x 510 x 270 mm | 400 x 510 x 270 mm
Net weight
Gross weight | 5.9 kg | 5.9 kg | 5.9 kg
Precision Balance CY 10002HC
Intuitive to operate with solid, durable construction. High-capacity precision balance with 10 kg capacity and 0.01 g readability. Large, bright LCD backlit display for easy viewing. Internal calibration and simple leveling for accurate results. USB and RS232 connectivity for seamless data transfer. Large pan size for high-capacity weighing applications.
| Weighing Capacity | 10 kg
| Readability | 0.01 g
| Minimum Weight (USP, 0.1%, typical) | 10 g
Brilliant Proven Technology | Solid & Reliable Design | Bright, High-Contrast Display | Automatic Internal Calibration (PSAC) | Seamless Connectivity | 10 Easy-to-Use Built-In Applications | ISO/GLP Documentation | Quick & Simple Leveling | Easy Maintenance & Cleaning
The Precision Balances use a MonoTech single-block weighing sensor ensuring uniform load transmission, thermal stability, rapid stabilization, accurate corner-load performance, high repeatability, minimal drift, and long-term reliability in continuous operation. | Built with a solid design and high-quality materials for long-term durability, ensuring reliable and consistent operation with accurate results. Supplied with a standard protective cover for added safety. | Excellent readable display with bright LED-backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | Motorised, built-in weight internal calibration ensures precise weighing under all conditions. The Perfect Self Automatic Calibration (PSAC) function automatically maintains accuracy by adjusting for time and temperature changes. | The USB and RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data. | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | The built-in real-time clock ensures accurate, time-stamped data, providing traceable documentation that reflects true test results and supports compliance with SOPs and GLP regulations. | An easy-to-view circular bubble level and large adjustable feet make balance leveling quick, simple, and precise for stable, accurate weighing. | Washable parts dismantle in seconds without tools. Dishwasher-safe components and a smooth housing design make routine cleaning quick, simple, and hygienic for everyday laboratory use.
| Specifications - Precision Balance CY 10002HC | Specifications - Precision Balance CY 12002HC
Maximum Capacity | 10 kg | 12 kg
Minimum load (OIML) | 0.5 g | 0.5 g
Readability | 0.01 g | 0.01 g
Verification unit [e] | 0.1 g | 0.1 g
Repeatability | 0.01 g | 0.01 g
Linearity | 0.03 g | 0.03 g
Tare Range (Subtractive) | -10 kg | -12 kg
Standard repeatability [5% Max] | 0.005 g | 0.005 g
Minimum Weight USP | 10 g | 10 g
Minimum weight ((U=1%, k=2) | 1 g | 1 g
Weighing Pan | 190 x 205 mm | 190 x 205 mm
Setting Time | 2- 3 Sec | 2- 3 Sec
Display | Backlit LCD Display | Backlit LCD Display
Calibration | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC)
Units of measure | gm, mg, ct, GN, mo, oz, dwt, t1T, t1H, t1S, mom, Bat, MS | gm, mg, ct, GN, mo, oz, dwt, t1T, t1H, t1S, mom, Bat, MS
Operating temperature | 15 to 35ºC | 15 to 35ºC
Relative humidity | 40% ÷ 80% | 40% ÷ 80%
Sensitivity Drift | ± 2ppm (2 x 10⁻ / ºC) | ± 2ppm (2 x 10⁻ / ºC)
Protection class | IP 43 | IP 43
Leveling system | manual | manual
Interfaces | RS-232/USB | RS-232/USB
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold
Power Supply | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz. | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz.
Compliance | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP
Balance Dimension (W x D X H) | 517 x 302 x 130 mm | 518 x 302 x 130 mm
Packaging dimensions (W x D X H) | 540 x 540 x 250 mm | 540 x 540 x 250 mm
Net weight | 15 kg | 15 kg
Gross weight | 16 kg | 16 kg
High Capacity Precision Balance AXR 10K2HC
Next-generation precision weighing with robust, durable construction for long-term reliability. High-capacity performance with 10 kg capacity and 0.01 g readability for precise micro-weighing. Vibrant color TFT touchscreen with internal calibration and easy leveling alerts for intuitive operation. 21 CFR Part 11 compliant user management ensures secure, traceable laboratory workflows. Multiple connectivity options — USB, RS232, and Ethernet — for seamless data transfer. Compact, safe, and efficient design with overload protection and power-saving mode, perfect for modern laboratories.
| Weighing Capacity | 10 kg
| Readability | 0.01 g
| Minimum Weight (USP, 0.1%, typical) | 10 g
Advanced Technology for Trusted Results | Intuitive Touchscreen | Automatic Internal Adjustment with PSAC | 10 Easy-to-use Built-in-Applications | 21 CFR Part 11 Compliance | Data Free Transfer | Compact & Safe Design
The Precision Balances use a MonoTech single-block weighing sensor ensuring uniform load transmission, thermal stability, rapid stabilization, accurate corner-load performance, high repeatability, minimal drift, and long-term reliability in continuous operation. | The large 4.3” color TFT touchscreen ensures clear visibility and smooth operation, even while wearing gloves. Its intuitive interface and wipe-clean surfaces allow quick, hygienic cleaning after every weighing task. | Automatic motorized, built-in weight internal adjustment ensures continuous accuracy using internal test weights. The balance performs PSAC at scheduled times or automatically whenever ambient temperature changes are detected. | The ten easy to use build in applications include density determination, check weighing, totaling, and formulation, simplify your routine weighing tasks. These applications make your analytical operations much easier and accurate. | Our machine fully meets the requirements of 21 CFR Part 11/ EU GMP Annex 11. It also provides features like Full data security, Electronic signatures, Validations, Electronic records and Audit trails. | Simply transfer task parameters and results to a PC via many interface ports like RS232, USB A, USB B, Ethernet and wifi. Transcription errors are completely eliminated and reliability of your data is preserved. | Space-efficient construction ensures easy integration in any workspace. Features overload protection and power-saving functionality for reliable and efficient operation.
| Specifications - High Capacity Precision Balance AXR 10K2HC | Specifications - High Capacity Precision Balance AXR 12K2HC
Maximum Capacity | 10 kg | 12 kg
Minimum load (OIML) | 0.5 g | 0.5 g
Readability | 0.01 g | 0.01 g
Verification unit [e] | 0.1 g | 0.1 g
Repeatability | 0.01 g | 0.01 g
Linearity | 0.03 g | 0.03 g
Tare Range (Subtractive) | -10 kg | -12 kg
Standard repeatability [5% Max] | 0.05 g | 0.05 g
Minimum Weight USP | 10 g | 10 g
Minimum weight ((U=1%, k=2) | 1 g | 1 g
Weighing Pan | 190 x 205 mm | 190 x 205 mm
Setting Time | 2- 3 Sec | 2- 3 Sec
Display | 4.3" color TFT touchscreen | 4.3" color TFT touchscreen
Calibration | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC)
User Management | Log History [21 CFR Part 11 Compliant ] 7 User management Password protection AuditTrai Minimum Sample Weight (MSW) | Log History [21 CFR Part 11 Compliant ] 7 User management Password protection AuditTrai Minimum Sample Weight (MSW)
Features | Touch-free operation Infrared Proximity Sensor [Automatic Door] 4 built-in languages English, German, French, Spanish Levelling guidance | Touch-free operation Infrared Proximity Sensor [Automatic Door] 4 built-in languages English, German, French, Spanish Levelling guidance
Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
Operating temperature | 15º to 35ºC | 15º to 35ºC
Relative humidity | 40% ÷ 80% | 40% ÷ 80%
Sensitivity Drift | ± 1ppm (1 x 10-6/ ºC) | ± 1ppm (1 x 10-6/ ºC)
Protection class | IP 43 | IP 43
Leveling system | manual | manual
Interfaces | RS232/USB | RS232/USB
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold
Power Supply | DC Adapter, input 100 ~ 240V 0.5A output 13V / 0.9A | DC Adapter, input 100 ~ 240V 0.5A output 13V / 0.9A
Compliance | ISO, GLP / GMP & USP | ISO, GLP / GMP & USP
Balance Dimension (W x D X H)
Packaging dimensions (W x D X H)
Net weight
Gross weight
High Capacity Precision Balance CY 15001HC
Intuitive to operate with solid, durable construction. High-capacity precision balance with 15 kg capacity and 0.1 g readability. Large, bright LCD backlit display for easy viewing. Internal calibration and simple leveling for accurate results. USB and RS232 connectivity for seamless data transfer. Large pan size for high-capacity weighing applications.
| Weighing Capacity | 15 kg
| Readability | 0.1 g
| Minimum Weight (USP, 0.1%, typical) | 82 g
Brilliant Proven Technology | Solid & Reliable Design | Bright, High-Contrast Display | Automatic Internal Calibration (PSAC) | Seamless Connectivity | 10 Easy-to-Use Built-In Applications | ISO/GLP Documentation | Quick & Simple Leveling | High Quality Design | Easy Maintenance & Cleaning
The Precision Balances use a MonoTech single-block weighing sensor ensuring uniform load transmission, thermal stability, rapid stabilization, accurate corner-load performance, high repeatability, minimal drift, and long-term reliability in continuous operation. | Built with a solid design and high-quality materials for long-term durability, ensuring reliable and consistent operation with accurate results. Supplied with a standard protective cover for added safety. | Excellent readable display with bright LED-backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | Motorised, built-in weight internal calibration ensures precise weighing under all conditions. The Perfect Self Automatic Calibration (PSAC) function automatically maintains accuracy by adjusting for time and temperature changes. | The USB and RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data. | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | The built-in real-time clock ensures accurate, time-stamped data, providing traceable documentation that reflects true test results and supports compliance with SOPs and GLP regulations. | An easy-to-view circular bubble level and large adjustable feet make balance leveling quick, simple, and precise for stable, accurate weighing. | An in-built four-point overload protection system safeguards the balance against excessive loading, ensuring safe and reliable operation. | Washable parts dismantle in seconds without tools. Dishwasher-safe components and a smooth housing design make routine cleaning quick, simple, and hygienic for everyday laboratory use.
| Specifications - Precision Balance CY 15001HC | Specifications - Precision Balance CY 25001HC | Specifications - Precision Balance CY 35001HC | Specifications - Precision Balance CY 50001HC | Specifications - Precision Balance CY 60000HC
Maximum Capacity | 15 kg | 25kg | 35 kg | 50 kg | 60 kg
Minimum load (OIML) | 5 g | 5 g | 5 g | 5 g | 10 g
Readability | 0.1 g | 0.1 g | 0.1 g | 0.1 g | 1 g
Verification unit [e] | 1 g | 1 g | 1 g | 1 g | 100 g
Repeatability | 0.1 g | 0.1 g | 0.1 g | 0.1 g | 1 g
Linearity | 0.2 g | 0.2 g | 0.2 g | 0.2 g | 2 g
Tare Range (Subtractive) | -15 kg | -25kg | -35 kg | -50 kg | -60 kg
Standard repeatability [5% Max] | 0.041 g | 0.041 g | 0.041 g | 0.041 g | 0.41 g
Minimum Weight USP | 82 g | 82 g | 82 g | 82 g | 820 g
Minimum weight ((U=1%, k=2) | 8.2 g | 8.2 g | 8.2 g | 8.2 g | 82 g
Weighing Pan | 370 x 360 mm | 370 x 360 mm | 370 x 360 mm | 370 x 360 mm | 370 x 360 mm
Setting Time | 2 - 3 Sec. | 2 - 3 Sec. | 2 - 3 Sec. | 2 - 3 Sec. | 2 - 3 Sec.
Display | Backlit LCD Display | Backlit LCD Display | Backlit LCD Display | Backlit LCD Display | Backlit LCD Display
Calibration | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC)
Units of measure | gm, mg, ct, GN, mo, oz, dwt, t1T, t1H, t1S, mom, Bat, MS | gm, mg, ct, GN, mo, oz, dwt, t1T, t1H, t1S, mom, Bat, MS | gm, mg, ct, GN, mo, oz, dwt, t1T, t1H, t1S, mom, Bat, MS | gm, mg, ct, GN, mo, oz, dwt, t1T, t1H, t1S, mom, Bat, MS | gm, mg, ct, GN, mo, oz, dwt, t1T, t1H, t1S, mom, Bat, MS
Operating temperature | 15 to 35ºC | 15 to 35ºC | 15 to 35ºC | 15 to 35ºC | 15 to 35ºC
Relative humidity | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80%
Sensitivity Drift | ± 2ppm (2 x 10⁻ / ºC) | ± 2ppm (2 x 10⁻ / ºC) | ± 2ppm (2 x 10⁻ / ºC) | ± 2ppm (2 x 10⁻ / ºC) | ± 2ppm (2 x 10⁻ / ºC)
Protection class | IP 43 | IP 43 | IP 43 | IP 43 | IP 43
Leveling system | manual | manual | manual | manual | manual
Interfaces | RS232/USB | RS232/USB | RS232/USB | RS232/USB | RS232/USB
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold
Power Supply | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz. | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz. | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz. | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz. | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz.
Compliance | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP
Balance Dimension (W x D X H) | 517 x 302 x 130 mm | 518 x 302 x 130 mm | 519 x 302 x 130 mm | 520 x 302 x 130 mm | 521 x 302 x 130 mm
Packaging dimensions (W x D X H) | 540 x 540 x 250 mm | 540 x 540 x 250 mm | 540 x 540 x 250 mm | 540 x 540 x 250 mm | 540 x 540 x 250 mm
Net weight | 15 kg | 15 kg | 15 kg | 15 kg | 15 kg
Gross weight | 16 kg | 16 kg | 16 kg | 16 kg | 16 kg
High Capacity Precision Balance CY 15001H
Intuitive to operate with solid, durable construction. High-capacity precision balance with 15 kg capacity and 0.1 g readability. Large, bright LCD backlit display for easy viewing. external calibration and simple leveling for accurate results. USB and RS232 connectivity for seamless data transfer. Large pan size for high-capacity weighing applications.
| Weighing Capacity | 15 kg
| Readability | 0.1 g
| Minimum Weight (USP, 0.1%, typical) | 82 g
Brilliant Proven Technology | Solid & Reliable Design | Bright, High-Contrast Display | Seamless Connectivity | 10 Easy-to-Use Built-In Applications | ISO/GLP Documentation | Quick & Simple Leveling | High Quality Design | Easy Maintenance & Cleaning
The Precision Balances use a MonoTech single-block weighing sensor ensuring uniform load transmission, thermal stability, rapid stabilization, accurate corner-load performance, high repeatability, minimal drift, and long-term reliability in continuous operation. | Built with a solid design and high-quality materials for long-term durability, ensuring reliable and consistent operation with accurate results. Supplied with a standard protective cover for added safety. | Excellent readable display with bright LED-backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | The USB and RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data. | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | The built-in real-time clock ensures accurate, time-stamped data, providing traceable documentation that reflects true test results and supports compliance with SOPs and GLP regulations. | An easy-to-view circular bubble level and large adjustable feet make balance leveling quick, simple, and precise for stable, accurate weighing. | An in-built four-point overload protection system safeguards the balance against excessive loading, ensuring safe and reliable operation. | Washable parts dismantle in seconds without tools. Dishwasher-safe components and a smooth housing design make routine cleaning quick, simple, and hygienic for everyday laboratory use.
| Specifications - Precision Balance CY 15001H | Specifications - Precision Balance CY 25001H | Specifications - Precision Balance CY 35001H | Specifications - Precision Balance CY 60000H
Maximum Capacity | 15 kg | 25kg | 35 kg | 60 kg
Minimum load (OIML) | 1 g | 1 g | 1 g | 10 g
Readability | 0.1 g | 0.1 g | 0.1 g | 1 g
Verification unit [e] | 5 g | 5 g | 5 g | 100 g
Repeatability | 0.1 g | 0.1 g | 0.1 g | 1 g
Linearity | 0.2 g | 0.2 g | 0.2 g | 2 g
Tare Range (Subtractive) | -15 kg | -25kg | -35 kg | -60 kg
Standard repeatability [5% Max] | 0.041 g | 0.041 g | 0.041 g | 0.41 g
Minimum Weight USP | 82 g | 82 g | 82 g | 820 g
Minimum weight ((U=1%, k=2) | 8.2 g | 8.2 g | 8.2 g | 82 g
Weighing Pan | 370 x 270 mm | 370 x 360 mm | 370 x 360 mm | 370 x 360 mm
Setting Time | 2 - 3 Sec. | 2 - 3 Sec. | 2 - 3 Sec. | 2 - 3 Sec.
Display | Backlit LCD Display | Backlit LCD Display | Backlit LCD Display | Backlit LCD Display
Calibration | External Calibration | External Calibration | External Calibration | External Calibration
Units of measure | gm, mg, ct, GN, mo, oz, dwt, t1T, t1H, t1S, mom, Bat, MS | gm, mg, ct, GN, mo, oz, dwt, t1T, t1H, t1S, mom, Bat, MS | gm, mg, ct, GN, mo, oz, dwt, t1T, t1H, t1S, mom, Bat, MS | gm, mg, ct, GN, mo, oz, dwt, t1T, t1H, t1S, mom, Bat, MS
Operating temperature | 15 to 35ºC | 15 to 35ºC | 15 to 35ºC | 15 to 35ºC
Relative humidity | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80%
Sensitivity Drift | ± 2ppm (2 x 10⁻ / ºC) | ± 2ppm (2 x 10⁻ / ºC) | ± 2ppm (2 x 10⁻ / ºC) | ± 2ppm (2 x 10⁻ / ºC)
Protection class | IP 43 | IP 43 | IP 43 | IP 43
Leveling system | manual | manual | manual | manual
Interfaces | RS232/USB | RS232/USB | RS232/USB | RS232/USB
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold
Power Supply | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz. | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz. | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz. | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz.
Compliance | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP
Balance Dimension (W x D X H) | 517 x 302 x 130 mm | 518 x 302 x 130 mm | 519 x 302 x 130 mm | 521 x 302 x 130 mm
Packaging dimensions (W x D X H) | 540 x 540 x 250 mm | 540 x 540 x 250 mm | 540 x 540 x 250 mm | 540 x 540 x 250 mm
Net weight | 15 kg | 15 kg | 15 kg | 15 kg
Gross weight | 16 kg | 16 kg | 16 kg | 16 kg
High Capacity Precision Balance AXR 15K1HC
Next-generation precision weighing with robust, durable construction for long-term reliability. High-capacity performance with 15 kg capacity and 0.1 g readability for precise micro-weighing. Vibrant color TFT touchscreen with internal calibration and easy leveling alerts for intuitive operation. 21 CFR Part 11 compliant user management ensures secure, traceable laboratory workflows. Multiple connectivity options — USB, RS232, for seamless data transfer. large size stainless steel pan , safe, and efficient design with overload protection and power-saving mode, perfect for modern laboratories.
| Weighing Capacity | 15 kg
| Readability | 0.1 g
| Minimum Weight (USP, 0.1%, typical) | 82 g
Advanced Technology for Trusted Results | Intuitive Touchscreen | Automatic Internal Adjustment with PSAC | 10 Easy-to-use Built-in-Applications | 21 CFR Part 11 Compliance | Data Free Transfer | Compact & Safe Design
The Precision Balances use a MonoTech single-block weighing sensor ensuring uniform load transmission, thermal stability, rapid stabilization, accurate corner-load performance, high repeatability, minimal drift, and long-term reliability in continuous operation. | The large 4.3” color TFT touchscreen ensures clear visibility and smooth operation, even while wearing gloves. Its intuitive interface and wipe-clean surfaces allow quick, hygienic cleaning after every weighing task. | Automatic motorized, built-in weight internal adjustment ensures continuous accuracy using internal test weights. The balance performs PSAC at scheduled times or automatically whenever ambient temperature changes are detected. | The ten easy to use build in applications include density determination, check weighing, totaling, and formulation, simplify your routine weighing tasks. These applications make your analytical operations much easier and accurate. | Our machine fully meets the requirements of 21 CFR Part 11/ EU GMP Annex 11. It also provides features like Full data security, Electronic signatures, Validations, Electronic records and Audit trails. | Simply transfer task parameters and results to a PC via many interface ports like RS232, USB A, USB B, Ethernet and wifi. Transcription errors are completely eliminated and reliability of your data is preserved. | Space-efficient construction ensures easy integration in any workspace. Features overload protection and power-saving functionality for reliable and efficient operation.
| Specifications -High Capacity Precision Balance AXR 15K1HC | Specifications -High Capacity Precision Balance AXR 35K1HC | Specifications -High Capacity Precision Balance AXR 60K0HC
Maximum Capacity | 15 kg | 35 kg | 60 kg
Minimum load (OIML) | 10 g | 10 g | 100 g
Readability | 0.1 g | 0.1 g | 1 g
Verification unit [e] | 5 g | 5 g | 100 g
Repeatability | 0.1 g | 0.1 g | 1 g
Linearity | 0.2 g | 0.2 g | 2 g
Tare Range (Subtractive) | -15 kg | -35 kg | -60 kg
Standard repeatability [5% Max] | 0.041 g | 0.041 g | 4.00 g
Minimum Weight USP | 82 g | 82 g | 820 g
Minimum weight ((U=1%, k=2) | 8.2 g | 8.2 g | 82 g
Weighing Pan | 370 x 360 mm | 370 x 360 mm | 370 x 360 mm
Setting Time | 2- 3 Sec | 2 - 3 Sec. | 2 - 3 Sec.
Display | 4.3" color TFT touchscreen | 4.3" color TFT touchscreen | 4.3" color TFT touchscreen
Calibration | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC)
User Management | Log History [21 CFR Part 11 Compliant ] 7 User management Password protection AuditTrai Minimum Sample Weight (MSW) | Log History [21 CFR Part 11 Compliant ] 7 User management Password protection AuditTrai Minimum Sample Weight (MSW) | Log History [21 CFR Part 11 Compliant ] 7 User management Password protection AuditTrai Minimum Sample Weight (MSW)
Features | Touch-free operation Infrared Proximity Sensor [Automatic Door] 4 built-in languages English, German, French, Spanish Levelling guidance | Touch-free operation Infrared Proximity Sensor [Automatic Door] 4 built-in languages English, German, French, Spanish Levelling guidance | Touch-free operation Infrared Proximity Sensor [Automatic Door] 4 built-in languages English, German, French, Spanish Levelling guidance
Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
Operating temperature | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC
Relative humidity | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80%
Sensitivity Drift | ± 1ppm (1 x 10-6/ ºC) | ± 1ppm (1 x 10-6/ ºC) | ± 1ppm (1 x 10-6/ ºC)
Protection class | IP 43 | IP 43 | IP 43
Leveling system | manual | manual | manual
Interfaces | RS232/USB | RS232/USB | RS232/USB
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold
Power Supply | DC Adapter, input 100 ~ 240V 0.5A output 13V / 0.9A | DC Adapter, input 100 ~ 240V 0.5A output 13V / 0.9A | DC Adapter, input 100 ~ 240V 0.5A output 13V / 0.9A
Compliance | ISO, GLP / GMP & USP | ISO, GLP / GMP & USP | ISO, GLP / GMP & USP
Balance Dimension (W x D X H)
Packaging dimensions (W x D X H) | 550 x 550 x 260 mm | 550 x 550 x 260 mm | 550 x 550 x 260 mm
Net weight
Gross weight | 16 kg | 16 kg | 16 kg
High Capacity Precision Balance AXR 15K1H
Next-generation precision weighing with robust, durable construction for long-term reliability. High-capacity performance with 15 kg capacity and 0.1 g readability for precise micro-weighing. Vibrant color TFT touchscreen with external calibration and easy leveling alerts for intuitive operation. 21 CFR Part 11 compliant user management ensures secure, traceable laboratory workflows. Multiple connectivity options — USB, RS232, and Ethernet — for seamless data transfer. , large size stanless steel pan safe, and efficient design with overload protection and power-saving mode, perfect for modern laboratories.
| Weighing Capacity | 15 kg
| Readability | 0.1 g
| Minimum Weight (USP, 0.1%, typical) | 82 g
Advanced Technology for Trusted Results | Intuitive Touchscreen | Designed for Easy Handling and Access | Advanced User Management & Access Control | Infrared Proximity Sensor | ALIBI Memory secure storage of measurements | 10 Easy-to-Use Built-In Applications | Easy Leveling for Accurate Results | Optional On-Screen Leveling Assistance | Audit Trial Report | Seamless Connectivity
The Precision Balances use a MonoTech single-block weighing sensor ensuring uniform load transmission, thermal stability, rapid stabilization, accurate corner-load performance, high repeatability, minimal drift, and long-term reliability in continuous operation. | The large 4.3” color TFT touchscreen ensures clear visibility and smooth operation, even while wearing gloves. Its intuitive interface and wipe-clean surfaces allow quick, hygienic cleaning after every weighing task. | The spacious weighing chamber provides extra height for large containers, convenient access to the weighing pan, smooth, removable top and side sliding doors, and a fully detachable draft shield for easy handling and cleaning. | User Management ensures secure operation by allowing administrators to create up to 7 users and assign specific access rights, ensuring controlled usage, data integrity, and protection against unauthorized actions. | Equipped with two programmable hands-free sensors, the balance allows convenient contactless operation. Each sensor can be assigned functions such as Zero, Tare, Print, Home Screen, or user-defined tasks, improving workflow efficiency and maintaining hygiene. | ALIBI memory securely stores up to 1000 reports. ensuring reliable data storage, full traceability, and long-term safety. It guarantees continuity and protection of vital measurement data over extended periods. USB-A (Optional for mass data storage in a pendrive ) | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | ALT Series balances ensure maximum stability with adjustable feet and a front bubble level, delivering perfectly leveled, accurate, and reliable weighing results. | Optional real-time digital level guidance uses an integrated sensor to continuously monitor the balance position and provides on-screen instructions, enabling quick and accurate leveling within seconds. | AuditTrail adds value by providing complete, gapless documentation with automatic recording of calibration and leveling processes, including secure date and time stamps for full traceability. | The RS232, USB-C, and Ethernet interfaces allow easy connection to a printer, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data.
| Specifications -High Capacity Precision Balance AXR 15K1H | Specifications -High Capacity Precision Balance AXR 35K1H | Specifications -High Capacity Precision Balance AXR 60K0H
Maximum Capacity | 15 kg | 35 kg | 60 kg
Minimum load (OIML) | 10 g | 10 g | 100 g
Readability | 0.1 g | 0.1 g | 1 g
Verification unit [e] | 5 g | 5 g | 100 g
Repeatability | 0.1 g | 0.1 g | 1 g
Linearity | 0.2 g | 0.2 g | 2 g
Tare Range (Subtractive) | -15 kg | -35 kg | -60 kg
Standard repeatability [5% Max] | 0.041 g | 0.041 g | 4.00 g
Minimum Weight USP | 82 g | 82 g | 820 g
Minimum weight ((U=1%, k=2) | 8.2 g | 8.2 g | 82 g
Weighing Pan | 370 x 360 mm | 370 x 360 mm | 370 x 360 mm
Setting Time | 2- 3 Sec | 2 - 3 Sec. | 2 - 3 Sec.
Display | 4.3" color TFT touchscreen | 4.3" color TFT touchscreen | 4.3" color TFT touchscreen
Calibration | External Calibration | External Calibration | External Calibration
User Management | Log History [21 CFR Part 11 Compliant ] 7 User management Password protection AuditTrai Minimum Sample Weight (MSW) | Log History [21 CFR Part 11 Compliant ] 7 User management Password protection AuditTrai Minimum Sample Weight (MSW) | Log History [21 CFR Part 11 Compliant ] 7 User management Password protection AuditTrai Minimum Sample Weight (MSW)
Features | Touch-free operation Infrared Proximity Sensor [Automatic Door] 4 built-in languages English, German, French, Spanish Levelling guidance | Touch-free operation Infrared Proximity Sensor [Automatic Door] 4 built-in languages English, German, French, Spanish Levelling guidance | Touch-free operation Infrared Proximity Sensor [Automatic Door] 4 built-in languages English, German, French, Spanish Levelling guidance
Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
Operating temperature | 15º to 35ºC | 15º to 35ºC | 15º to 35ºC
Relative humidity | 40% ÷ 80% | 40% ÷ 80% | 40% ÷ 80%
Sensitivity Drift | ± 1ppm (1 x 10-6/ ºC) | ± 1ppm (1 x 10-6/ ºC) | ± 1ppm (1 x 10-6/ ºC)
Protection class | IP 43 | IP 43 | IP 43
Leveling system | manual | manual | manual
Interfaces | RS232/USB | RS232/USB | RS232/USB
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold
Power Supply | DC Adapter, input 100 ~ 240V 0.5A output 13V / 0.9A | DC Adapter, input 100 ~ 240V 0.5A output 13V / 0.9A | DC Adapter, input 100 ~ 240V 0.5A output 13V / 0.9A
Compliance | ISO, GLP / GMP & USP | ISO, GLP / GMP & USP | ISO, GLP / GMP & USP
Balance Dimension (W x D X H)
Packaging dimensions (W x D X H) | 550 x 550 x 260 mm | 550 x 550 x 260 mm | 550 x 550 x 260 mm
Net weight
Gross weight | 16 kg | 16 kg | 16 kg
Precision Balance CG 302
Portable Lab Balance with 300 g capacity and 0.01 g readability ensures accurate, cost-effective weighing. The compact design with inbuilt rechargeable battery and bright red LED display enables easy use anywhere, even during power outages.
| Weighing Capacity | 300 g
| Readability | 0.01 g
| Calibration | External
Precision Load Cell Technology | Solid & Reliable Design | Bright LED Red Display | Power Efficiency
Equipped with a precision aluminum load cell (IP65) and advanced 8-bit microprocessor design for accurate and stable performance. The level indicator with adjustable feet ensures proper setup, while 150% overload protection enhances safety and reliability. | Designed with durable, high-quality housing to provide a solid and reliable structure, ensuring long-term stability and repeatable performance. | High-visibility display with bright LED red digits ensures clear readability. The high-contrast screen offers wide viewing angles, enabling quick and accurate readings in all lighting conditions. | Built-in rechargeable backup battery and automatic shutoff features extend battery life, ensuring reliable performance and convenient everyday operation anywhere anytime.
| Specifications - Precision Balance CG 302 | Specifications - Precision Balance CG 602
Maximum Capacity | 300 g | 600 g
Readability | 0.01 g | 0.01 g
Repeatability | 0.01 g | 0.01 g
Linearity | 0.02 g | 0.02 g
Tare Range (Subtractive) | -300 g | -600 g
Weighing Pan Diameter | Ø 118 mm | Ø 118 mm
Setting Time | 2-3 sec | 2-3 sec
Display | Bright LED Red | Bright LED Red
Calibration | External Calibration | External Calibration
Units of measure | g, ct, ozt, tl, tola, gn, dwt | g, ct, ozt, tl, tola, gn, dwt
Operating temperature | 10ºC to 40ºC | 10ºC to 40ºC
Interfaces | RS232 | RS232
Built-in Application | Counting | Counting
Battery | 6V 1.2Ah rechargeable battery | 6V 1.2Ah rechargeable battery
Power Supply | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz
Balance Dimension (W x D X H) | 190 x 230 x 155 mm | 190 x 230 x 155 mm
Packaging dimensions (W x D X H)
Net weight | 2 kg | 2 kg
Gross weight | 3 kg | 3kg
Precision Balance CG 602L
Portable Lab Balance with 600 g capacity and 0.01 g readability ensures accurate, cost-effective weighing. Its compact design with inbuilt rechargeable battery and bright LCD backlit display allows easy use anywhere, even during power outages.
| Weighing Capacity | 600 g
| Readability | 0.01 g
| Calibration | External
Precision Load Cell Technology | Solid & Reliable Design | Bright, High-Contrast Display | Power Efficiency
Equipped with a precision aluminum load cell (IP65) and advanced 8-bit microprocessor design for accurate and stable performance. The level indicator with adjustable feet ensures proper setup, while 150% overload protection enhances safety and reliability. | Designed with durable, high-quality housing to provide a solid and reliable structure, ensuring long-term stability and repeatable performance. | Excellent readable display with bright backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | Built-in rechargeable backup battery and automatic shutoff features extend battery life, ensuring reliable performance and convenient everyday operation anywhere anytime.
| Specifications - Precision Balance CG 602L | Specifications - Precision Balance CG 1202L | Specifications - Precision Balance CG 2002L
Maximum Capacity | 600 g | 1200 g | 2000 gm
Readability | 0.01 g | 0.01 g | 0.01 g
Repeatability | 0.01 g | 0.01 g | 0.01 g
Linearity | 0.02 g | 0.02 g | 0.02 g
Tare Range (Subtractive) | -600 g | -1200 g | -2000 gm
Weighing Pan | 120 x 120 mm | 120 x 120 mm | 120 x 120 mm
Setting Time | 2 - 3 sec | 2 - 3 sec | 2 - 3 sec
Display | Backlite LCD Display | Backlite LCD Display | Backlite LCD Display
Calibration | External Calibration | External Calibration | External Calibration
Units of measure | g, ct, ozt, tl, tola, gn, dwt | g, ct, ozt, tl, tola, gn, dwt | g, ct, ozt, tl, tola, gn, dwt
Operating temperature | 10ºC to 40ºC | 10ºC to 40ºC | 10ºC to 40ºC
Interfaces | RS232 | RS232 | RS232
Built-in Application | Counting, Percent Weighing, Check weighing | Counting, Percent Weighing, Check weighing | Counting, Percent Weighing, Check weighing
Battery | 6V 1.2Ah rechargeable battery | 6V 1.2Ah rechargeable battery | 6V 1.2Ah rechargeable battery
Power Supply | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz
Balance Dimension (W x D X H) | 200 x 285 x 170 mm | 200 x 285 x 170 mm | 200 x 285 x 170 mm
Packaging dimensions (W x D X H) | 330 x 230 x 160 mm | 330 x 230 x 160 mm | 330 x 230 x 160 mm
Net weight | 2 kg | 2 kg | 2 kg
Gross weight | 3 Kg | 3 Kg | 3 Kg
Precision Balance CG 3001L
Portable Lab Balance with 3000 g capacity and 0.1 g readability ensures accurate, cost-effective weighing. Its compact design with inbuilt rechargeable battery and bright LCD backlit display allows easy use anywhere, even during power outages.
| Weighing Capacity | 3000 g
| Readability | 0.1 g
| Calibration | External
Precision Load Cell Technology | Solid & Reliable Design | Bright, High-Contrast Display | Power Efficiency
Equipped with a precision aluminum load cell (IP65) and advanced 8-bit microprocessor design for accurate and stable performance. The level indicator with adjustable feet ensures proper setup, while 150% overload protection enhances safety and reliability. | Designed with durable, high-quality housing to provide a solid and reliable structure, ensuring long-term stability and repeatable performance. | Excellent readable display with bright backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | Built-in rechargeable backup battery and automatic shutoff features extend battery life, ensuring reliable performance and convenient everyday operation anywhere anytime.
| Specifications - Precision Balance CG 3001L | Specifications - Precision Balance CG 6001L | Specifications - Precision Balance CG 10001L | Specifications - Precision Balance CG 15001L | Specifications - Precision Balance CG 30001L
Maximum Capacity | 3000 g | 6000 g | 10000 g | 15000 g | -30000 g
Readability | 0.1 g | 0.1 g | 0.1 g | 0.1 g | 0.1 g
Repeatability | 0.1 g | 0.1 g | 0.1 g | 0.1 g | 0.1 g
Linearity | 0.2 g | 0.2 g | 0.2 g | 0.2 g | 0.2 g
Tare Range (Subtractive) | -3000 g | -6000 g | -10000 g | -15000 g | -30000 g
Weighing Pan | 175 x 155 mm | 175 x 155 mm | 175 x 155 mm | 175 x 155 mm | 175 x 155 mm
Setting Time | 2-3 sec | 2-3 sec | 2-3 sec | 2-3 sec | 2-3 sec
Display | LCD Backlit | LCD Backlit | LCD Backlit | LCD Backlit | LCD Backlit
Calibration | External Calibration | External Calibration | External Calibration | External Calibration | External Calibration
Units of measure | g, ct, ozt, tl, tola, gn, dwt | g, ct, ozt, tl, tola, gn, dwt | g, ct, ozt, tl, tola, gn, dwt | g, ct, ozt, tl, tola, gn, dwt | g, ct, ozt, tl, tola, gn, dwt
Operating temperature | 10ºC to 40ºC | 10ºC to 40ºC | 10ºC to 40ºC | 10ºC to 40ºC | 10ºC to 40ºC
Interfaces | RS232 | RS232 | RS232 | RS232 | RS232
Built-in Application | Counting, Percent Weighing, Check weighing | Counting, Percent Weighing, Check weighing | Counting, Percent Weighing, Check weighing | Counting, Percent Weighing, Check weighing | Counting, Percent Weighing, Check weighing
Battery | 6V 1.2Ah rechargeable battery | 6V 1.2Ah rechargeable battery | 6V 1.2Ah rechargeable battery | 6V 1.2Ah rechargeable battery | 6V 1.2Ah rechargeable battery
Power Supply | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz
Balance Dimension (W x D X H) | 190 x 290 x 90 mm | 190 x 290 x 90 mm | 190 x 290 x 90 mm | 190 x 290 x 90 mm | 190 x 290 x 90 mm
Packaging dimensions (W x D X H) | 380 x 320 x 150 mm | 380 x 320 x 150 mm | 380 x 320 x 150 mm | 380 x 320 x 150 mm | 380 x 320 x 150 mm
Net weight | 2 kg | 2 kg | 2 kg | 2 kg | 2 kg
Gross weight | 3 Kg | 3 Kg | 3 Kg | 3 Kg | 3 Kg
Precision Balance CG 203L
Portable Lab Balance with 200 g capacity and 0.001 g readability ensures accurate, cost-effective weighing. Its compact design with inbuilt rechargeable battery and bright LCD backlit display allows easy use anywhere, even during power outages.
| Weighing Capacity | 200 g
| Readability | 0.001 g
| Calibration | External
Precision Load Cell Technology | Solid & Reliable Design | Bright, High-Contrast Display | Power Efficiency
Equipped with a precision aluminum load cell (IP65) and advanced 8-bit microprocessor design for accurate and stable performance. The level indicator with adjustable feet ensures proper setup, while 150% overload protection enhances safety and reliability. | Designed with durable, high-quality housing to provide a solid and reliable structure, ensuring long-term stability and repeatable performance. | Excellent readable display with bright backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | Built-in rechargeable backup battery and automatic shutoff features extend battery life, ensuring reliable performance and convenient everyday operation anywhere anytime.
| Specifications - Precision Balance CG 203L | Specifications - Precision Balance CG 303L | Specifications - Precision Balance CG 503L
Maximum Capacity | 200 g | 300 g | 500 g
Readability | 0.001 g | 0.001 g | 0.001 g
Repeatability | 0.001 g | 0.001 g | 0.001 g
Linearity | 0.002 g | 0.002 g | 0.002 g
Tare Range (Subtractive) | -200 g | -300 g | -500 g
Weighing Pan Diameter | Ø 90 mm | Ø 90 mm | Ø 90 mm
Setting Time | 2-3 sec | 2-3 sec | 2-3 sec
Display | LCD Backlit | LCD Backlit | LCD Backlit
Calibration | External Calibration | External Calibration | External Calibration
Units of measure | g, ct, ozt, tl, tola, gn, dwt | g, ct, ozt, tl, tola, gn, dwt | g, ct, ozt, tl, tola, gn, dwt
Operating temperature | 10ºC to 40ºC | 10ºC to 40ºC | 10ºC to 40ºC
Interfaces | RS232 | RS232 | RS232
Built-in Application | Counting, Percent Weighing, Check weighing | Counting, Percent Weighing, Check weighing | Counting, Percent Weighing, Check weighing
Battery | 6V 1.2Ah rechargeable battery | 6V 1.2Ah rechargeable battery | 6V 1.2Ah rechargeable battery
Power Supply | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz
Balance Dimension (W x D X H) | 200 x 285 x 170 mm | 200 x 285 x 170 mm | 200 x 285 x 170 mm
Packaging dimensions (W x D X H) | 330 x 230 x 160 mm | 330 x 230 x 160 mm | 330 x 230 x 160 mm
Net weight | 3 kg | 3 kg | 3 kg
Gross weight | 3 Kg | 3 Kg | 3 Kg
Precision Balance CZ 602 NTEP
Accurate and reliable weighing with 600 g capacity and 0.01 g readability. NTEP-certified for guaranteed precision. Compact and portable, with an inbuilt rechargeable battery. The bright, backlit LCD makes readings easy anywhere, even during power outages.
| Weighing Capacity | 600 g
| Readability | 0.01 g
| Calibration | External
Capacity Bar Graph | Long-Life Battery Operation | Compact Space-Saving Design | Backlit LCD Display | User-Friendly Keypad Operation | Protective Draft Shield | NTEP Certified Accuracy | Backlit LCD Display | Stainless Steel Weighing Pan
Integrated capacity bar graph provides visual indication of available weighing range. Enables better load monitoring and helps prevent overload during operation. | Built-in rechargeable battery supports portable operation with extended runtime. Provides up to 40 hours of operation for uninterrupted weighing performance. | Compact footprint makes the balance ideal for limited workspace environments. Designed for convenience without compromising weighing performance. | High-visibility backlit LCD with large digits ensures clear and easy reading. Enhanced display readability improves operator comfort and efficiency. | Dedicated function keys provide quick access to primary weighing applications. Simplifies operation and improves workflow efficiency. | Draft shield minimizes environmental influence for improved weighing stability. Removable metal cover with pipette opening supports convenience and precision. | NTEP verified system ensures accurate, compliant, and reliable weighing performance. Provides confidence for trade, industrial, and regulated applications. | High-visibility backlit LCD with large digits ensures clear and easy reading. Enhanced display readability improves operator comfort and efficiency. | Durable stainless steel weighing pan offers hygienic and reliable sample handling. Designed for routine use with easy cleaning and long-term durability.
| Specifications - Precision Balance CZ 602 | Specifications - Precision Balance CZ 1202 | Specifications - Precision Balance CZ 1200 | Specifications - Precision Balance CZ 6000H
Maximum Capacity | 600 g | 1200 g | 1200 g | 6000 g
Readability | 0.01 g | 0.01 g | 0.1 g | 0.1 g
Repeatability | 0.01 g | 0.01 g | 0.1 g | 0.1 g
Linearity | 0.02 g | 0.02 g | 0.2 g | 0.2 g
Tare Range (Subtractive) | - 600 g | - 1200 g | - 1200 g | - 6000 g
Weighing Pan | 120 Ø mm | 120 Ø mm | 140 x 150mm | 140 x 150mm
Setting Time | 2-3 sec | 2-3 sec | 2-3 sec | 2-3 sec
Display | LCD Backlit | LCD Backlit | LCD Backlit | LCD Backlit
Calibration | External Calibration | External Calibration | External Calibration | External Calibration
Units of measure | g, ct, lb, oz, d, gn, ozt, dwt, mom, tl.t, tl.h, tlj, t, bt, n | g, ct, lb, oz, d, gn, ozt, dwt, mom, tl.t, tl.h, tlj, t, bt, n | g, ct, lb, oz, d, gn, ozt, dwt, mom, tl.t, tl.h, tlj, t, bt, n | g, ct, lb, oz, d, gn, ozt, dwt, mom, tl.t, tl.h, tlj, t, bt, n
Operating temperature | 10ºC to 40ºC | 10ºC to 40ºC | 10ºC to 40ºC | 10ºC to 40ºC
Built-in Application | Counting, Percent Weighing, accumulation | Counting, Percent Weighing, accumulation | Counting, Percent Weighing, accumulation | Counting, Percent Weighing, accumulation
Battery | 6V 1.2Ah rechargeable battery | 6V 1.2Ah rechargeable battery | 6V 1.2Ah rechargeable battery | 6V 1.2Ah rechargeable battery
Power Supply | Input: 100~240VAC 50-60Hz/ Output: DC 12V 500mA | Input: 100~240VAC 50-60Hz/ Output: DC 12V 500mA | Input: 100~240VAC 50-60Hz/ Output: DC 12V 500mA | Input: 100~240VAC 50-60Hz/ Output: DC 12V 500mA
Balance Dimension (W x D X H) | 158 x 143 x 64mm | 158 x 143 x 64mm | 158 x 143 x 64mm | 158 x 143 x 64mm
Packaging dimensions (W x D X H) | 34x26x12 cm | 34x26x12 cm | 34x26x12 cm | 34x26x12 cm
Net weight | 2 kg | 2 kg | 2 kg | 2 kg
Gross weight | 3 Kg | 3 Kg | 3 Kg | 3 Kg

## Pipette Check
Pipette Check Balance CY 224CP
Intuitive and robust balance with 220 g capacity and 0.1 mg readability. Features a bright backlit LCD, internal check, easy leveling, and USB/RS232 connectivity. Volumes can be verified in compliance with ISO 8655 standards CY–CP models are designed for gravimetric pipette verification, with integrated Pipette Check and an evaporation trap that reduces liquid loss and ensures accurate, reliable results.
| Weighing Capacity | 220 g
| Readability | 0.1 mg
Brilliant Proven Technology | Evaporation trap | Volume* of the pipette to be checked V | Automatic Internal Calibration (PSAC) | Seamless Connectivity | ISO/GLP Documentation | Easy Maintenance & Cleaning
The sensor truly stands out with its unparalleled design based on the electromagnetic force compensation (EMFC). It provides an exceptional performance that meets the highest of the set international standards. | Included with CY-CP models, the evaporation trap minimizes liquid loss during pipette testing by creating a saturated vapor zone, reducing errors and ensuring accurate, reliable results for small volumes. | 100 μl < V ≤ 1000 μl 1 ml < V ≤ 10 ml | Motorised, built-in weight internal calibration ensures precise weighing under all conditions. The Perfect Self Automatic Calibration (PSAC) function automatically maintains accuracy by adjusting for time and temperature changes. | The USB and RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data. | The built-in real-time clock ensures accurate, time-stamped data, providing traceable documentation that reflects true test results and supports compliance with SOPs and GLP regulations. | Washable parts dismantle in seconds without tools. Dishwasher-safe components and a smooth housing design make routine cleaning quick, simple, and hygienic for everyday laboratory use.
Specifications - Pipette Check Balance CY 224CP
Maximum Capacity | 220 g
Readability | 0.1 mg
Repeatability | 0.1 mg
Linearity | 0.2 mg
Tare Range (Subtractive) | -220 g
Setting Time | 2- 3 Sec
Display | LCD Backlit
Calibration | Motorised Internal Calibration (automatic/PSAC)
Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
Operating temperature | 15º to 35ºC
Interfaces | RS232/USB
Power Supply | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A
Compliance | ISO, GLP, GMP, USP
Balance Dimension (W x D X H) | 342.5 x 212 x 350 mm
Packaging dimensions (W x D X H) | 490 × 400 × 520 mm
Net weight | 8 kg
Gross weight | 12 kg

## flame proof Balances
Flameproof Balance CY 224FC
CY-FC Series flameproof analytical balances are built with a solid, chemical-resistant cast aluminum alloy LM6 body and epoxy polyester powder coating. Offering 220 g capacity and 0.1 mg readability, they feature a bright backlit LCD, internal calibration, easy leveling, and push-button operation. With IP66 ingress protection and optional RS232, they are CMFR-certified for Zone 1 and 2, Gas Group IIB, ensuring safe and reliable performance in hazardous environments.
| Weighing Capacity | 220 g
| Readability | 0.1 mg
| Minimum Weight (USP, 0.1%, typical) | 160 mg
Brilliant Proven Technology | Solid & Reliable Design | Bright, High-Contrast Display | 10 Easy-to-Use Built-In Applications | Quick & Simple Leveling
The sensor truly stands out with its unparalleled design based on the electromagnetic force compensation (EMFC). It provides an exceptional performance that meets the highest of the set international standards. | Built with a solid, flameproof enclosure of cast aluminum alloy LM6 and epoxy polyester powder coating, featuring push-button operation and IP66 ingress protection, suitable for Zone 1 and 2, Gas Group IIB, ensuring safe and reliable performance in hazardous environments. | Excellent readable display with bright LED-backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | An easy-to-view circular bubble level and large adjustable feet make balance leveling quick, simple, and precise for stable, accurate weighing.
Specifications - Flameproof Balance CY 224FC
Maximum Capacity | 220 g
Readability | 0.1 mg
Verification unit [e] | 1 mg
Repeatability | 0.1 mg
Linearity | 0.2 mg
Tare Range (Subtractive) | -220 g
Standard repeatability [5% Max] | 0.08 mg
Minimum Weight USP | 160 mg
Minimum weight ((U=1%, k=2) | 16 mg
Weighing Pan Diameter | Ø 90 mm
Setting Time | 2- 3 Sec
Display | LCD Backlit
Calibration | Motorised Internal Calibration (automatic/PSAC)
Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
Operating temperature | 15º to 35ºC
Relative humidity | 40% ÷ 80%
Sensitivity Drift | ± 2ppm (2 x 10-6 / ºC)
Protection class | IP 43
Interface | RS232 Optional
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold
Power Supply | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A
Compliance | ISO, GLP, GMP, USP
Balance Dimension (W x D X H) | 342.5 x 212 x 350 mm
Packaging dimensions (W x D X H) | 490 × 400 × 520 mm
Net weight | 8 kg
Gross weight | 12 kg
Flameproof Balance CY 224F
CY-FC Series flameproof analytical balances are built with a solid, chemical-resistant cast aluminum alloy LM6 body and epoxy polyester powder coating. Offering 220 g capacity and 0.1 mg readability, they feature a bright backlit LCD, external calibration, easy leveling, and push-button operation. With IP66 ingress protection and optional RS232, they are CMFR-certified for Zone 1 and 2, Gas Group IIB, ensuring safe and reliable performance in hazardous environments.
| Weighing Capacity | 220 g
| Readability | 0.1 mg
| Minimum Weight (USP, 0.1%, typical) | 160 mg
Brilliant Proven Technology | Solid & Reliable Design | Bright, High-Contrast Display | 10 Easy-to-Use Built-In Applications | Quick & Simple Leveling
The sensor truly stands out with its unparalleled design based on the electromagnetic force compensation (EMFC). It provides an exceptional performance that meets the highest of the set international standards. | Built with a solid, flameproof enclosure of cast aluminum alloy LM6 and epoxy polyester powder coating, featuring push-button operation and IP66 ingress protection, suitable for Zone 1 and 2, Gas Group IIB, ensuring safe and reliable performance in hazardous environments. | Excellent readable display with bright LED-backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | An easy-to-view circular bubble level and large adjustable feet make balance leveling quick, simple, and precise for stable, accurate weighing.
Specifications - Flameproof Balance CY 224F
Maximum Capacity | 220 g
Minimum load (OIML) | 10 mg
Readability | 0.1 mg
Verification unit [e] | 1 mg
Repeatability | 0.1 mg
Linearity | 0.2 mg
Tare Range (Subtractive) | -220 g
Standard repeatability [5% Max] | 0.08 mg
Minimum Weight USP | 160 mg
Minimum weight ((U=1%, k=2) | 16 mg
Weighing Pan Diameter | Ø 90 mm
Setting Time | 2- 3 Sec
Display | LCD Backlit
Calibration | External Calibration
Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
Operating temperature | 15º to 35ºC
Relative humidity | 40% ÷ 80%
Sensitivity Drift | ± 2ppm (2 x 10-6 / ºC)
Protection class | IP 43
Interface | RS232 Optional
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold
Power Supply | AC 220V / 50-60 Hz
Compliance | ISO, GLP, GMP, USP
Balance Dimension (W x D X H) | 342.5 x 212 x 350 mm
Packaging dimensions (W x D X H) | 490 × 400 × 520 mm
Net weight | 8 kg
Gross weight | 12 kg
Flameproof Balance CY 513FC
CY-FC Series flameproof precision balances are built with a solid, chemical-resistant cast aluminum alloy LM6 body and epoxy polyester powder coating. Offering 510 g capacity and 0.001 g readability, they feature a bright backlit LCD, internal calibration, easy leveling, and push-button operation. With IP66 ingress protection and optional RS232, they are CMFR-certified for Zone 1 and 2, Gas Group IIB, ensuring safe and reliable performance in hazardous environments.
| Weighing Capacity | 510 g
| Readability | 0.001g
| Minimum Weight (USP, 0.1%, typical) | 1 g
Brilliant Proven Technology | Solid & Reliable Design | Bright, High-Contrast Display | 10 Easy-to-Use Built-In Applications | Quick & Simple Leveling
The sensor truly stands out with its unparalleled design based on the electromagnetic force compensation (EMFC). It provides an exceptional performance that meets the highest of the set international standards. | Built with a solid, flameproof enclosure of cast aluminum alloy LM6 and epoxy polyester powder coating, featuring push-button operation and IP66 ingress protection, suitable for Zone 1 and 2, Gas Group IIB, ensuring safe and reliable performance in hazardous environments. | Excellent readable display with bright LED-backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | An easy-to-view circular bubble level and large adjustable feet make balance leveling quick, simple, and precise for stable, accurate weighing.
Specifications - Flameproof Balance CY 513FC
Maximum Capacity | 510 g
Readability | 0.001 g
Verification unit [e] | 10 mg
Repeatability | 0.001 g
Linearity | 0.002 g
Tare Range (Subtractive) | -510 g
Standard repeatability [5% Max] | 0.5 mg
Minimum Weight USP | 1 g
Minimum weight ((U=1%, k=2) | 0.1 g
Weighing Pan Diameter | 128 X 128 mm
Setting Time | 2- 3 Sec
Display | LCD Backlit
Calibration | Motorised Internal Calibration (automatic/PSAC)
Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
Operating temperature | 15º to 35ºC
Relative humidity | 40% ÷ 80%
Sensitivity Drift | ± 2ppm (2 x 10-6 / ºC)
Protection class | IP 43
Interface | RS232 Optional
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold
Power Supply | AC 220V / 50-60 Hz
Compliance | ISO, GLP, GMP, USP
Balance Dimension (W x D X H) | 342.5 x 212 x 350 mm
Packaging dimensions (W x D X H) | 490 × 400 × 520 mm
Net weight | 8 kg
Gross weight | 12 kg
Flameproof Balance CY 513F
CY-F Series flameproof precision balances are built with a solid, chemical-resistant cast aluminum alloy LM6 body and epoxy polyester powder coating. Offering 510 g capacity and 0.001 g readability, they feature a bright backlit LCD, external calibration, easy leveling, and push-button operation. With IP66 ingress protection and optional RS232, they are CMFR-certified for Zone 1 and 2, Gas Group IIB, ensuring safe and reliable performance in hazardous environments.
| Weighing Capacity | 510 g
| Readability | 0.001g
| Minimum Weight (USP, 0.1%, typical) | 1 g
Brilliant Proven Technology | Solid & Reliable Design | Bright, High-Contrast Display | 10 Easy-to-Use Built-In Applications | Quick & Simple Leveling
The sensor truly stands out with its unparalleled design based on the electromagnetic force compensation (EMFC). It provides an exceptional performance that meets the highest of the set international standards. | Built with a solid, flameproof enclosure of cast aluminum alloy LM6 and epoxy polyester powder coating, featuring push-button operation and IP66 ingress protection, suitable for Zone 1 and 2, Gas Group IIB, ensuring safe and reliable performance in hazardous environments. | Excellent readable display with bright LED-backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | An easy-to-view circular bubble level and large adjustable feet make balance leveling quick, simple, and precise for stable, accurate weighing.
Specifications - Flameproof Balance CY 513F
Maximum Capacity | 510 g
Readability | 0.001 g
Verification unit [e] | 10 mg
Repeatability | 0.001 g
Linearity | 0.002 g
Tare Range (Subtractive) | -510 g
Standard repeatability [5% Max] | 0.5 mg
Minimum Weight USP | 1 g
Minimum weight ((U=1%, k=2) | 0.1 g
Weighing Pan Diameter | 128 X 128 mm
Setting Time | 2- 3 Sec
Display | LCD Backlit
Calibration | External Calibration
Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
Operating temperature | 15º to 35ºC
Relative humidity | 40% ÷ 80%
Sensitivity Drift | ± 2ppm (2 x 10-6 / ºC)
Protection class | IP 43
Interface | RS232 Optional
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold
Power Supply | AC 220V / 50-60 Hz
Compliance | ISO, GLP, GMP, USP
Balance Dimension (W x D X H) | 342.5 x 212 x 350 mm
Packaging dimensions (W x D X H) | 490 × 400 × 520 mm
Net weight | 8 kg
Gross weight | 12 kg
Flameproof Balance CY 4102F
CY-F Series flameproof precision balances are built with a solid, chemical-resistant cast aluminum alloy LM6 body and epoxy polyester powder coating. Offering 4100 g capacity and 0.01 g readability, they feature a bright backlit LCD, external calibration, easy leveling, and push-button operation. With IP66 ingress protection and optional RS232, they are CMFR-certified for Zone 1 and 2, Gas Group IIB, ensuring safe and reliable performance in hazardo:us environments.
| Weighing Capacity | 4100 g
| Readability | 0.01g
| Minimum Weight (USP, 0.1%, typical) | 10 g
Brilliant Proven Technology | Solid & Reliable Design | Bright, High-Contrast Display | 10 Easy-to-Use Built-In Applications | Quick & Simple Leveling
The sensor truly stands out with its unparalleled design based on the electromagnetic force compensation (EMFC). It provides an exceptional performance that meets the highest of the set international standards. | Built with a solid, flameproof enclosure of cast aluminum alloy LM6 and epoxy polyester powder coating, featuring push-button operation and IP66 ingress protection, suitable for Zone 1 and 2, Gas Group IIB, ensuring safe and reliable performance in hazardous environments. | Excellent readable display with bright LED-backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | An easy-to-view circular bubble level and large adjustable feet make balance leveling quick, simple, and precise for stable, accurate weighing.
Specifications - Flameproof Balance CY 4102F
Maximum Capacity | 4100 g
Readability | 0.01 g
Verification unit [e] | 0.1 g
Repeatability | 0.01 g
Linearity | 0.02 g
Tare Range (Subtractive) | -4100 g
Standard repeatability [5% Max] | 0.05 g
Minimum Weight USP | 10 g
Minimum weight ((U=1%, k=2) | 1 g
Weighing Pan Diameter | 190 x 205 mm
Setting Time | 2- 3 Sec
Display | LCD Backlit
Calibration | External Calibration
Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
Operating temperature | 15º to 35ºC
Relative humidity | 40% ÷ 80%
Sensitivity Drift | ± 2ppm (2 x 10-6 / ºC)
Protection class | IP 43
Interface | RS232 Optional
Built-in Application | Simple weighing Counting %percentage weighting Custom Unit Dynamic (Animal Weighing) Check weighting Formulation Totalization Density Statistics Peak Hold
Power Supply | AC 220V / 50-60 Hz
Compliance | ISO, GLP, GMP, USP
Balance Dimension (W x D X H) | 342.5 x 212 x 350 mm
Packaging dimensions (W x D X H) | 490 × 400 × 520 mm
Net weight | 8 kg
Gross weight | 12 kg

## Moisture Analyzers
Moisture Analyzers MB 54
Designed for fast and precise moisture determination, this halogen moisture analyzer offers ultra-high accuracy (0.0001 % MC / 0.1 mg), adjustable 50–175 °C heating profiles, and a large backlit graphical LCD.
| Weighing Capacity | 50 g
| Readability | 0.1 mg
| Readability MC | 0.0001%
Reliable and Precise Results | Ultra Fast and Uniform Heating | Advanced Temperature & Analysis Control | Various heating profile & Switch off criteria | Data Management | Graphical Display | GLP/ISO Compliance | Real-Time Graphical Analysis | Seamless Connectivity | Easy Maintenance & Cleaning
On-screen guidance and a simple three-step process ensure easy and accurate moisture analysis. EMFC technology delivers high precision, while overload protection ensures long-term reliability, Quick & Simple Leveling | Integrated halogen SRA and reflector design ensures rapid and uniform heating. Temperature range from 30°C to 175°C (optional up to 250°C) for precise moisture analysis. | Adjustable temperature settings enable accurate and repeatable drying performance. Offers versatile analysis modes such as Auto, Manual, Time-based, and user-defined criteria for flexible operation. | Standard, Gentle, Steps, Rapid & High. It ia a great choice for drying process and achieving reliable outcome. Various switch off criteria Auto, Manual, Time-defined, User Def Weight / time, User def %U / Time, Intelligent for accurate and reliable result. | Supports up to 4 user IDs for customized and secure operation. Allows creation of up to 40 methods with storage of 25 results per method. Batch-wise storage enables up to 100 batches with a maximum 1000 data storage for efficient tracking and traceability. | Bright graphical LED-backlit LCD display with large digits ensures clear and easy readability. High-contrast LCD glass provides wide viewing angles for better visibility from different positions. Intuitive user interface enables simple, quick, and comfortable operation. | Built-in real-time clock ensures accurate time-stamped data for traceable documentation in compliance with SOPs and GLP regulations. | Graphical display enables real-time visualization of the drying process with auto-updating moisture trends. Helps users monitor changes in moisture levels for better control and accurate analysis. | The USB and RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data. | Washable parts dismantle in seconds without tools. Dishwasher-safe components and a smooth housing design make routine cleaning quick, simple, and hygienic for everyday laboratory use.
| Specifications -MB 54 | Specifications - MB 50 | Specifications - MB 120 | Specifications -MB 200
Maximum Capacity | 50 g | 50 g | 120 g | 200 g
Readability | 0.1 mg | 0.001 g | 0.001 g | 0.001 g
Moisture Content Readability | 0.0001% | 0.001% | 0.001% | 0.001%
Repeatability (Std Dev) initial sample weight = 2g | 0.05% | 0.1% | 0.1% | 0.1%
Repeatability (Std Dev)for initial sample weight = 10g | 0.02% | 0.03% | 0.03% | 0.03%
Statistics | Method wise for each individual 40 methods or Batch wise for each individual 100 batches for last 1000 results | Method wise for each individual 40 methods or Batch wise for each individual 100 batches for last 1000 results | Method wise for each individual 40 methods or Batch wise for each individual 100 batches for last 1000 results | Method wise for each individual 40 methods or Batch wise for each individual 100 batches for last 1000 results
Data Storage | Last 1000 Moisture Determination results Data Storage | Last 1000 Moisture Determination results Data Storage | Last 1000 Moisture Determination results Data Storage | Last 1000 Moisture Determination results Data Storage
Display units of Measurement | %M, %D, %R (ratio), gm/lt. (g/I), %M ATRO, %D ATRO | %M, %D, %R (ratio), gm/lt. (g/I), %M ATRO, %D ATRO | %M, %D, %R (ratio), gm/lt. (g/I), %M ATRO, %D ATRO | %M, %D, %R (ratio), gm/lt. (g/I), %M ATRO, %D ATRO
Heating Source | Halogen | Halogen | Halogen | Halogen
Temperature range | 30ºC to 175ºC (*Optional 250ºC) | 30ºC to 175ºC (*Optional 250ºC) | 30ºC to 175ºC (*Optional 250ºC) | 30ºC to 175ºC (*Optional 250ºC)
Temperature increment | 1ºC | 1ºC | 1ºC | 1ºC
Weighing Pan Diameter | Ø 95 mm | Ø 95 mm | Ø 95 mm | Ø 95 mm
Display | Graphical Alpha Numeric LCD Backlit | Graphical Alpha Numeric LCD Backlit | Graphical Alpha Numeric LCD Backlit | Graphical Alpha Numeric LCD Backlit
Heating Profile | Standard, Gentle, Rapid, Steps, High Temperature | Standard, Gentle, Rapid, Steps, High Temperature | Standard, Gentle, Rapid, Steps, High Temperature | Standard, Gentle, Rapid, Steps, High Temperature
Switch off Criteria | Auto, Manual, Time, User Def Weight /time, User Def % unit/time, Intelligent | Auto, Manual, Time, User Def Weight /time, User Def % unit/time, Intelligent | Auto, Manual, Time, User Def Weight /time, User Def % unit/time, Intelligent | Auto, Manual, Time, User Def Weight /time, User Def % unit/time, Intelligent
Auto Method Development | Intelligent | Intelligent | Intelligent | Intelligent
Method Memory | 40 | 40 | 40 | 40
Batches Memory | 100 | 100 | 100 | 100
Method Designation | Alpha Numeric | Alpha Numeric | Alpha Numeric | Alpha Numeric
Serial Number | Yes | Yes | Yes | Yes
Test | Weighing and Heater | Weighing and Heater | Weighing and Heater | Weighing and Heater
Leveling system | Manual | Manual | Manual | Manual
Interface | RS232 & USB , PS/2 | RS232 & USB , PS/2 | RS232 & USB , PS/2 | RS232 & USB , PS/2
Power Supply | 230V/50Hz or 115V/60Hz | 230V/50Hz or 115V/60Hz | 230V/50Hz or 115V/60Hz | 230V/50Hz or 115V/60Hz
Compliance | GLP / GMP Compliance, User Configurable | GLP / GMP Compliance, User Configurable | GLP / GMP Compliance, User Configurable | GLP / GMP Compliance, User Configurable
Balance Dimension (W x D X H) | 211 x 342 x 180 mm | 211 x 342 x 180 mm | 211 x 342 x 180 mm | 211 x 342 x 180 mm
Packaging dimensions (W x D X H) | 420 x 350 x 320 mm | 420 x 350 x 320 mm | 420 x 350 x 320 mm | 420 x 350 x 320 mm
Net weight | 5 kg | 5 kg | 5 kg | 5 kg
Gross weight | 9 kg | 9 kg | 9 kg | 9 kg

## Density Scales
Density Scales CY 613GT
Digital density scales offer high-precision measurements with a 610 g capacity and 0.001 g/cm³ resolution. Using Archimedes’ principle, they accurately determine density, purity, and material quality. Designed for labs and industrial use, they feature automatic calculations, guided operation, temperature compensation, and data connectivity via USB or RS-232 for efficiency.
| Weighing Capacity | 610 g
| Readability | 0.001 g
| Resolution | 0.001 g/cm³
Quick and Easy Operation | Brilliant Proven Technology | Changing Liquid Medium | Easy to read LCD Screen | Water Temperature Compensation | Seamless Connectivity | ISO/GLP Documentation | Easy Maintenance & Cleaning
The CY-GT Series Electronic Densimeter is engineered for simplified and accurate density analysis. Featuring intuitive on-screen guidance, it enables reliable results through a streamlined two-step measurement process: • Step 1 – Measurement of weight in air • Step 2 – Measurement of weight in water | The sensor is designed based on electromagnetic force compensation (EMFC), ensuring high precision and stable performance. It delivers reliable results in accordance with established international standards. | The densimeter supports the use of alternative liquid media to accommodate challenging samples, such as those with air bubbles or complex geometries, ensuring accurate and reliable density measurements. H₂O / Ethanol / Other | The balance features a large backlit LCD display that ensures clear visibility under all lighting conditions. The integrated backlight enhances readability, enabling accurate viewing even in low-light environments. | Designed to ensure high accuracy by compensating for temperature variations, delivering precise density values at the defined reference temperature. | The USB and RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data. | The built-in real-time clock ensures accurate, time-stamped data, providing traceable documentation that reflects true test results and supports compliance with SOPs and GLP regulations. | The draft shields are designed for easy dismantling and reassembly, ensuring convenient cleaning and maintaining a hygienic working environment.
| Specifications - Density Scales CY 204GT | Specifications - Density Scales CY 613GT | Specifications - Density Scales CY 3102GT
Maximum Capacity | 220 g | 610 g | 3100 g
Readability | 0.0001 g | 0.001 g | 0.01 g
Density Resolution | 0.0001 g/m³ | 0.001 g/m³ | 0.01 g/m³
Repeatability | 0.1 mg | 0.001 g | 0.01 g
Linearity | 0.3 g | 0.003 g | 0.03 g
Tare Range (Subtractive) | -220g | - 610 g | -3100 g
Weight in Air Pan Size | 30 Ø mm | 85 x 106 mm | 124 x 160 mm
Weight in Water Pan Size | 60 Ø mm | 92 x 94 mm | 128 x 141 mm
Setting Time | 2- 3 Sec | 2- 3 Sec | 2- 3 Sec
Display | LCD Backlit | LCD Backlit | LCD Backlit
Calibration | External Calibration | External Calibration | External Calibration
Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
Operating temperature | 15° to 35°C | 15° to 35°C | 15° to 35°C
Leveling system | manual | manual | manual
Interface | RS232/USB | RS232/USB | RS232/USB
Power Supply | DC Adaptor, input 100-240 0.8A output 13V/15.A, 50-60Hz | DC Adaptor, input 100-240 0.8A output 13V/15.A, 50-60Hz | DC Adaptor, input 100-240 0.8A output 13V/15.A, 50-60Hz
Compliance | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP | ISO, GLP, GMP, USP
Balance Dimension (W x D X H) | 342.5 x 212 x 270 mm | 342.5 x 212 x 270 mm | 342.5 x 212 x 270 mm
Packaging dimensions (W x D X H) | 500 x 350 x 540 mm | 500 x 350 x 540 mm | 500 x 350 x 540 mm
Net weight | 10 kg | 10 kg | 10 kg
Gross weight | 14 Kg | 14 Kg | 14 Kg

## Desk Scales
Desk Scales CG 6
A desk scale is a compact digital weighing device designed for accurate measurement, featuring durable construction, clear display, easy operation, multiple units, and reliable performance for commercial and industrial applications.
| Capacity | 6 kg
| Readability | 0.5 gm
| Pan Size | 275 x 220 mm
Precision Sensor | Weighing Performance | Robust Housing | High-Contrast Display | Power Continuity
Precision aluminum load cell, IP65 rated, delivers excellent repeatability and measurement stability for routine and industrial weighing applications. Engineered for measurement consistency with robust performance in routine applications. | Fast weighing response time ensures instant stabilization, while the weight stability indicator confirms stable readings for accurate and reliable measurement results. | Rugged ABS housing with a corrosion-resistant stainless steel platter ensures mechanical strength and durability. A 6-key color-coded keyboard enables ergonomic operation, while 150% overload protection and adjustable feet ensure stable, accurate setup and safe industrial use. | High-contrast backlit LCD with graphical guidance ensures clear readability, while auto power-off standby mode minimizes energy consumption during periods of inactivity. | In-built 6V / 1.3 Ah rechargeable battery ensures uninterrupted operation during power interruptions, while the low battery indicator provides advance alerts to prevent unexpected downtime.
| Specifications - Desk Scales CG 3 | Specifications - Desk Scales CG 6 | Specifications - Desk Scales CG 30
Maximum Capacity | 3 kg | 6 kg | 30 kg
Readability | 0.1 g | 0.5 g | 1 g
Repeatability | 0.1 g | 0.5 g | 1 g
Linearity | 0.2 g | 1 g | 2 g
Tare Range (Subtractive) | - 3 kg | - 6 kg | - 30 kg
Weighing Pan Diameter | 275 x 220 mm | 275 x 220 mm | 275 x 220 mm
Setting Time | 2- 3 Sec | 2- 3 Sec | 2- 3 Sec
Display | LCD Backlit | LCD Backlit | LCD Backlit
Calibration | Eternal Calibration | Eternal Calibration | Eternal Calibration
Units of measure | kg, g, lb, oz | kg, g, lb, oz | kg, g, lb, oz
Operating temperature | 10ºC to 40ºC | 10ºC to 40ºC | 10ºC to 40ºC
Protection class | IP43 | IP43 | IP43
Interface | RS232 | RS232 | RS232
Built-in Application | Counting, Percent Weighing, Dynamic Weighing, Check weighing | Counting, Percent Weighing, Dynamic Weighing, Check weighing | Counting, Percent Weighing, Dynamic Weighing, Check weighing
Battery | 6V 1.2Ah rechargeable battery | 6V 1.2Ah rechargeable battery | 6V 1.2Ah rechargeable battery
Power Supply | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz
Balance Dimension (W x D X H) | 300 x 290 x 90 mm | 300 x 290 x 90 mm | 300 x 290 x 90 mm
Packaging dimensions (W x D X H) | 320 x 390 x 150 mm | 320 x 390 x 150 mm | 320 x 390 x 150 mm
Net weight | 4 kg | 4 kg | 4 kg
Gross weight | 4.5 kg | 4.5 kg | 4.5 kg
Counting Desk Scales CG 6N
Counting desk scale designed for precise weighing and counting, featuring clear display, stable platform, easy leveling, durable construction, and RS-232C connectivity for accurate, reliable, efficient industrial measurement operations use.
| Capacity | 6 kg
| Readability | 0.5 gm
| Pan Size | 275 x 220 mm
Advanced Sensor Technology | Three-View LCD Display | Robust Housing | Weighing Performance | Power Continuity
Precision aluminum load cell, IP65 rated, delivers excellent repeatability and measurement stability for routine and industrial weighing applications. Engineered for measurement consistency with robust performance in routine applications. | Dedicated three window LCD displays unit weight, total weight, and piece count simultaneously, enabling faster reading, improved counting efficiency, clear visibility, and accurate monitoring during continuous industrial operation use. | Rugged ABS housing with a corrosion-resistant stainless steel platter ensures mechanical strength and durability. A key colour-coded keyboard enables ergonomic operation, while 150% overload protection and adjustable feet ensure stable, accurate setup and safe industrial use. | Fast weighing response time ensures instant stabilization, while the weight stability indicator confirms stable readings for accurate and reliable measurement results. | In-built 6V / 1.3 Ah rechargeable battery ensures uninterrupted operation during power interruptions, while the low battery indicator provides advance alerts to prevent unexpected downtime.
| Specifications - Desk Scales CG 3N | Specifications - Desk Scales CG 6N | Specifications - Desk Scales CG 30N
Maximum Capacity | 3 kg | 6 kg | 30 kg
Readability | 0.1 g | 0.5 g | 1 g
Repeatability | 0.1 g | 0.5 g | 1 g
Linearity | 0.2 g | 1 g | 2 g
Tare Range (Subtractive) | - 3 kg | - 6 kg | - 30 kg
Weighing Pan Diameter | 275 x 220 mm | 275 x 220 mm | 275 x 220 mm
Setting Time | 2- 3 Sec | 2- 3 Sec | 2- 3 Sec
Display | LCD Backlit | LCD Backlit | LCD Backlit
Calibration | Eternal Calibration | Eternal Calibration | Eternal Calibration
Units of measure | kg, g, lb, oz | kg, g, lb, oz | kg, g, lb, oz
Operating temperature | 10ºC to 40ºC | 10ºC to 40ºC | 10ºC to 40ºC
Protection class | IP43 | IP43 | IP43
Interface | RS232 | RS232 | RS232
Built-in Application | Counting, Percent Weighing, Check weighing | Counting, Percent Weighing, Check weighing | Counting, Percent Weighing, Check weighing
Battery | 6V 1.2Ah rechargeable battery | 6V 1.2Ah rechargeable battery | 6V 1.2Ah rechargeable battery
Power Supply | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz
Balance Dimension (W x D X H) | 300 x 290 x 90 mm | 300 x 290 x 90 mm | 300 x 290 x 90 mm
Packaging dimensions (W x D X H) | 320 x 390 x 150 mm | 320 x 390 x 150 mm | 320 x 390 x 150 mm
Net weight | 4 kg | 4 kg | 4 kg
Gross weight | 4.5 kg | 4.5 kg | 4.5 kg
Stainless Steel Desk Scales CG 3S
Stainless steel desk scale engineered for accuracy and durability, with IP67 protection, clear display, precise weighing, easy leveling, and multiple unit support for reliable performance in industrial and commercial applications.
| Capacity | 3 kg
| Readability | 0.1 gm
| Pan Size | 260 x 210 mm
Advanced Sensor Technology | Rugged Industrial Design | Weighing Performance | High-Contrast Display | Power Continuity
Precision aluminum load cell, IP66 rated, delivers excellent repeatability and measurement stability for routine and industrial weighing applications. Engineered for measurement consistency with robust performance in routine applications. | Rugged grade 304 stainless steel housing ensures durability, sealed keypad prevents dirt and spills, stainless steel weighing platter ensures hygienic stable measurement surface, and overload safety supports 150 percent capacity reliability. | Fast weighing response time ensures instant stabilization, while the weight stability indicator confirms stable readings for accurate and reliable measurement results. | High-contrast backlit LCD with graphical guidance ensures clear readability, while auto power-off standby mode minimizes energy consumption during periods of inactivity. | In-built 6V / 1.3 Ah rechargeable battery ensures uninterrupted operation during power interruptions, while the low battery indicator provides advance alerts to prevent unexpected downtime.
| Specifications - Desk Scales CG 3S | Specifications - Desk Scales CG 6S | Specifications - Desk Scales CG 30S
Maximum Capacity | 3 kg | 6 kg | 30 kg
Readability | 0.1 g | 0.5 g | 1 g
Repeatability | 0.1 g | 0.5 g | 1 g
Tare Range (Subtractive) | -3 kg | -6 kg | -30 kg
Weighing Pan Diameter | 260 x 210 mm | 260 x 210 mm | 260 x 210 mm
Setting Time | 2 - 3 sec | 2 - 3 sec | 2 - 3 sec
Display | LCD Backlit | LCD Backlit | LCD Backlit
Calibration | External | External | External
Units of measure | Counting, Custom Unit | Counting, Custom Unit | Counting, Custom Unit
Operating temperature | 10ºC to 40ºC | 10ºC to 40ºC | 10ºC to 40ºC
Protection class | IP 67 | IP 67 | IP 67
Interface | RS232 | RS232 | RS232
Built-in Application | Counting | Counting | Counting
Battery | 6V 1.2Ah rechargeable battery | 6V 1.2Ah rechargeable battery | 6V 1.2Ah rechargeable battery
Power Supply | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz
Balance Dimension (W x D X H) | 330 x 290 x 110 mm | 330 x 290 x 110 mm | 330 x 290 x 110 mm
Packaging dimensions (W x D X H) | 310 x 410 x 140 mm | 310 x 410 x 140 mm | 310 x 410 x 140 mm
Net weight | 3.5 kg | 3.5 kg | 3.5 kg
Gross weight | 4 kg | 4 kg | 4 kg
Water Proof Desk Scales CG 6W
A waterproof desk scale is a digital weighing device with IP68 protection, designed for wet, humid, and hygienic environments such as food processing, fish markets, pharma, and chemical industries applications.
| Capacity | 6 kg
| Readability | 2 gm
| Pan Size | 190 x 260 mm
Advanced Sensor Technology | Rugged Industrial Design | Weighing Performance | High-Contrast Display | Power Efficiency
Precision aluminum load cell, IP657 rated, delivers excellent repeatability and measurement stability for routine and industrial weighing applications. Engineered for measurement consistency with robust performance in routine applications. | Rugged grade 304 stainless steel housing ensures durability, sealed keypad prevents dirt and spills, stainless steel weighing platter ensures hygienic stable measurement surface, and overload safety supports 150 percent capacity reliability. | Fast weighing response time ensures instant stabilization, while the weight stability indicator confirms stable readings for accurate and reliable measurement results. | High-contrast backlit LED display ensures clear readability in all lighting conditions, while auto power-off standby mode reduces energy consumption during inactivity, improving efficiency and extending overall battery life performance. | Built-in rechargeable backup battery and automatic shutoff features extend battery life, ensuring reliable performance and convenient everyday operation anywhere anytime.
| Specifications - Desk Scales CG 6W | Specifications - Desk Scales CG 15W
Maximum Capacity | 6 kg | 15 kg
Readability | 1 g | 5g
Repeatability | 1 g | 5g
Linearity | 2 g | 10 g
Tare Range (Subtractive) | - 6 kg | - 15 kg
Weighing Pan Diameter | 190 x 260 mm | 190 x 260 mm
Setting Time | 2 - 3 sec | 2 - 3 sec
Display | Red LED | Red LED
Calibration | External | External
Units of measure | kg, g, lb, oz | kg, g, lb, oz
Operating temperature | 10ºC to 40ºC | 10ºC to 40ºC
Protection class | IP68 | IP68
Built-in Application | Counting | Counting
Battery | 6V 1.2Ah rechargeable battery | 6V 1.2Ah rechargeable battery
Power Supply | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz
Balance Dimension (W x D X H) | 297 x 246 x 127 mm | 297 x 246 x 127 mm
Packaging dimensions (W x D X H) | 320 x 390 x 150 mm | 320 x 390 x 150 mm
Net weight | 3.5 kg | 3.5 kg
Gross weight | 4 kg | 4 kg
Flame-Proof Desk Scales CG-3SF
Designed for hazardous areas, this flameproof desk scale delivers precise weighing with 15,000 counts resolution, IP66 protection, and certified safe operation for Zone 1 and Zone 2 environments.
| Capacity | 3 kg
| Readability | 0.1 gm
| Pan Size | 250 x 210 mm
Advanced Sensor Technology | Flameproof certified indicator | Weighing Performance | High-Contrast Display
Precision aluminum load cell, IP65 rated, delivers excellent repeatability and measurement stability for routine and industrial weighing applications. Engineered for measurement consistency with robust performance in routine applications. | Built with a solid, flameproof enclosure of cast aluminum alloy LM6 and epoxy polyester powder coating, featuring push-button operation and IP66 ingress protection, suitable for Zone 1 and 2, Gas Group IIB, ensuring safe and reliable performance in hazardous environments. | Fast weighing response time ensures instant stabilization, while the weight stability indicator confirms stable readings for accurate and reliable measurement results. | High-contrast backlit LED display ensures clear readability, auto power-off saves energy, improves efficiency, extends battery life performance overall.
| Specifications - Desk Scales CG 3SF | Specifications - Desk Scales CG 6SF | Specifications - Desk Scales CG 30SF
Maximum Capacity | 3 kg | 6 kg | 30 kg
Readability | 0.1 g | 0.5 g | 1 g
Repeatability | 0.1 g | 0.5 g | 1 g
Linearity | 0.2 g | 1 g | 2 g
Tare Range (Subtractive) | -3 kg | - 6 kg | -30 kg
Weighing Pan Diameter | 260 x 210 mm | 260 x 210 mm | 260 x 210 mm
Setting Time | 2 - 3 sec | 2 - 3 sec | 2 - 3 sec
Display | Red LED | Red LED | Red LED
Calibration | External | External | External
Units of measure | kg | kg | kg
Operating temperature | 10ºC to 40ºC | 10ºC to 40ºC | 10ºC to 40ºC
Protection class | IP 43 | IP 43 | IP 43
Power Supply | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz
Balance Dimension (W x D X H) | 330 x 290 x 110 mm | 297 x 246 x 127 mm | 330 x 290 x 110 mm
Packaging dimensions (W x D X H)
Net weight | 5 kg | 5 kg | 5 kg
Gross weight | 6 kg | 6 kg | 6 kg

## Bench & Platform Scales
Industrial Precision Scale CTG 100K1
CTG 100K1 Series Industrial Precision Scales combine rugged construction, high-capacity weighing, and precise measurements, making them ideal for heavy-duty applications in food, chemical, and manufacturing industries.
| Capacity | 100 kg
| Readability | 1 g
| Pan Size | 400 x 500 mm
Rugged Construction | Precision Loadcell | LCD Display | Fast & Flexible | Safety & Power | Compact & User-Friendly
Mild steel frame, powder-coated finish, stainless steel platter, corrosion-resistant, durable, industrial-grade, compact bench model, robust design, long-lasting, reliable operation. | Aluminum sensor, IP65 protection, high accuracy, stable readings, repeatable results, top-to-bottom measurement, external calibration option, sensitive detection, weight stability, reliable calibration. | Backlit, high-visibility, fast response, numeric clarity, user-friendly, large digits, alpha numeric display, stable readings, easy operation, bright contrast. | Quick measurement (2–3 sec), real-time display, efficient workflow, instant weighing, dynamic weighing support, multiple modes: counting, percent, checkweighing, custom units, optimized load response, industrial-speed performance. | Overload protection (150% capacity), protective design, safe operation, energy-saving standby mode, built-in rechargeable battery (6V 1.3 Ah), emergency-ready, uninterrupted operation, low-maintenance, reliable power. | Space-saving bench model, easy placement, lightweight, ergonomic design, 6-key color-coded keypad, intuitive interface, smooth navigation, minimal training, organized layout, consistent operation.
| Specifications - CTG 100K1 | Specifications - CTG 200K2 | Specifications - CTG 500K10
Maximum Capacity | 100 kg | 200 kg | 500 kg
Readability | 1 g | 2 g | 10 g
Repeatability | 1 g | 2 g | 10 g
Linearity | 2 g | 4 g | 20 g
Tare Range (Subtractive) | - 100 kg | - 200 kg | - 500 kg
Weighing Pan Diameter | 400 x 500 mm | 400 x 500 mm | 400 x 500 mm
Setting Time | 2 - 3 sec | 2 - 3 sec | 2 - 3 sec
Display | LCD Backlit | LCD Backlit | LCD Backlit
Calibration | External | External | External
Units of measure | kg,lb,oz,g | kg,lb,oz,g | kg,lb,oz,g
Operating temperature | 5ºC to 40ºC | 5ºC to 40ºC | 5ºC to 40ºC
Protection class | IP 65 | IP 65 | IP 65
Interface | RS232 | RS232 | RS232
Built-in Application | Counting, Totalization, Check weighing | Counting, Totalization, Check weighing | Counting, Totalization, Check weighing
Power Supply | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz
Balance Dimension (W x D X H) | 450 x 550 x 100 mm | 450 x 550 x 100 mm | 450 x 550 x 100 mm
Packaging dimensions (W x D X H)
Net weight | 15 kg | 15 kg | 15 kg
Gross weight
Bench & platform Scale (MS Version) CTG 60B
Heavy-duty industrial bench & platform scale with stainless steel platter and powder-coated mild steel frame, offering high accuracy, corrosion resistance, and durability. Features LCD display, RS-232 interface, rechargeable battery backup, ideal for warehouse, manufacturing, and commercial use.
| Capacity | 60 kg
| Readability | 5 g
| Pan Size | 500 x 500 mm
Advanced Sensor Technology | Precision & Reliable Operation | Robust Housing | High-Contrast Display | Power Continuity | Seamless Connectivity
Precision aluminum load cell, IP65 rated, delivers excellent repeatability and measurement stability for routine and industrial weighing applications. Engineered for measurement consistency with robust performance in routine applications. | Advanced load cell technology ensures stable, accurate, and repeatable measurements. Fast response time and weight stability indication improve operational efficiency. | Heavy-duty bench platform scale with stainless steel platter and powder-coated mild steel bottom frame, offering durability, corrosion resistance, accurate weighing, and stable, reliable performance for industrial, warehouse and commercial applications. | High-contrast backlit LCD with graphical guidance ensures clear readability, while auto power-off standby mode minimizes energy consumption during periods of inactivity. | In-built 6V / 1.3 Ah rechargeable battery ensures uninterrupted operation during power interruptions, while the low battery indicator provides advance alerts to prevent unexpected downtime. | The RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data.
| Specifications - CTG 60B | Specifications - CTG 100B | Specifications - CTG 150 | Specifications - CTG 300 | Specifications - CTG 500 | Specifications - CTG 501
Maximum Capacity | 60 kg | 100 kg | 150 kg | 300 kg | 500 kg | 500 kg
Readability | 5 g | 10 g | 10 g | 20 g | 50 g | 100 g
Repeatability | 5 g | 10 g | 10 g | 20 g | 50 g | 100 g
Linearity | 10 g | 20 g | 20 g | 40 g | 100 g | 200 g
Tare Range (Subtractive) | - 60 kg | - 100 kg | - 100 kg | -300 | -500 | -500
Weighing Pan Diameter | 500 x 500 mm | 500 x 500 mm | 600 x 600 mm | 600 x 600 mm | 600 x 600 mm | 750 x 750 mm
Setting Time | 2 - 3 sec | 2 - 3 sec | 2 - 3 sec | 2 - 3 sec | 2 - 3 sec | 2 - 3 sec
Display | LCD Backlit | LCD Backlit | LCD Backlit | LCD Backlit | LCD Backlit | LCD Backlit
Calibration | External | External | External | External | External | External
Units of measure | kg,lb,oz,g | kg,lb,oz,g | kg,lb,oz,g | kg,lb,oz,g | kg,lb,oz,g | kg,lb,oz,g
Operating temperature | 5ºC to 40ºC | 5ºC to 40ºC | 5ºC to 40ºC | 5ºC to 40ºC | 5ºC to 40ºC | 5ºC to 40ºC
Protection class | IP 43 | IP 43 | IP 43 | IP 43 | IP 43 | IP 43
Interface | RS232 | RS232 | RS232 | RS232 | RS232 | RS232
Built-in Application | Counting, Totalization, Check weighing | Counting, Totalization, Check weighing | Counting, Totalization, Check weighing | Counting, Totalization, Check weighing | Counting, Totalization, Check weighing | Counting, Totalization, Check weighing
Power Supply | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz
Balance Dimension (W x D X H) | 500 x 500 x 100 mm | 500 x 500 x 100 mm | 600 x 600 x 100 | 600 x 600 x 100 | 600 x 600 x 100 | 750 x 750 x 100
Packaging dimensions (W x D X H) | 570 x 570 x 260 mm | 570 x 570 x 260 mm | 700 x 690 x 230 mm | 700 x 690 x 230 mm | 700 x 690 x 230 mm | 700 x 690 x 230 mm
Net weight | 15 kg | 15 kg | 20 kg | 20 kg | 20 kg | 25 kg
Gross weight
Fully stainless steel Bench Scale (SS Version) CTG 60BS
Heavy-duty industrial bench & platform scale with full stainless steel structure and a capsulated load cell is a fully sealed to protect internal strain gauges from water, dust, moisture, and contaminants, ensuring reliable, long-term performance in harsh or washdown environments.. offering high accuracy, corrosion resistance, and durability. Features LCD display, RS232 interface, and rechargeable battery backup, ideal for food, pharma, chemical, warehouse, manufacturing, and commercial use.
| Capacity | 60 kg
| Readability | 5 g
| Pan Size | 500 x 500 mm
Advanced Sensor Technology | Precision & Reliable Operation | Robust Housing | High-Contrast Display | Power Continuity | Seamless Connectivity
Precision aluminum load cell, IP66 rated, delivers excellent repeatability and measurement stability for routine and industrial weighing applications. Engineered for measurement consistency with robust performance in routine applications. | Advanced load cell technology ensures stable, accurate, and repeatable measurements. Fast response time and weight stability indication improve operational efficiency. | Built for durability with a corrosion-resistant stainless steel platform, ensuring easy cleaning, hygiene, and long-term performance. Ideal for food processing, pharmaceutical, chemical, and industrial applications requiring reliable and precise weighing. | High-contrast backlit LCD with graphical guidance ensures clear readability, while auto power-off standby mode minimizes energy consumption during periods of inactivity. | In-built 6V / 1.3 Ah rechargeable battery ensures uninterrupted operation during power interruptions, while the low battery indicator provides advance alerts to prevent unexpected downtime. | The RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data.
| Specifications - CTG 60BS | Specifications - CTG 150BS | Specifications - CTG 300BS
Maximum Capacity | 60 kg | 150 kg | 300 kg
Readability | 5 g | 10 g | 20 g
Repeatability | 5 g | 10 g | 20 g
Linearity | 10 g | 20 g | 40 g
Tare Range (Subtractive) | - 60 kg | - 150 kg | - 300 kg
Weighing Pan Diameter | 500 x 500 mm | 600 x 600 mm | 600 x 600 mm
Setting Time | 2 - 3 sec | 2 - 3 sec | 2 - 3 sec
Display | LCD Backlit | LCD Backlit | LCD Backlit
Calibration | External | External | External
Units of measure | kg,lb,g | kg,lb,g | kg,lb,g
Operating temperature | 5ºC to 40ºC | 5ºC to 40ºC | 5ºC to 40ºC
Protection class | IP67 | IP67 | IP67
Interface | RS232 | RS232 | RS232
Built-in Application | Counting | Counting | Counting
Power Supply | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz
Balance Dimension (W x D X H) | 500 x 500 x 100 mm | 600 x 600 x 100 mm | 600 x 600 x 100 mm
Packaging dimensions (W x D X H)
Net weight | 15 kg | 15 kg | 15 kg
Gross weight
Counting Platform Scale CTG-300N
Heavy-duty industrial bench & platform scale with stainless steel platter and powder-coated mild steel frame, three-window LCD display, RS232 interface, rechargeable battery backup, ideal for warehouse, manufacturing, and commercial use.
| Capacity | 300 kg
| Readability | 20 g
| Pan Size | 600 x 600 mm
Advanced Sensor Technology | Three-View LCD Display | Robust Housing | Weighing Performance | Power Continuity | Seamless Connectivity
Precision aluminum load cell, IP65 rated, delivers excellent repeatability and measurement stability for routine and industrial weighing applications. Engineered for measurement consistency with robust performance in routine applications. | Dedicated three window LCD displays unit weight, total weight, and piece count simultaneously, enabling faster reading, improved counting efficiency, clear visibility, and accurate monitoring during continuous industrial operation use. | Heavy-duty bench platform scale with stainless steel platter and powder-coated mild steel bottom frame, offering durability, corrosion resistance, accurate weighing, and stable, reliable performance for industrial, warehouse and commercial applications. | Fast weighing response time ensures instant stabilization, while the weight stability indicator confirms stable readings for accurate and reliable measurement results. | In-built 6V / 1.3 Ah rechargeable battery ensures uninterrupted operation during power interruptions, while the low battery indicator provides advance alerts to prevent unexpected downtime. | The RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data.
| Specifications - CTG 60N | Specifications - CTG 150N | Specifications - CTG 300N
Maximum Capacity | 60 kg | 150 kg | 300 kg
Readability | 5 g | 10 g | 20 g
Repeatability | 5 g | 10 g | 20 g
Linearity | 10 g | 20 g | 40 g
Tare Range (Subtractive) | - 60 kg | - 100 kg | -300
Weighing Pan Diameter | 500 x 500 mm | 600 x 600 mm | 600 x 600 mm
Setting Time | 2 - 3 sec | 2 - 3 sec | 2 - 3 sec
Display | LCD Backlit | LCD Backlit | LCD Backlit
Calibration | External | External | External
Units of measure | kg,lb,oz,g | kg,lb,oz,g | kg,lb,oz,g
Operating temperature | 5ºC to 40ºC | 5ºC to 40ºC | 5ºC to 40ºC
Protection class | IP43 | IP43 | IP43
Interface | RS232 | RS232 | RS232
Built-in Application | Counting, Totalization, Check weighing | Counting, Totalization, Check weighing | Counting, Totalization, Check weighing
Power Supply | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz
Balance Dimension (W x D X H) | 500 x 500 x 100 mm | 600 x 600 x 100 mm | 600 x 600 x 100 mm
Packaging dimensions (W x D X H)
Net weight | 15 kg | 20 kg | 20 kg
Gross weight
Airport Luggage Scales
Airport Baggage Scale CTG Series provides precision weighing, reliable load management, and operational stability. Engineered for airports, it ensures accurate baggage measurement, streamlined passenger processing, and robust performance.
| Capacity | 300 kg
| Readability | 20 g
| Pan Size | 600 x 600 mm
High Precision | Stable Performance | Efficient Cooling | Operator-Friendly Interface | Safe & Compliant
Equipped with advanced load cells, the scale delivers ±100 g readability and ultra-accurate measurements, ensuring compliance with airport standards. | Optimized electronics and mechanical design provide consistent output and reliable operation across high-traffic check-in counters. | Integrated intelligent cooling system maintains optimal operating temperature, preventing sensor drift and ensuring sustained performance under continuous load. | High-contrast LED display, intuitive keypad, and menu-driven software allow seamless operation and quick access to functional settings. | OIML-approved design with overload protection, audible alarms, and RS-232 interface for safe, regulatory-compliant baggage handling.
| Specifications -CTG 150L | Specifications - CTG 300L
Maximum Capacity | 150 kg | 300 kg
Readability | 100g | 100g
Repeatability
Linearity
Tare Range (Subtractive) | - 150 kg | - 300 kg
Platform size | 600 x 800 mm | 600 x 800 mm
Setting Time | 2 - 3 sec | 2 - 3 sec
Display |  |  |  | `
Calibration | External | External
Units of measure | kg,lb,oz,g | kg,lb,oz,g
Operating temperature | 5ºC to 40ºC | 5ºC to 40ºC
Protection class
Interface
Built-in Application
Power Supply | 230Vac, Single Phase | 230Vac, Single Phase
Balance Dimension (W x D X H)
Packaging dimensions (W x D X H)
Net weight
Gross weight
Floor scales - Mild Steel CTG-1 T4L
Heavy-duty floor scale with MS chequered plate platform and powder-coated structure, equipped with IP67-rated 4 single-ended shear beam load cells to combine four load cell signals using an IP68-rated waterproof junction box, for high accuracy and durability. Features LCD display and RS232 interface, ideal for warehouse and industrial applications.
| Capacity | 1 Ton
| Readability | 100 g
| Pan Size | 1000 x 1000 mm
Advanced Sensor Technology | Robust Housing | Precision & Reliable Operation | Overload Protection | High-Contrast Display | Power Continuity | Seamless Connectivity
IP67-rated shear beam load cells with stainless steel construction, to combine four load cell signals using an IP68-rated waterproof junction box offering high precision, excellent corrosion resistance, strong overload protection, and long-term stability. Designed for reliable and consistent weighing performance in demanding industrial environments. | Heavy-duty floor scale with MS chequered plate platform and powder-coated structure, ensuring robust strength, long-lasting durability, corrosion resistance, and precise weighing with stable, reliable performance for industrial, warehouse, and commercial applications. | Advanced load cell technology ensures stable, accurate, and repeatable measurements. Fast response time and weight stability indication improve operational efficiency. | Equipped with built-in overload protection up to 150% of the scale capacity to prevent damage and ensure long-term durability, even under extreme load conditions. | High-contrast backlit LCD with graphical guidance ensures clear readability, while auto power-off standby mode minimizes energy consumption during periods of inactivity. | In-built 6V / 1.3 Ah rechargeable battery ensures uninterrupted operation during power interruptions, while the low battery indicator provides advance alerts to prevent unexpected downtime. | The RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data.
| Specifications - CTG 1T4L | Specifications - CTG 1T4L2 | Specifications - CTG 2T4L | Specifications - CTG 2T4L5 | Specifications - CTG 3T4L | Specifications - CTG 5T4L | Specifications - CTG 5T4L2
Maximum Capacity | 1 Ton | 1 Ton | 2 Ton | 2 Ton | 3 Ton | 5 Ton | 5 Ton
Readability | 100 g | 100 g | 500 g | 500 g | 500 g | 1 kg | 2 kg
Repeatability | 100 g | 100 g | 500 g | 500 g | 500 g | 1 kg | 2 kg
Linearity | 200 g | 200 g | 1 kg | 1 kg | 1 kg | 2 kg | 4 kg
Tare Range (Subtractive) | -1 Ton | -1 Ton | -2 Ton | -2 Ton | -3 Ton | -5 Ton | -5 Ton
Weighing Pan Diameter | 1000 x 1000 mm | 1200 x 1200 mm | 1200 x 1200 mm | 1500 x 1500 mm | 1500 x 1500 mm | 1500 x 1500 mm | 2000 x 2000 mm
Setting Time | 2 - 3 sec | 2 - 3 sec | 2 - 3 sec | 2 - 3 sec | 2 - 3 sec | 2 - 3 sec | 2 - 3 sec
Display | LCD Backlit | LCD Backlit | LCD Backlit | LCD Backlit | LCD Backlit | LCD Backlit | LCD Backlit
Calibration | External | External | External | External | External | External | External
Units of measure | kg,lb,oz,g | kg,lb,oz,g | kg,lb,oz,g | kg,lb,oz,g | kg,lb,oz,g | kg,lb,oz,g | kg,lb,oz,g
Operating temperature | 5ºC to 40ºC | 5ºC to 40ºC | 5ºC to 40ºC | 5ºC to 40ºC | 5ºC to 40ºC | 5ºC to 40ºC | 5ºC to 40ºC
Protection class | IP43 | IP43 | IP43 | IP43 | IP43 | IP43 | IP43
Interface | RS232 | RS232 | RS232 | RS232 | RS232 | RS232 | RS232
Built-in Application | Counting, Totalization, Check weighing | Counting, Totalization, Check weighing | Counting, Totalization, Check weighing | Counting, Totalization, Check weighing | Counting, Totalization, Check weighing | Counting, Totalization, Check weighing | Counting, Totalization, Check weighing
Power Supply | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz
Balance Dimension (W x D X H) | 1000 x 1000 x 100 mm | 1200 x 1200 x 100 mm | 1200 x 1200 x 100 mm | 1500 x 1500 x 100 mm | 1500 x 1500 x 100 mm | 1500 x 1500 x 100 mm | 2000 x 2000 x 100 mm
Packaging dimensions (W x D X H)
Net weight | 50 kg | 60 kg | 60 kg | 75 kg | 75 kg | 75 kg | 85 kg
Gross weight
Floor scales - Stainless Steel CTG- 1T4L S
Heavy-duty floor scale with fully stainless steel platform structure, equipped with IP67-rated 4 single-ended shear beam load cells to combine four load cell signals using an IP68-rated waterproof junction box, for high accuracy and durability. Features LCD display and RS232 interface, ideal for warehouse and industrial applications.
| Capacity | 1 Ton
| Readability | 100 g
| Pan Size | 1200 x 1200 mm
Advanced Sensor Technology | Robust Housing | Precision & Reliable Operation | Overload Protection | High-Contrast Display | Power Continuity | Seamless Connectivity
IP67-rated shear beam load cells with stainless steel construction, to combine four load cell signals using an IP68-rated waterproof junction box offering high precision, excellent corrosion resistance, strong overload protection, and long-term stability. Designed for reliable and consistent weighing performance in demanding industrial environments. | Built for durability with a corrosion-resistant stainless steel platform, ensuring easy cleaning, hygiene, and long-term performance. Ideal for food processing, pharmaceutical, chemical, and industrial applications requiring reliable and precise weighing. | Advanced load cell technology ensures stable, accurate, and repeatable measurements. Fast response time and weight stability indication improve operational efficiency. | Equipped with built-in overload protection up to 150% of the scale capacity to prevent damage and ensure long-term durability, even under extreme load conditions. | High-contrast backlit LCD with graphical guidance ensures clear readability, while auto power-off standby mode minimizes energy consumption during periods of inactivity. | In-built 6V / 1.3 Ah rechargeable battery ensures uninterrupted operation during power interruptions, while the low battery indicator provides advance alerts to prevent unexpected downtime. | The RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data.
| Specifications - CTG 1T4LS | Specifications - CTG 1T4L2S | Specifications - CTG 2T4LS | Specifications - CTG 2T4L5S | Specifications - CTG 3T4LS | Specifications - CTG 5T4LS | Specifications - CTG 5T4L2S
Maximum Capacity | 1 Ton | 1 Ton | 2 Ton | 2 Ton | 3 Ton | 5 Ton | 5 Ton
Readability | 100 g | 100 g | 500 g | 500 g | 500 g | 1 kg | 2 kg
Repeatability | 100 g | 100 g | 500 g | 500 g | 500 g | 1 kg | 2 kg
Linearity | 200 g | 200 g | 1 kg | 1 kg | 1 kg | 2 kg | 4 kg
Tare Range (Subtractive) | -1 Ton | -1 Ton | -2 Ton | -2 Ton | -3 Ton | -5 Ton | -5 Ton
Weighing Pan Diameter | 1200 x 1200 mm | 1200 x 1200 mm | 1200 x 1200 mm | 1500 x 1500 mm | 1500 x 1500 mm | 1500 x 1500 mm | 2000 x 2000 mm
Setting Time | 2 - 3 sec | 2 - 3 sec | 2 - 3 sec | 2 - 3 sec | 2 - 3 sec | 2 - 3 sec | 2 - 3 sec
Display | LCD Backlit | LCD Backlit | LCD Backlit | LCD Backlit | LCD Backlit | LCD Backlit | LCD Backlit
Calibration | External | External | External | External | External | External | External
Units of measure | kg,lb,oz,g | kg,lb,oz,g | kg,lb,oz,g | kg,lb,oz,g | kg,lb,oz,g | kg,lb,oz,g | kg,lb,oz,g
Operating temperature | 5ºC to 40ºC | 5ºC to 40ºC | 5ºC to 40ºC | 5ºC to 40ºC | 5ºC to 40ºC | 5ºC to 40ºC | 5ºC to 40ºC
Protection class | IP68 | IP68 | IP68 | IP68 | IP68 | IP68 | IP68
Interface | RS232 | RS232 | RS232 | RS232 | RS232 | RS232 | RS232
Built-in Application | Counting, Totalization, Check weighing | Counting, Totalization, Check weighing | Counting, Totalization, Check weighing | Counting, Totalization, Check weighing | Counting, Totalization, Check weighing | Counting, Totalization, Check weighing | Counting, Totalization, Check weighing
Power Supply | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz
Balance Dimension (W x D X H) | 1000 x 1000 x 100 | 1200 x 1200 x 100 | 1200 x 1200 x 100 | 1500 x 1500 x 100 | 1500 x 1500 x 100 | 1500 x 1500 x 100 | 2000 x 2000 x 100
Packaging dimensions (W x D X H)
Net weight | 50 kg | 60 kg | 60 kg | 75 kg | 75 kg | 75 kg | 85 kg
Gross weight
Cylinder Scale CTG-T4LC
Cylinder Scale is designed for accurate and safe gas cylinder weighing with robust mild steel construction and high-precision shear beam load cells, ensuring stable performance, durability, and reliable weight measurement in industrial environments.
| Capacity | 2000 kg
| Readability | 500 g
| Pan Size | 1200 x 2000 mm
Industrial-Grade Construction | Precision Performance | Precision & Reliable Operation | Overload Protection | High-Contrast Display | Power Continuity | Seamless Connectivity
Beam Construction – Provides strength for repeated high-impact loading in harsh environments Engineered to handle continuous use in harsh industrial environments without compromising on accuracy. | Double-Ended Shear Beam Load Cells ensure accurate and stable weight measurement under heavy industrial loads. Equipped with 4 load cells for high-capacity applications and integrated junction box support for reliable signal transmission and performance. | Advanced load cell technology ensures stable, accurate, and repeatable measurements. Fast response time and weight stability indication improve operational efficiency. | Equipped with built-in overload protection up to 150% of the scale capacity to prevent damage and ensure long-term durability, even under extreme load conditions. | High-contrast backlit LCD with graphical guidance ensures clear readability, while auto power-off standby mode minimizes energy consumption during periods of inactivity. | In-built 6V / 1.3 Ah rechargeable battery ensures uninterrupted operation during power interruptions, while the low battery indicator provides advance alerts to prevent unexpected downtime. | The RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data.
| Specifications - CTG 2T 4LC | Specifications - CTG 3T 4LC
Maximum Capacity | 2000 kg | 3000 kg
Readability | 500 g | 1 kg
Repeatability | 500 g | 1 kg
Linearity | 1200 x 2000 mm | 1200 x 2000 mm
Tare Range (Subtractive) | 5 sec | 5 sec
Weighing Pan Diameter | LCD Backlit | LCD Backlit
Setting Time | External | External
Display | kg,lb,g | kg,lb,g
Calibration | 5ºC to 40ºC | 5ºC to 40ºC
Units of measure | IP 43 | IP 43
Operating temperature | RS232 | RS232
Protection class | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz
Leveling system | 1200 x 2000 x 150 mm | 1200 x 2000 x 150 mm
Interface | RS232 | RS232
Built-in Application | Counting, Totalization, Check weighing | Counting, Totalization, Check weighing
Power Supply | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz
Balance Dimension (W x D X H) | 1200 x 2000 x 150 mm | 1200 x 2000 x 150 mm
Packaging dimensions (W x D X H)
Net weight | 100 kg | 100 kg
Gross weight
Coil Weighing Scale WB 30CW
The HI-IMPACT Heavy-Duty Industrial Scale is a high-capacity weighing system built for extreme manufacturing environments. Engineered specifically for steel mills, foundries, and scrap yards, this system eliminates structural warping caused by dense, concentrated point loads. With a robust 30-ton capacity and advanced shock-absorption engineering, it delivers repeatable, legal-for-trade precision despite daily high-impact drops from overhead cranes and heavy forklifts.
| Capacity | 30 Ton
| Readability | 5 kg
| Pan Size | 1500 x 2000 mm
Durable Industrial Construction | Precision Performance | Robust Housing | Overload Protection | High-Contrast Display | Power Continuity | Seamless Connectivity
WF-Beam Weighbridge Construction – Provides strength for repeated high-impact loading in harsh environments Engineered to handle continuous use in harsh industrial environments without compromising on accuracy. | Double-Ended Shear Beam Load Cells ensure accurate and stable weight measurement under heavy industrial loads. Equipped with 4 load cells for high-capacity applications and integrated junction box support for reliable signal transmission and performance. | Heavy-duty floor scale with MS chequered plate platform and powder-coated structure, ensuring robust strength, long-lasting durability, corrosion resistance, and precise weighing with stable, reliable performance for industrial, warehouse, and commercial applications. | Equipped with built-in overload protection up to 150% of the scale capacity to prevent damage and ensure long-term durability, even under extreme load conditions. | High-contrast backlit LCD with graphical guidance ensures clear readability, while auto power-off standby mode minimizes energy consumption during periods of inactivity. | In-built 6V / 1.3 Ah rechargeable battery ensures uninterrupted operation during power interruptions, while the low battery indicator provides advance alerts to prevent unexpected downtime. | The RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data.
Specifications - WB 30CW
Maximum Capacity | 30 Ton
Readability | 5 kg
Repeatability | 5 kg
Linearity | 10 kg
Tare Range (Subtractive) | 5 kg
Weighing Pan Diameter | 1500 x 2000
Setting Time | 5 sec
Display | LCD Backlit
Calibration | External
Units of measure | kg, lb
Operating temperature | 5ºC to 40ºC
Protection class | IP 43
Interface | RS232
Built-in Application | Counting, Totalization, Check weighing
Power Supply | AC 220V / 50-60 Hz
Balance Dimension (W x D X H) | 1500 x 2000 x 150 mm
Packaging dimensions (W x D X H)
Net weight | 700 kg
Gross weight
Low Profile Floor Scale - Mild steel CTG- 600 4L U.M.S
The Ultra Low Profile Scale in mild Steel is designed for easy loading, high durability, and precise weighing in industrial environments. Its compact low-height platform ensures smooth material handling, while the stainless steel construction provides excellent corrosion resistance, hygiene, and long-lasting performance for heavy-duty applications.
| Capacity | 600 kg
| Readability | 100 g
| Pan Size | 600 x 600 mm
Advanced Sensor Technology | Robust Housing | Precision & Reliable Operation | Overload Protection | High-Contrast Display | Power Continuity | Seamless Connectivity
IP67-rated shear beam load cells with stainless steel construction, to combine four load cell signals using an IP68-rated waterproof junction box offering high precision, excellent corrosion resistance, strong overload protection, and long-term stability. Designed for reliable and consistent weighing performance in demanding industrial environments. | Heavy-duty floor scale with MS chequered plate platform and powder-coated structure, ensuring robust strength, long-lasting durability, corrosion resistance, and precise weighing with stable, reliable performance for industrial, warehouse, and commercial applications. | Advanced load cell technology ensures stable, accurate, and repeatable measurements. Fast response time and weight stability indication improve operational efficiency. | Equipped with built-in overload protection up to 150% of the scale capacity to prevent damage and ensure long-term durability, even under extreme load conditions. | High-contrast backlit LCD with graphical guidance ensures clear readability, while auto power-off standby mode minimizes energy consumption during periods of inactivity. | In-built 6V / 1.3 Ah rechargeable battery ensures uninterrupted operation during power interruptions, while the low battery indicator provides advance alerts to prevent unexpected downtime. | The RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data.
| Specifications - CTG 600K4LUMS | Specifications - CTG 1T4LUMS | Specifications - CTG 2T4LUMS
Maximum Capacity | 600 kg | 1 Ton | 2 Ton
Readability | 100 g | 200 g | 500 g
Repeatability | 200 g | 200 g | 500 g
Linearity | 200 g | 400 g | 1 kg
Weighing Pan Diameter | 600 x 600 mm | 1000 x 1000 mm | 1200 x 1200 mm
Setting Time | 5 sec | 5 sec | 5 sec
Display | LCD Backlit | LCD Backlit | LCD Backlit
Calibration | External | External | External
Units of measure | Kg, lb, oz, g | Kg, lb, oz, g | Kg, lb, oz, g
Operating temperature | 5ºC to 40ºC | 5ºC to 40ºC | 5ºC to 40ºC
Protection class | IP67 | IP67 | IP67
Interface | RS232 | RS232 | RS232
Built-in Application | Counting, Totalization, Check weighing | Counting, Totalization, Check weighing | Counting, Totalization, Check weighing
Power Supply | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz
Balance Dimension (W x D X H) | 600 x 900 x 55 mm | 1000 x 1300 x 55 mm | 1200 x 1500 x 55 mm
Packaging dimensions (W x D X H)
Net weight | 30 kg | 70 kg | 90 kg
Gross weight
Ultra Low Profile Scale - Stainless Steel CTG 600K 4L USS
The Ultra Low Profile Scale in Stainless Steel is designed for easy loading, high durability, and precise weighing in industrial environments. Its compact low-height platform ensures smooth material handling, while the stainless steel construction provides excellent corrosion resistance, hygiene, and long-lasting performance for heavy-duty applications.
| Capacity | 600 kg
| Readability | 100 g
| Pan Size | 600 x 600 mm
Advanced Sensor Technology | Robust Housing | Precision & Reliable Operation | Overload Protection | High-Contrast Display | Power Continuity | Seamless Connectivity
IP67-rated shear beam load cells with stainless steel construction, to combine four load cell signals using an IP68-rated waterproof junction box offering high precision, excellent corrosion resistance, strong overload protection, and long-term stability. Designed for reliable and consistent weighing performance in demanding industrial environments. | Built for durability with a corrosion-resistant stainless steel platform, ensuring easy cleaning, hygiene, and long-term performance. Ideal for food processing, pharmaceutical, chemical, and industrial applications requiring reliable and precise weighing. | Advanced load cell technology ensures stable, accurate, and repeatable measurements. Fast response time and weight stability indication improve operational efficiency. | Equipped with built-in overload protection up to 150% of the scale capacity to prevent damage and ensure long-term durability, even under extreme load conditions. | High-contrast backlit LCD with graphical guidance ensures clear readability, while auto power-off standby mode minimizes energy consumption during periods of inactivity. | In-built 6V / 1.3 Ah rechargeable battery ensures uninterrupted operation during power interruptions, while the low battery indicator provides advance alerts to prevent unexpected downtime. | The RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data.
| Specifications - CTG 600K 4L USS | Specifications - CTG 1 T4L USS | Specifications - CTG 2 T4L USS
Maximum Capacity | 600 kg | 1 Ton | 2 Ton
Readability | 100 g | 200 g | 500 g
Repeatability | 200 g | 200 g | 500 g
Linearity | 200 g | 400 g | 1 kg
Weighing Pan Diameter | 600 x 600 mm | 1000 x 1000 mm | 1200 x 1200 mm
Setting Time | 5 sec | 5 sec | 5 sec
Display | LCD Backlit | LCD Backlit | LCD Backlit
Calibration | External | External | External
Units of measure | Kg, lb, oz, g | Kg, lb, oz, g | Kg, lb, oz, g
Operating temperature | 5ºC to 40ºC | 5ºC to 40ºC | 5ºC to 40ºC
Protection class | IP67 | IP67 | IP67
Interface | RS232 | RS232 | RS232
Built-in Application | Counting, Totalization, Check weighing | Counting, Totalization, Check weighing | Counting, Totalization, Check weighing
Power Supply | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz
Balance Dimension (W x D X H) | 600 x 900 x 55 mm | 1000 x 1300 x 55 mm | 1200 x 1500 x 55 mm
Packaging dimensions (W x D X H)
Net weight | 30 kg | 70 kg | 90 kg
Gross weight
Pallet Scales CTG 2PT
Pallet Jack Scales and Pallet Scales with robust mild steel construction provide accurate weighing up to 2-ton capacity, ensuring durability, smooth pallet handling, enhanced safety, space-saving operation, and reliable long-term performance in demanding industrial and warehouse environments.
| Capacity | 2 Ton
| Readability | 2 Kg
| Construction | Mild Steel
Advanced Sensor Technology | Robust Housing | Precision & Reliable Operation | Overload Protection | High-Contrast Display | Power Continuity | Seamless Connectivity
IP67-rated shear beam load cells with stainless steel construction, offering high precision, excellent corrosion resistance, strong overload protection, and long-term stability. Designed for reliable and consistent weighing performance in demanding industrial environments. | Heavy-duty floor scale with mild steel construction and powder-coated structure, ensuring robust strength, long-lasting durability, corrosion resistance, Ensures easy pallet movement and efficient material handling for warehouse and logistics operations. | Advanced load cell technology ensures stable, accurate, and repeatable measurements. Fast response time and weight stability indication improve operational efficiency. | Equipped with built-in overload protection up to 150% of the scale capacity to prevent damage and ensure long-term durability, even under extreme load conditions. | High-contrast backlit LCD with graphical guidance ensures clear readability, while auto power-off standby mode minimizes energy consumption during periods of inactivity. | In-built 6V / 1.3 Ah rechargeable battery ensures uninterrupted operation during power interruptions, while the low battery indicator provides advance alerts to prevent unexpected downtime. | The RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data.
Specifications - CTG 2PT
Maximum Capacity | 2 Ton
Readability | 2 Kg
Repeatability | 2 Kg
Fork Length (mm) | 1150
Outer Width (mm) | 540
Fork Max Lifting Height | 85 ~ 200 mm
Setting Time | 5 sec
Display | LCD Backlit
Calibration | External
Units of measure | Kg, lb
Operating temperature | 5ºC to 40ºC
Protection class | IP67
Leveling system
Interface | RS232
Power Supply | AC 220V / 50-60 Hz
Balance Dimension (W x D X H) | 540 x 1150 x 85 ~ 200 mm
Packaging dimensions (W x D X H)
Net weight | 90 kg
Gross weight
Pallet Scales CTG 2PTS
Pallet Jack Scales and Pallet Scales with durable stainless steel construction provide accurate weighing up to 2-ton capacity, ensuring corrosion resistance, hygienic operation, smooth pallet handling, and reliable performance in industrial, warehouse, and pharmaceutical applications.
| Capacity | 2 Ton
| Readability | 2 Kg
| Construction | Stainlees steal
Advanced Sensor Technology | Robust Housing | Precision & Reliable Operation | Overload Protection | High-Contrast Display | Power Continuity | Seamless Connectivity
IP67-rated shear beam load cells with stainless steel construction, offering high precision, excellent corrosion resistance, strong overload protection, and long-term stability. Designed for reliable and consistent weighing performance in demanding industrial environments. | Heavy-duty floor scale with mild steel construction and powder-coated structure, ensuring robust strength, long-lasting durability, corrosion resistance, Ensures easy pallet movement and efficient material handling for pharma , chemical , food warehouse and logistics operations. | Advanced load cell technology ensures stable, accurate, and repeatable measurements. Fast response time and weight stability indication improve operational efficiency. | Equipped with built-in overload protection up to 150% of the scale capacity to prevent damage and ensure long-term durability, even under extreme load conditions. | High-contrast backlit LCD with graphical guidance ensures clear readability, while auto power-off standby mode minimizes energy consumption during periods of inactivity. | In-built 6V / 1.3 Ah rechargeable battery ensures uninterrupted operation during power interruptions, while the low battery indicator provides advance alerts to prevent unexpected downtime. | The RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data.
Specifications - CTG 2PTS
Maximum Capacity | 2 Ton
Readability | 2 Kg
Repeatability | 2 Kg
Fork Length (mm) | 1150
Outer Width (mm) | 540
Fork Max Lifting Height | 85 ~ 200 mm
Setting Time | 5 sec
Display | LCD Backlit
Calibration | External
Units of measure | Kg, lb
Operating temperature | 5ºC to 40ºC
Protection class | IP68
Leveling system
Interface | RS232
Power Supply | AC 220V / 50-60 Hz
Balance Dimension (W x D X H) | 540 x 1150 x 85 ~ 200 mm
Packaging dimensions (W x D X H)
Net weight | 90 kg
Gross weight
Hanging Crane Scale HA 2T
Designed for heavy-duty industrial applications, this weighing system offers capacities up to 50 tons with rugged construction, 15,000 counts resolution, and wireless remote operation for accurate, reliable, and durable performance in steel, logistics, manufacturing, shipping, and warehouse environments.
| Capacity | 2 Ton
| Readability | 500 g
| Display | (1.2 ”) Bright Red LED display
Heavy-Duty Industrial Construction | High-Visibility LED Display | Wireless Remote Operation | High-Precision Weighing Performance | Advanced Safety & Reliability | Industrial Application Efficiency
Engineered with a rugged mechanical structure, high-tensile hook, and robust shackle for secure and dependable suspension weighing. Designed to perform reliably in demanding industrial and construction environments with capacities up to 100 tons. | Large bright red LED display ensures clear and accurate weight readings from long viewing distances. Optimized for overhead crane operations and industrial weighing applications. | Supplied with an infrared wireless remote control for safe and convenient operation from a distance. Supports remote tare, zero, hold, and calibration functions for enhanced operational efficiency. | Equipped with sealed high-precision IP67 protected S-Type load cells for stable, accurate, and repeatable measurements. Fast response time and weight hold functionality ensure reliable performance in continuous industrial operations. | High-strength hook assembly with automatic lock latch enhances safety during heavy load handling. Integrated overload alarm and indication system provide additional protection against excessive loading conditions. | Ideal for steel plants, warehouses, foundries, logistics, shipping yards, and heavy material handling industries. Provides accurate and efficient suspension weighing for cranes, hoists, and overhead lifting systems.
| Specifications - HA2T | Specifications - HA3T | Specifications - HA5T | Specifications - HA10T | Specifications - HA20T | Specifications - HA30T | Specifications - HA50T
Maximum Capacity | 2 Ton | 3 Ton | 5 Ton | 10 Ton | 20 Ton | 30 Ton | 50 Ton
Readability | 500 g | 1 kg | 2 kg | 5 kg | 10 kg | 10 kg | 20 kg
Display | RED LED | RED LED | RED LED | RED LED | RED LED | RED LED | RED LED
Units of measure | Kg | Kg | Kg | Kg | Kg | Kg | Kg
Protection class | IP67 | IP67 | IP67 | IP67 | IP67 | IP67 | IP67
Battery | Rechargeable Battery 6V / 10Ah, Approx 100 Hours Backup Time | Rechargeable Battery 6V / 10Ah, Approx 100 Hours Backup Time | Rechargeable Battery 6V / 10Ah, Approx 100 Hours Backup Time | Rechargeable Battery 6V / 10Ah, Approx 100 Hours Backup Time | Rechargeable Battery 6V / 10Ah, Approx 100 Hours Backup Time | Rechargeable Battery 6V / 10Ah, Approx 100 Hours Backup Time | Rechargeable Battery 6V / 10Ah, Approx 100 Hours Backup Time
Power Supply | AC 230V + 10% 50-60Hz | AC 230V + 10% 50-60Hz | AC 230V + 10% 50-60Hz | AC 230V + 10% 50-60Hz | AC 230V + 10% 50-60Hz | AC 230V + 10% 50-60Hz | AC 230V + 10% 50-60Hz
Balance Dimension (W x D X H)
Packaging dimensions (W x D X H)
Net weight | 16kg | 16kg | 18kg | 20kg | 25Kg | 30Kg | 40Kg
Gross weight
Wireless Crane Scales HA 5TR
Wireless Crane Scale is designed for safe and accurate suspended load weighing in industrial environments. Featuring wireless remote operation, rugged construction, high-capacity performance, and bright digital display, it ensures reliable weighing for steel, shipping, logistics, warehouse, and heavy material handling applications.
| Capacity | 5 Ton
| Readability | 500 g
| Display | (1.2 ”) Bright Red LED
Heavy-Duty Industrial Construction | High-Contrast Display | Wireless Remote Operation | High-Precision Weighing Performance | Advanced Safety & Reliability | Industrial Application Efficiency
Engineered with a rugged mechanical structure, high-tensile hook, and robust shackle for secure and dependable suspension weighing. Designed to perform reliably in demanding industrial and construction environments with capacities up to 100 tons. | High-contrast backlit LCD with graphical guidance ensures clear readability, while auto power-off standby mode minimizes energy consumption during periods of inactivity. | Supplied with an infrared wireless remote control for safe and convenient operation from a distance. Supports remote tare, zero, hold, and calibration functions for enhanced operational efficiency. | Equipped with sealed high-precision IP67 protected S-Type load cells for stable, accurate, and repeatable measurements. Fast response time and weight hold functionality ensure reliable performance in continuous industrial operations. | High-strength hook assembly with automatic lock latch enhances safety during heavy load handling. Integrated overload alarm and indication system provide additional protection against excessive loading conditions. | Ideal for steel plants, warehouses, foundries, logistics, shipping yards, and heavy material handling industries. Provides accurate and efficient suspension weighing for cranes, hoists, and overhead lifting systems.
| Specifications - HA 5TR | Specifications - HA 10TR | Specifications - HA 20TR | Specifications - HA 30TR
Maximum Capacity | 5 Ton | 10 Ton | 20 Ton | 30 Ton
Readability | 2 kg | 5 kg | 10 kg | 10 kg
Display | LCD backlite | LCD backlite | LCD backlite | LCD backlite
Units of measure | Kg | Kg | Kg | Kg
Protection class | IP67 | IP67 | IP67 | IP67
Battery | Rechargeable Battery 6V / 10Ah, Approx 100 Hours Backup Time | Rechargeable Battery 6V / 10Ah, Approx 100 Hours Backup Time | Rechargeable Battery 6V / 10Ah, Approx 100 Hours Backup Time | Rechargeable Battery 6V / 10Ah, Approx 100 Hours Backup Time
Power Supply | AC 230V + 10% 50-60Hz | AC 230V + 10% 50-60Hz | AC 230V + 10% 50-60Hz | AC 230V + 10% 50-60Hz
Balance Dimension (W x D X H)
Packaging dimensions (W x D X H)
Net weight | 16kg | 16kg | 18kg | 20kg
Gross weight
Tank Weighing Systems WB 1TN
Tank Weighing Scale is designed for accurate and reliable weighing of storage tanks, mixing vessels, hoppers, and silos in industrial applications. Equipped with high-precision load cells and a robust structure, it ensures stable performance, efficient process control, and long-term durability in manufacturing, chemical, food, and pharmaceutical industries.
| Capacity | 1 Ton
| Readability | 500 g
| Display | (1.2 ”) Bright Red LED
Advanced Sensor Technology | Load cell Mounting | High-Contrast Display | Optional Indicator | Power Continuity | Seamless Connectivity
IP67-rated shear beam load cells with stainless steel construction, to combine four load cell signals using an IP68-rated waterproof junction box offering high precision, excellent corrosion resistance, strong overload protection, and long-term stability. Designed for reliable and consistent weighing performance in demanding industrial environments. | high-precision load cells with robust mild steel and stainless steel mounting assemblies, ensuring accurate weighing, structural stability, corrosion resistance, and reliable long-term performance in demanding industrial applications. | High-contrast LED display ensures clear readability, while auto power-off standby mode minimizes energy consumption during periods of inactivity. | Stainless Steel Indicator 4–20 mA Output Interface Flame Proof Indicator | In-built 6V / 1.3 Ah rechargeable battery ensures uninterrupted operation during power interruptions, while the low battery indicator provides advance alerts to prevent unexpected downtime. | The RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data.
| Specifications - WB 1TN | Specifications - WB 2TN | Specifications - WB 5TN | Specifications - WB 10TN | Specifications - WB 20TN | Specifications - WB 30TN | Specifications - WB 50TN | Specifications - WB 100TN | Specifications -WB 150TN
Maximum Capacity | 1 Ton | 2 Ton | 5 Ton | 10 Ton | 20 Ton | 30 Ton | 50 Ton | 100 Ton | 150 Ton
Readability | 500 g | 1 kg | 2 kg | 5 kg | 10 kg | 10 kg | 20 kg | 20 Kg | 30 Kg
Display | (1.2 ”) Bright Red LED display | (1.2 ”) Bright Red LED display | (1.2 ”) Bright Red LED display | (1.2 ”) Bright Red LED display | (1.2 ”) Bright Red LED display | (1.2 ”) Bright Red LED display | (1.2 ”) Bright Red LED display | (1.2 ”) Bright Red LED display | (1.2 ”) Bright Red LED display
Units of measure | Kg, ltr | Kg, ltr | Kg, ltr | Kg, ltr | Kg, ltr | Kg, ltr | Kg, ltr | Kg, ltr | Kg, ltr
Protection class | IP67 | IP67 | IP67 | IP67 | IP67 | IP67 | IP67 | IP67 | IP67
Battery | Rechargeable Battery 6V / 10Ah, Approx 100 Hours Backup Time | Rechargeable Battery 6V / 10Ah, Approx 100 Hours Backup Time | Rechargeable Battery 6V / 10Ah, Approx 100 Hours Backup Time | Rechargeable Battery 6V / 10Ah, Approx 100 Hours Backup Time | Rechargeable Battery 6V / 10Ah, Approx 100 Hours Backup Time | Rechargeable Battery 6V / 10Ah, Approx 100 Hours Backup Time | Rechargeable Battery 6V / 10Ah, Approx 100 Hours Backup Time | Rechargeable Battery 6V / 10Ah, Approx 100 Hours Backup Time | Rechargeable Battery 6V / 10Ah, Approx 100 Hours Backup Time
Power Supply | AC 230V + 10% 50-60Hz | AC 230V + 10% 50-60Hz | AC 230V + 10% 50-60Hz | AC 230V + 10% 50-60Hz | AC 230V + 10% 50-60Hz | AC 230V + 10% 50-60Hz | AC 230V + 10% 50-60Hz | AC 230V + 10% 50-60Hz | AC 230V + 10% 50-60Hz
Balance Dimension (W x D X H)
Packaging dimensions (W x D X H)
Net weight | 16kg | 16kg | 18kg | 20kg | 25Kg | 30Kg | 40Kg
Gross weight
Weigh in Motion Scales WB 30WIM
Weigh-In-Motion systems accurately weigh moving vehicles, improving traffic flow, preventing overloading, enhancing road safety, ensuring regulatory compliance, and enabling efficient real-time monitoring across highways, toll plazas, logistics, mining, and industries.
| Capacity | 30 Ton
| Readability | 10 kg
| Display | Green VFD
Dynamic Weighing | System Overview | Weighing sensor | Software Features | Space Efficient | Advanced System Features
The weigh-in-motion system measures vehicle weight while moving, overcoming delays and limitations associated with traditional static weighing systems. | System integrates weighing indicator, infrared detection, card reader, cameras, traffic lights, software, and displays to measure trucks, record data securely. | Robust steel structure supports dynamic loads, while high-accuracy load cells ensure stable, precise weight measurement, durability, and reliable performance under continuous vehicle movement. | Advanced software enables timed operations, flexible formats, error correction, anti-cheating, automatic weighing, networking, remote monitoring, high-accuracy conversion, and communication protocols. | Compact platform design minimizes installation space, ideal for retrofit projects and sites where traditional static weighbridges are impractical or difficult to install. | System supports four video inputs, WLAN, large storage, mobile alerts, multi-sensor monitoring, real-time transmission, data security, anti-cheating, and user management.
| Specifications - WB 30WIM | Specifications - WB 50WIM
Maximum Capacity | 30 Ton | 50 Ton
Graduation | 10 kg | 20 kg
Platform size | 3.0 x 0.8 m | 3.2 x 2.2 m
Total Weighing Capacity | 160 Ton | 200 Ton
Verification Class | Class I | Class I
Display | Green VFD | Green VFD
Units of measure | Kg | Kg
Protection class | IP67 | IP67
Number of load cell | 4 | 4
Power Supply | 230 VAC, 50 Hz, Max 15 W | 230 VAC, 50 Hz, Max 15 W
Balance Dimension (W x D X H)
Packaging dimensions (W x D X H)
Net weight
Gross weight
Weighbridge Scales WB 30WIM
Weigh-In-Motion systems accurately weigh moving vehicles, improving traffic flow, preventing overloading, enhancing road safety, ensuring regulatory compliance, and enabling efficient real-time monitoring across highways, toll plazas, logistics, mining, and industries.
| Capacity | 30 Ton
| Readability | 5 kg
| Platform size | 9 x 3 mtr.
Pitless Weighbridge | Pit Type Weighbridge | Weighing sensor | Smart Indicator | Network Integration | Weighbridge Software
Pitless weighbridge offers cost-effective installation without excavation, faster setup, easy access for maintenance, reduced downtime, and reliable performance, enabling businesses to start operations quickly with lower civil and operational costs. | Pit weighbridges save space with ground-level installation, allow smooth ramp-free vehicle movement, reduce wear and tear, and provide a clean, professional appearance suitable for sites with limited space. | A high-accuracy double-ended shear beam load cell ensures superior strength, stability, precise measurement, excellent repeatability, and long-term reliability under heavy loads. | weighbridge indicator with intelligent terminal, digital calibration, large memory, LED/LCD displays, auto zero tracking, report generation, jumbo display support, and CIMER-certified accuracy for reliable operations. | LAN-based wired or wireless connectivity enables long-distance data transfer, eliminates RS232 dependency, supports SAP/ERP integration, and allows multiple indicators on one network. | Advanced weighbridge software offers international standards with statistics and summary reporting features for accurate data analysis and efficient operational management.
| Specifications - WB 30WIM | Specifications - WB 50WIM
Maximum Capacity | 30 Ton | 50 Ton
Graduation | 10 kg | 20 kg
Platform size | 3.0 x 0.8 m | 3.2 x 2.2 m
Total Weighing Capacity | 160 Ton | 200 Ton
Verification Class | Class I | Class I
Display | Green VFD | Green VFD
Units of measure | Kg | Kg
Protection class | IP67 | IP67
Number of load cell | 4 | 4
Power Supply | 230 VAC, 50 Hz, Max 15 W | 230 VAC, 50 Hz, Max 15 W
Balance Dimension (W x D X H)
Packaging dimensions (W x D X H)
Net weight
Gross weight
Portable Truck Scale WB 15PT
A portable weigh pad is a compact, movable weighing system used for axle load measurement of vehicles. It enables quick setup, accurate readings, and flexible roadside or field deployment without permanent installation
| Maximum Capacity ( per pad) | 15 Ton
| Readability | 10 Kg
| Capacity Per Axle | 30 Ton
Precision Vehicle Weighing | Portable Weigh Pad System | Smart Axle Weighing Solution | Heavy Duty Weigh Pads | Fast & Accurate Vehicle Weighing
Reliable axle weighing solution designed for accurate staticvehicle measurement. Lightweight aluminum weigh pads ensure easy portability, fast installation, and dependable performance in industrial and transport applications. | Compact and rugged weigh pads built for efficient on-site vehicle weighing. High-precision load cells, IP65 protection, and durable construction provide accurate results even in demanding working environments. | Engineered for fast and precise axle load measurement with excellent stability and overload protection. Ideal for transportation, logistics, mining, and industrial weighing applications requiring dependable performance. | Advanced portable weigh pads offering high load capacity, reliable accuracy, and simple operation. Designed with durable aluminum structure and weather-resistant protection for long-term field performance. | Portable weighing pads deliver quick setup and accurate axle weight measurement for commercial and industrial vehicles. Built with robust construction, wide temperature operation, and seamless connectivity support.
| Specifications - WB 15PT | Specifications - WB 20PT | Specifications - WB 30PT
Maximum Capacity ( per pad) | 15 Ton | 20 Ton | 30 Ton
Capacity Per Axle | 30 Ton | 40 Ton | 60 Ton
Readability | 10 Kg | 20 Kg | 20 Kg
Linearity | 0.3% | 0.3% | 0.3%
Display | LCD Backlit | LCD Backlit | LCD Backlit
Units of measure | Kg | Kg | Kg
Protection class | IP 65 | IP 65 | IP 65
Interface | RS232 | RS232 | RS232
Weigh Pad Size (LxWxH) | 765 × 468 × 18 (±5mm) | 950 × 550 × 21.5 (±5mm) | 950 × 550 × 21.5 (±5mm)
Active Surface Size (LxWxH) | 714 × 419 × 18 (±5mm) | 898 × 499 × 21.5 (±5mm) | 898 × 499 × 21.5 (±5mm)
Balance Dimension (W x D X H)
Packaging dimensions (W x D X H)
Net weight | 28kg | 28kg | 28kg
Gross weight

## Laboratory Instruments
Analogue Abbe Refractometer 2 AWAJ
The 2WAJ analog Abbe refractometer is a precision benchtop optical instrument for measuring refractive index (nD) and %Brix of liquids and solids, delivering accurate, repeatable results through manual eyepiece observation.
| Refractive index range (nD) | 1.300 – 1.700 nD
| Brix (sugar %) range | 0 – 95 %
| Refractive index accuracy | ± 0.0003 nD
Optical System Design | Digital Temperature Display | Dissolved Solids Brix: | Temperature Control
The optical system uses high-precision primary and secondary glass prisms. The sample is held as a thin layer between them, where light refraction creates a sharp critical-angle boundary for accurate visual measurement. | The digital thermometer provides precise temperature readings, ensuring accurate measurements even under varying environmental conditions. | A Brix measurement range of 0-95%, this refractometer is suitable for measuring the mass fraction of sugar solutions, making it highly useful in the food and beverage industry. | built-in thermometer and provisions for optional connection to an external thermostatic bath to maintain stable and accurate measurement conditions.
Specifications - 2AWAJ
Refractive Index nD Range | 1.300~ 1.700nD
Refractive Index nD Accuracy | ±0.0003nD
Refractive e Index nD Min. | 0.0005nD
Sugar Measuring Range | 0~95% Brix
Sugar Concentration | 0.25%
Observation Method | Binocular
Temperature | RT~70°C
Temperature Division | 1°C
Temperature Display | Digital Thermometer
Dimension (W x D X H) | 100 x 140 × 235 mm
Packaging dimensions (W x D X H) | 240 x200 x 390 mm
Net weight | 2 kg
Gross weight | 3 kg
Semi-Automatic Digital Abbe Refractometer ARM2S
The ARM-2S semi-automatic digital Abbe refractometer is a precision benchtop optical instrument for measuring refractive index (nD) and %Brix of liquid, viscous, and solid samples, delivering accurate and repeatable results through electronic evaluation and digital display.
| Refractive index range (nD) | 13000 – 1.7000 nD
| Brix (sugar %) range | 0 – 100%
| Refractive index accuracy | ± 0.0002 nD
Optical System Design | Digital Display | Digital Temperature | Temperature Control
Precision dual-prism optical system with hard glass prism ensures sharp boundary definition, high light transmission, and stable, repeatable refractive index measurements. | A clear, backlit digital display provides easy viewing of refractive index, Brix values, and temperature, ensuring quick, accurate readings and improved operator convenience in laboratory environments. | Built-in temperature sensor displays real-time sample temperature, improving measurement reliability and ensuring accurate refractive index and Brix determination. | Optional external thermostatic bath enables precise temperature control of the prism, ensuring stable, accurate refractive index and Brix measurements across varying conditions.
Specifications - ARM-2S
Refractive Index nD Range | 1.3000~ 1.7000nD
Measurement range dissolved solids Brix | 0-100%
Refractive Index nD Accuracy | ±0.0002nD
Accuracy Dissolved Solids Brix | 0.1%
Temperature Resolution | 0.1°C
Temperature Range (Min 0.1°C) | 0.-50°C
Correcting Brix Versus Temperature | 15-45°C
Operating Temperature Range | Ambient to 35°C
Display | LCD Bcklit
Interface | RS232
Power Supply | AC 220V / 50Hz
Dimension (W x D X H)
Packaging dimensions (W x D X H) | 370 x 240 x 470 mm
Net weight | 4 kg
Gross weight | 5 kg
Automatic Digital Abbe Refractometer ARM-Z
ARM-Z automatic refractometer delivers fully automated refractive index and Brix measurement with touchscreen control, fast results, temperature compensation, data storage, and reliable accuracy for liquid, viscous, and solid samples analysis.
| Refractive index range (nD) | 13000 – 1.7000 nD
| Brix (sugar %) range | 0 – 100%
| Refractive index accuracy | ± 0.0002 nD
Optical System Design | Fast Measurement Time | Touchscreen Display | Data Storage & Memory
Precision dual-prism optical system with hard glass prism ensures sharp boundary definition, high light transmission, and stable, repeatable refractive index measurements. | Delivers stable measurement results within seconds, improving laboratory efficiency and supporting high-throughput quality control and production monitoring. | Large color touchscreen enables intuitive operation, quick parameter setup, sample identification entry, and clear visualization of measurement results and temperature. | Built-in memory stores up to 1000 results, enabling traceability, historical comparison, and efficient quality documentation without external devices.
Specifications - ARM-Z
Refractive Index nD Range | 1.3000~ 1.7000nD
Measurement range dissolved solids Brix | 0-100%
Measuring Accuracy nD | ±0.0002nD
Measuring Accuracy Brix | 0.1%
Measuring Accuracy Resolution nD | ± 0.00001 nD
Measuring Accuracy Resolution Brix | ± 0.01% Brix
Correcting Brix Versus Temperature | 15-45°C
Operating Temperature Range | Ambient to 50°C
Display | 7" Touchscreen
Interface | RS232/USB
Data Storage | 1000 Data
Power Supply | AC 220V / 50Hz
Dimension (W x D X H)
Packaging dimensions (W x D X H) | 365 x 245 x 445mm
Net weight | 8 kg
Gross weight | 10 kg
Analog Polarimeter APM-4
APM-4 Series Analog Polarimeter is an analog instrument for accurate optical rotation measurement, using half-shade principle, sodium light source, precise vernier design, ideal for education, training, food, sugar, chemical laboratories.
| Range | ±180°
| Scale Value | 1°
| Light Source | Sodium Lamp
Analog Measurement System | High Measurement Accuracy | Sodium Vapor Light Source | Clear Optical Viewing System
Traditional manual polarimeter designed for accurate optical rotation measurement using the proven half-shade principle | Delivers precise and repeatable optical rotation results, ideal for routine laboratory and educational applications. | Equipped with a stable sodium lamp ensuring consistent wavelength output for reliable measurements. | High-quality optics provide sharp field contrast for accurate end-point detection.
Specifications - APM-4
Range | ±180°
Scale Value | 1°
Vernier | 0.05°
Magnifier | 3X
Light Source | Sodium Lamp
Wavelength | 589 nm
Tube Length | 100/200 mm
Power Supply | AC 220V / 50Hz
Dimension (W x D X H)
Packaging dimensions (W x D X H) | 135 x 500 × 330 mm
Net weight | 3 kg
Gross weight | 5 kg
Automatic Polarimeter APM-2
Designed for high-precision optical rotation analysis, the APM Series Automatic Digital Polarimeter from Aczet delivers fast, accurate, and fully automated measurements for industrial and research laboratories.
| Range | -45°+45°
| Resolution | 0.001°
| Display | 8" Color Datamatrix Touchscreen
High Accuracy & Resolution | Touchscreen Color Display | Temperature-Controlled Sample Chamber APM-3 | Wide Sample Compatibility
Advanced digital polarimeter based on the optical null principle for accurate and repeatable results. Measures optical rotation, specific rotation, sugar (°Z), and concentration with resolutions up to 0.001° and excellent repeatability. | Large 8" color touchscreen provides clear data visualisation, easy navigation, and intuitive operation. Stores up to 5000 data sets with date, time, measuring mode, tube length, and temperature records | Built-in Peltier temperature control maintains sample stability for consistent and reliable results. | Performs up to 6 automatic repeat measurements and calculates the average value for improved accuracy. Capable of measuring dark-colored samples with high transmittance sensitivity.
| Specifications - APM 2 | Specifications - APM 3
No. of Measuring Mode (4) | Optical Rotation, Specific Rotation, Sugar, Concentration | Optical Rotation, Specific Rotation, Sugar, Concentration
Rotation Measuring Range | -45°+45° | -90°+90°
Angle Resolution | 0.001° | 0.001°
Angle Resolution Accuracy | ±0.01° | ±0.004°
Measuring Mode Sugar | -120° +120°Z | -259° +259°Z
Sugar Resolution | ±0.01°Z | ±0.001°Z
Sugar Resolution Accuracy | ±0.02°Z | ±0.004°Z
Measuring Range Concentration | 0-99.9 g / ml | 0-99.9 g / ml
Concentration unit | g / 100 ml | g / 100 ml
Concentration Resolution Accuracy | ±0.5 g / 100 ml | ±0.5 g / 100 ml
Temperature Control | - | Peltier (Built in)
Temperature Range | - | 10oC~50oC
Temperature Accuracy | - | ±0.1oC & ±0.2oC
Constant Temp. Test Tube | - | 100 mm
Memory | - | 5000 data sets
Display | 8" Color Datamatrix Touch Screen | 8" Color Datamatrix Touch Screen
Sample Transmittance | > 1 % | > 1 %
Light Source | LED (2.4V, 1W, 5000 hours) | LED (2.4V, 1W, 5000 hours)
Operation Wavelength | 589.3nm | 589.44nm
Length of Test Tube | 100 mm & 200 mm | 100 mm & 200 mm
Interface | USB & RS 232 | USB & RS 232
Power Supply | AC 220V / 50Hz | AC 220V / 50Hz
Dimension (W x D X H)
Packaging dimensions (W x D X H)
Net weight
Gross weight
Visual Colorimeter- AT 2
The Visual Colorimeter AT 2 Series from Aczet is a reliable and easy-to-operate instrument designed for precise visual color matching based on internationally accepted color scale units.
| Measurement Range | Red R 0.1-R79.9 Yellow Y 0.1-Y79.9 Blue B 0.1-B49.9 Neutral N 0.1-N3.9
| Minimum Readout | 0.1
| Whiteboard Reectivity | 1.9X
Analog Measurement System | High Measurement Accuracy | Sodium Vapor Light Source | Clear Optical Viewing System
Traditional manual polarimeter designed for accurate optical rotation measurement using the proven half-shade principle | Delivers precise and repeatable optical rotation results, ideal for routine laboratory and educational applications. | Equipped with a stable sodium lamp ensuring consistent wavelength output for reliable measurements. | High-quality optics provide sharp field contrast for accurate end-point detection.
Specifications - AT 2
Measurement Range | Red R 0.1-R79.9 Yellow Y 0.1-Y79.9 Blue B 0.1-B49.9 Neutral N 0.1-N3.9
Minimum Readout | 0.1
Whiteboard Relectivity | 1.9X
Magnifying glass power | >80%
Dimension of color plate | 10 mm x 20 mm x 40 mm 25.4 mm x 20 mm x 40 mm 133.4 mm x 20 mm x 40 mm
Power Supply | AC 220V / 50Hz
Dimension (W x D X H)
Packaging dimensions (W x D X H) | 530 x 440 x 330 mm
Net weight | 6 kg
Gross weight | 8 kg
Milk Analyser - Milk Pro+
The Milk Pro+ Milk Analyser from Aczet is a high-speed ultrasonic milk testing instrument designed for accurate and reliable analysis of milk composition. Compact, robust, and chemical-free, it is ideal for dairies, milk collection centers, laboratories, and quality control applications.
| Sample Volume | 20 ml.
| Sample Speed | 60/40/30sec.
| Sample Temperature | 1˚ C – ˚42 C
Ultrasonic Measurement Technology | Rapid Analysis Time | Wide Parameter Measurement | Built-in Data Storage
Advanced double ultrasonic sensor ensures fast, precise, and repeatable milk analysis without the use of chemicals. Double Measuring Sensor Design Enhances measurement stability and accuracy during continuous operation. | Delivers complete milk composition results in 18 seconds at 28°C and 15 seconds at 38°C, improving sample throughput. | Accurately measures FAT, SNF, density, protein, lactose, salts, added water, freezing point, conductivity, and temperature. Low Sample 20 ml Volume Requirement | Memory capacity for 500 measurements with date and time recording. RS232 interfaces for seamless data transfer and system integration.
Specifications - Milk Pro+
Sample Volume | 20 ml.
Sample Speed | 60/40/30sec.
Sample Temperature | 1˚ C – ˚42 C
Humidity | 0 – 80℅ RH
Interface | RS232
Interface | USB
Power Supply | AC 95 – 250 V DC 12 V
Dimension (W x D X H) | 110X280X195 mm
Packaging dimensions (W x D X H)
Net weight | 2 kg
Gross weight | 3 kg
Digital Rotational Viscosity Meter - VISCO-L
The VISCO Series Digital Rotational Viscosity Meter from Aczet is a reliable and cost-effective solution for fast and accurate viscosity measurement of Newtonian and non-Newtonian fluids. Designed for flexibility, ease of operation, and consistent performance, the VISCO Series is ideal for quality control, research, and industrial applications.
| Sample Volume | 20 ml.
| Sample Speed | 60/40/30sec.
| Sample Temperature | 1˚ C – ˚42 C
Digital Rotational Viscometer Technology | High Measurement Accuracy & Repeatability | Wide Viscosity Measurement Range | Multiple Speed Options
Direct digital reading system enables precise measurement of absolute and apparent viscosity. | Offers ±1% full-scale accuracy with repeatability up to 0.2%, ensuring reliable test results. | Capable of measuring from 20 to 6,000,000 mPa·s, depending on model and spindle selection. | Wide speed range with multiple selectable speeds provides flexibility for various sample types.
| Specifications - VISCO-L | Specifications - VISCO-R | Specifications - VISCO-H
Measuring Range mpa.s | 20-2,000,000 | 20-2,000,000 | 20-6,000,000
Speed | 0.3 to 60 | 0.3 to 100 | 0.3 to 200
No. of Speed | 8 | 18 | 54
No. of Standard Spindle | 4 | 4 | 4
Display | LCD Backlit | LCD Backlit | LCD Backlit
Temperature Monitoring | Yes | Yes | Yes
Measurement Accuracy | ±1% (FS) | ±1% (FS) | ±1% (FS)
Repeatability | ±0.5% (FS) | ±0.5% (FS) | ±0.5% (FS)
Interface | RS232 | RS232 | RS232
Power Supply | AC 220V / 50Hz | AC 220V / 50Hz | AC 220V / 50Hz
Dimension (W x D X H)
Packaging dimensions (W x D X H) | 400 x 350 x 360mm | 400 x 350 x 360mm | 400 x 350 x 360mm
Net weight | 8 kg | 8 kg | 8 kg
Gross weight | 10 kg | 10 kg | 10 kg
Portable Digital Tablet Hardness - ATH 500
The ATH 500 Series Portable Digital Tablet Hardness Tester is an advanced, compact instrument designed for fast, accurate measurement of tablet breaking strength — ideal for pharmaceutical labs, quality control departments, and production environments. Tablet hardness testing is essential to ensure the mechanical strength, uniformity, and quality of tablets during manufacturing, storage, and handling.
| Hardness Range | 5 - 500N
| Hardness Accuracy | 0.05% FS
| Display | LCD Backlit
High-Precision Load Cell Measurement | Digital Display with Multiple Units | Programmable Tolerance Limits | Average & Statistical Calculation Functions
Measures tablet hardness up to 500 N with excellent accuracy and resolution. | Clear LCD screen shows results in Newton (N), Kilogram Force (kgf), and Pound Force (lbf) for versatile reporting. | Set upper (T1) and lower (T2) hardness thresholds to flag out-of-range results during testing. | Stores multiple test values and automatically computes averages and peak values.
Specifications - ATH 500
Hardness Range | 5 - 500N
Hardness Accuracy | 0.05% FS
Measurable Units | N, Kgf and lbf
Display | LCD Backlit
Peak Force Measurement | Yes
Acrylic Safety Glass Cover | Yes
Stored Results | Upto 10 (also calculate the average value of test results)
Zero point adjustment for Calibration | Yes
Auto Unit Conversion | Yes
Programmable T1, T2 tolerance | Yes
Tablet Size | 2 mm upto 28 mm
Interface | RS232
Battery | 1200 mAH, 8.4V, Ni-Hi Rechargeable Battery
Power Supply | AC 220V / 50Hz
Dimension (W x D X H) | 355 x 60 x 105 mm
Packaging dimensions (W x D X H)
Net weight | 2 kg
Gross weight | 3 kg
Ultrasonic Probe Sonicator - APU 500
The APU 500 Ultrasonic Probe Sonicator is a high-performance ultrasonic processor designed for efficient sample preparation, cell disruption, homogenization, emulsification, and dispersion in laboratory and industrial applications.
| Frequency (Khz) | 20-25KHz - Automatic Tracking
| Power (W) adjustable | 10-500W
| Probe (mm) | Ø6
High-Power Ultrasonic Processing | Probe-Based Direct Sonication | Adjustable Power & Amplitude Control | Digital Control Interface
Generates intense ultrasonic energy for rapid and effective sample treatment. High-Quality Titanium / Stainless Steel Probe Ensures excellent acoustic transmission, chemical resistance, and long service life. | Delivers ultrasonic energy directly into the sample for maximum efficiency and faster results compared to bath sonicators. | Allows precise control of sonication intensity to suit different sample types and volumes. | Easy-to-use digital panel for setting time, power, and pulse parameters with real-time monitoring.
Specifications - APU 500
Frequency (Khz) | 20-25KHz - Automatic Tracking
Power (W) adjustable | 10-500W
Probe (mm) | Ø6
Optional Probe (mm) | Ø2, Ø3, Ø10, Ø12
Process Capacity (ml) | 0.2-500
Store Data | 20 sets
Temperature Range | Ambient to 99oC
Ultrasonic Setting Time | 0.1-99.9 sec.
Alarm Function | Temp, time, overload, no-load,
Duty Ratio | 0.1-99.9%
Display | 4.3" Touchscreen
Power Supply | AC 220V / 50Hz
Dimension (W x D X H)
Packaging dimensions (W x D X H) | 600 x 300 x 500mm
Net weight | 10 kg
Gross weight | 12 kg
Ultrasonic Cleaner -CUB 2.0
The CUB Series Ultrasonic Cleaner is a powerful and reliable cleaning system designed to remove contaminants from complex, delicate, and hard-to-reach surfaces with exceptional efficiency. Using advanced ultrasonic cavitation technology, this cleaner delivers deep, uniform, and damage-free cleaning for laboratory, industrial, medical, and precision applications.
| Frequency (Khz) | 20-25KHz - Automatic Tracking
| Power (W) adjustable | 10-500W
| Probe (mm) | Ø6
Advanced Ultrasonic Cleaning Technology | Efficient, Uniform & Gentle Cleaning | User-Friendly & Robust Design | Degassing Function
High frequency ultrasonic waves create cavitation bubbles that remove grease, residues, and contaminants from tiny crevices, delivering consistent, damage-free cleaning. | Uniform ultrasonic energy ensures consistent, non-abrasive cleaning of delicate and precision components, while remaining powerful enough to remove heavy contamination effectively. | Designed for ease of use, intuitive controls with LED indicators, allowing precise time and temperature setting. The stainless steel tank and housing ensure durability, corrosion resistance, and reliable continuous laboratory and industrial operation. | Effectively removes dissolved gases and trapped air bubbles from liquids using ultrasonic cavitation, improving cleaning efficiency, solution stability, and accuracy for laboratory and industrial applications.
| Specifications - CUB 2.0 | Specifications - CUB 5 | Specifications - CUB 10 | Specifications - CUB 15 | Specifications - CUB 22 | Specifications - CUB 30 | Specifications - CUB 60 | Specifications - CUB 100
Ltr. Capacity | 1.8 Ltr. | 4.5 Ltr. | 9.8 Ltr. | 13.8 Ltr. | 20 Ltr. | 27 Ltr. | 53 Ltr. | 99 Ltr.
Frequency | 40 Khz | 40 Khz | 40 Khz | 40 Khz | 40 Khz | 40 Khz | 40 Khz | 40 Khz
Adjustable Time | 1 to 99 minutes | 1 to 99 minutes | 1 to 99 minutes | 1 to 99 minutes | 1 to 99 minutes | 1 to 99 minutes | 1 to 99 minutes | 1 to 99 minutes
Adjustable Temperature | 0 to 80oC | 0 to 80oC | 0 to 80oC | 0 to 80oC | 0 to 80oC | 0 to 80oC | 0 to 80oC | 0 to 80oC
Display | LED | LED | LED | LED | LED | LED | LED | LED
Ultrasonic Power | 50 W | 150 W | 200 W | 300 W | 400 W | 600 W | 900 W | 1200 W
Heating Power | 100 W | 300 W | 300 W | 400 W | 500 W | 1000 W | 2000 W | 3000 W
Tank Dimension (LxWxH) mm | 150 x 140 x 100 | 300 x 155 x 100 | 300 x 240 x 150 | 330 x 300 x 150 | 500 x 300 x 150 | 500 x 300 x 200 | 500 x 350 x 300 | 550 x 450 x 400
Power Supply | AC 220V / 50Hz | AC 220V / 50Hz | AC 220V / 50Hz | AC 220V / 50Hz | AC 220V / 50Hz | AC 220V / 50Hz | AC 220V / 50Hz | AC 220V / 50Hz
Dimension (W x D X H) mm | 190 x 170 x 220 | 312 x 162 x 240 | 330 x 270 x 310 | 360 x 330 x 330 | 550 x 340 x 330 | 500 x 330 x 360 | 760 x 530 x 730 | 820 x 630 x 850
Packaging dimensions (W x D X H) mm |  | 430 x 270 x 320 mm | 430 x 270 x 320 mm | 430 x 270 x 320 mm | 430 x 270 x 320 mm | 430 x 270 x 320 mm | 430 x 270 x 320 mm | 430 x 270 x 320 mm
Net weight | 3.2 kg | 5 kg | 10 kg | 18 kg | 18 kg | 20 kg | 45 kg | 50 kg
Gross weight | 4.2 kg | 6 kg | 11 kg | 21 kg | 21 kg | 31 kg | 50 kg | 62 kg
Ice Flake Maker -AMS 30
The AMS Series Ice Flake Maker is a reliable and efficient solution for producing hygienic ice flakes for laboratory, medical, pharmaceutical, food, and industrial applications. Designed for continuous operation, it delivers uniform ice flakes with high cooling efficiency and low energy consumption.
| Capacity | 30 kg /24h
| Ice storage | 10 kg
| Water consumption (L/H) | 0.8
Fast & Continuous Ice Production | Energy-Efficient Operation | Stainless Steel Construction | Automatic Operation
Produces dry, soft ice flakes continuously to meet demanding cooling requirements. Uniform Ice Flake Structure Thin, flat ice flakes provide maximum surface contact for rapid and efficient cooling. | Optimized refrigeration system delivers high output with low power consumption. | Corrosion-resistant stainless steel body ensures durability, hygiene, and easy cleaning. Compact & Space-Saving Design Ideal for laboratories, hospitals, and industrial facilities with limited space. | Fully automatic ice-making cycle with minimal operator intervention.
| Specifications -AMS 30 | Specifications -AMS 50 | Specifications -AMS 100 | Specifications -AMS 200
Capacity | 30 kg /24h | 50 kg /24h | 100 kg /24h | 200 kg /24h
Ice storage | 10 kg | 15 kg | 25 kg | 55 kg
Water consumption (L/H) | 0.8 | 2 | 4.1 | 8.3
Power (W) | 180 | 280 | 420 | 805
Power Supply | AC 220V / 50Hz | AC 220V / 50Hz | AC 220V / 50Hz | AC 220V / 50Hz
Dimension (W x D X H) mm | 300 x 493 x 547 | 380 x 543 x 722 | 548 x 611 x 883 | 680 x 641 x 1102
Packaging dimensions (W x D X H) mm
Net weight | 32 kg | 43 kg | 60kg | 88 kg
Gross weight | 36 kg | 45 kg | 64 kg | 91 kg
Coating Thickness Analyzer
A versatile XRF solution delivering elemental and coating analysis using micro-focus X-ray, proportional counter detection, precise CCD spot targeting, and an extensive calibration library with certified safety.
| Detector | Gas Filled Prop Counter
| Measuring Direction | Top to Bottom
| Measurement Time & Accuracy | 10 to 60 sec & ±0.3%
User Friendly Software | Safe | Coating Thickness | Bath Analysis | Material Analysis | Micro focus X-ray | Measuring Directions | Diverse Calibration Sample | Digital Pulse Processor [DPP] | Silicon Drift Detector [FSDD] | Silicon PIN Technology
Most comprehensive software for material analysis including Coating thickness measurement as well as bath analysis. Convenient calibration function High resolution CCD camera with optical zoom help positioning the spot for measurement. | The shutter is a safety device. It prevents primary radiation from entering the measuring chamber when measurement is in progress, thus safeguarding operator from risk of X-Ray exposures. X-Ray Auto cut off if accidentally door open. AERB approved | From simple coating thickness measurements in the electroplating industry, such as zinc on iron, multi-coating applications, sophisticated precious metal coating | Bath analysis in the electroplating industry. Measurement of e.g. thin gold, palladium in jewellery industry. | Versatile Material Analysis software x-Master based on both Empirical and Fundamental method help yield accurate Measurement as well as identify and measure any metal standard free | The X-ray generates the primary X-ray radiation. Advanced micro-focus tubes are used to concentrate a large part of the primary radiation onto a very small measuring spot. | User-friendly precision: position the spot anywhere you see fit. Ideal for uneven surface jewellery and rings with inner & outer surface testing. | More than 3 decade of field experience help us widening calibration library made it possible to perform diverse calibration for each unit leave our works. | DPP generates very high pulse and count rates amplifying the signal detected by the detector. Working together with the detector, the DPP plays a crucial role in maintaining very high stability and energy resolution. | Gold Testing machine installed with AMPTEK by USA Suitable for simple measurement. Ideal for measuring complex shape samples and different measuring distance. | Gold Testing machine installed with original AMPTEK by USA. Suitable for simple measurement. Ideal for measuring complex shape samples and different measuring distance.
| Specifications - Compact Eco, prop counter | Specifications - Compact Eco, High Resolution Si-PIN | Specifications - Compact Eco, Fast Silicon Drift Detector(FSDD)
Detector | Gas Filled Proportional Counter | High Resolution Si-PIN | Fast Silicon Drift Detector(FSDD)
Measuring Direction | Top to Bottom | Top to Bottom | Top to Bottom
Applied Application | Gold, Silver, Platinium Analysis in Jewellery Alloys | Coating Thickness, Bath Analysis | Coating Thickness, Bath Analysis
Detector Range | Au, Ag, Cu, Zn, Ni, Pd, Cd, Pt, Fe, Co, Mn,Pb, Cr & more | Au, Ag, Cu, Zn, Ni, Pd, Cd, Rh, Ru, Pt, Fe, Co, In. Os, Ir, Mn,Pb,Ti, Hf, Cr, V, W, Ta & more | Au, Ag, Cu, Zn, Ni, Pd, Cd, Rh, Ru, Pt, Fe, Co, In. Os, Ir, Mn,Pb,Ti, Hf, Cr, V, W, Ta & more
Measurement Time | 10 to 60 sec. | 10 to 60 sec. | 10 to 60 sec.
Focus | Auto focus | Auto focus | Auto focus
X-ray Tube | Power 50 Watts Oil cooling X-ray Tube | Power 50 Watts Oil cooling X-ray Tube | Power 50 Watts Oil cooling X-ray Tube
High Voltage | 50kV (1 mA), Software Controlled Optimized | 50kV (1 mA), Software Controlled Optimized | 50kV (1 mA), Software Controlled Optimized
Collimator | 0.3 mm | 0.3 mm / 0.5 mm (Optional four positions collimator charger) | 0.3 mm / 0.5 mm (Optional four positions collimator charger)
Sample Stage | Manual scissors-type Z-stage | Manual scissors-type Z-stage | Manual scissors-type Z-stage
Power Supply | 230 VAC, 50/60Hz, 120W / 100W | 230 VAC, 50/60Hz, 120W / 100W | 230 VAC, 50/60Hz, 120W / 100W
Inside Chamber | 330 x 200 x 170 mm (W x D x H) | 330 x 200 x 170 mm (W x D x H) | 330 x 200 x 170 mm (W x D x H)
Dimension | 630 x 430 x 420 mm (W x D x H) | 630 x 430 x 420 mm (W x D x H) | 630 x 430 x 420 mm (W x D x H)
Operating System | Windows 10 | Windows 10 | Windows 10
Sample View | Crosshair-guided zoomable video & image capture for precise targeting | Crosshair-guided zoomable video & image capture for precise targeting | Crosshair-guided zoomable video & image capture for precise targeting
System | External PC set | Standard PC Set | Standard PC Set
Operating Conditions | 15°–25° with <90% humidity | 15°–25°C with <90% humidity | 15°–25°C with <90% humidity
Integrations | Compatible with electronic balances | Compatible with electronic balances | Compatible with electronic balances
Approvals | ISO Approved, AERB | ISO Approved, AERB | ISO Approved, AERB
Accuracy | ±0.3% | ±0.2% | ±0 1%
Weight | 27 kg | 40 kg | 40 kg
Hand Held XRF Nanoris
A portable, non-destructive analyzer for rapid elemental testing of metals, alloys, precious metals and coatings. It identifies alloy grades in seconds and checks plating thickness without damaging samples. Lightweight for real-time on-site PMI and scrap sorting. Imported detector ensures low-noise multi-element performance, with PC/mobile remote-control data sync.
| Detector | Gas Filled Prop Counter
| Measuring Direction | (RSD) ≤ 1%
| Measurement Element Range | Mg (12) - U (92)
NDT | Computer Software | Portability | High quality imported detectors | Automobile Industry | XRF Precious Metal Analyzer | Characteristics of aluminum alloy | Alloy Field waste recyling (Scrap)
Quickly and accurately analyze the plating thickness and sample element content of the plating sample without damaging the test object. | Remote control system, C # data synchronization control system based on Windows system, Both computer and mobile software can operate instruments and synchronize data through software. . | The instrument has a smaller volume and lighter weight, making it convenient for real-time on-site testing and can easily handle samples of various shapes | The instrument adopts imported detectors suitable for analyzing multi-element coatings, with low noise and ?exible response to coating structure. | In the automotive industry, the Tektronix spectrometer is used in the field of ternary catalysis to evaluate whether automobile exhaust emissions meet standards. It can detect harmful gases and particulate matter in car exhaust, as well as their concentration and composition, providing data support for formulating car emission standards | The detection of the content of dozens of elements such as gold, silver, and platinum group metals (ruthenium, rhodium, palladium, osmium, iridium, platinum) from K to U shows that the content can be both% and K gold | Aluminum alloy is a lightweight metal material composed of aluminum as the base with a certain amount of other alloying elements added. In addition to the general properties of aluminum, aluminum alloys exhibit speci?c characteristics due to the types and quantities of alloying elements added. The density of aluminum alloys ranges from 2.63 to 2.85g/cm3, with high strength (b ranging from 110 to 650MPa), comparable strength to high-alloy steel, higher stiffness than steel, good casting performance, plastic processing performance, excellent electrical and thermal conductivity, corrosion resistance, and weldability. | For the scrap metal recycling industry, an accurate and portable handheld spectrometer can help identify alloys and analyze material chemical composition, while the Aczet handheld spectrometer can provide fast and accurate elemental analysis for scrap metal recycling. Relying on Aczet digital and intelligent industry solutions, Aczet’s handheld spectral analyzer can measure element content as low as several parts per million.
Specifications - Hand Held XRF Nanoris Si-PIN
Detector | Si - Pin
Element Range | Ti (22) - U (92)
Measurement Accuracy | (RSD) < 1%
Data Transmission | USB, E-mail, Wi-Fi, Bluetooth
Network Connectivity | Wi-Fi, Bluetooth
Control System | Customized Industrial - Grade Control System : CPU : 2.4 Ghz, RAM : 2GB, Internal Storage : 4GB (Expandable upto 32GB)
Application mode | Precious Alloy, RoHs, Alloys
Camera | High Definition Camera
Detector Window Protection | Shield Window
Analysis Methods | Empirical coefficient method, Intelligent FP Method
Excitation Source | 50kV / 200mA, X-ray fluorescence / 500mA (Optional)
Detector Cooling Method | Two-stage thermoelectric cooling
Safety | Proximity Sensor for Radiation Safety / Radiation Protection (Optional)
Battery Life | 6 - 7 Hours.
Dimensions | 254 mm (L) x 91 mm (W) x 319 mm (H)
Material | Alloy, polycarbonate
Weight | 1.75 kg (Including battery)
Laser Spot Welding Machine Stout 225
Stout is a high-performance laser spot welding system designed for precision joining of fine jewellery and micro components. It combines high pulse energy, stable beam control, and real-time visual monitoring to deliver consistent, high-quality welds with minimal thermal impact.
| Laser Source | Nd:Ce:YAG
| Wavelength | 1064 nm
| Max Pulse Energy | 225 J
Spacious Chamber Access | Innovative Smoother Spot | Double Cooling System | Precision Targeting Microscope | Capacitive Touch Display | 100 Program Memory | Sanitization Safety
Stout features a large welding chamber with a wide 180° opening for easy handling of larger components. It includes joystick control for dynamic parameter adjustment, integrated LED lighting for clear visibility, and gas/argon provision, ensuring flexible, precise, and comfortable operation for intricate jewellery work. | The advanced optical design ensures an even and stable energy spot, resulting in cleaner weld surfaces and better distribution of filler material. This helps achieve smoother finishing and stronger, more refined weld joints. | Advanced dual cooling design efficiently manages heat during welding, ensuring stable performance and continuous operation across varying working conditions. | Leica 10X stereo microscope provides advanced optics with clear and sharp visualization, enabling accurate micro-welding with reduced eye strain and improved productivity. | Equipped with a vibrant 8-inch colour capacitive touch screen, the system offers a smooth and intuitive user experience. The highly responsive display makes navigation easy and adjustments quick, ensuring comfortable and seamless operation. | Allows storage of up to 100 welding programs with customizable parameters, eliminating the need to manually remember or adjust settings for different metals. | Ensures a clean and safe working environment with an integrated dust filter that removes fumes, dust, and odors during operation.
Specifications | Stout 225
Laser Source | Nd:Ce:YAG
Wavelength | 1064 nm
Maximum Laser Energy | 225 J
Peak Power | 9 kW
Average Power | 100 W
Pulse Duration | 0.1 – 25 ms
Frequency of Repetition | 0.5 to 30 Hz
Spot Setting | Motorized
Spot Diameter | 0.2 to 2 mm (0.1 mm optional)
Microscope | 10× Leica Stereo Microscope
Display Theme | Light / Dark
Wave Program | 6, Preset Pulse Shapes
Program Memory | 100
Internal Control | Joystick
External Control / Display | 8" Colour TFT Touch Screen
Argon / Air Nozzle | Yes, Double Nozzle
Compressed Air Nozzle | Yes
Pre/Post Mode for Valve | Yes
Cooling System | Double Cooling System 4L
Provision for External Chiller | Yes
Camera | In-screen Camera
HDMI (Output) | Yes
Input Power | 230V AC, 50/60 Hz, Single Phase (110V on request)
Dimension (W × D × H) | 850 × 500 × 590 mm
Weight | 50 kg
Laser Marking Machine - CFM 20
Laser marking machines deliver high-quality cutting and engraving for precise marking of logos, text, and graphics on metal surfaces in minimal time. With versatile power configurations, they support both marking and foil cutting, ensuring precision, flexibility, and higher production efficiency.
| Laser Source | Ytterbium Pulsed Fiber Laser
| Laser Power | 20 W
| Wavelength | 1064 nm
Simple and Versatile Software | Motorized Z-Axis | Precision and Accuracy | Quality Components | Air-Cooled Fiber Laser
Intuitive and easy-to-use software allows smooth operation while supporting engraving of logos, text, images, and vector files on various metals, with compatibility for formats like SVG, DXF, BMP, JPG, and more. | The motorized Z-axis helps you quickly and accurately adjust focus, making the overall operation easier and more efficient. | Delivers high precision and accuracy with a marking area of up to 160 × 160 mm and resolution up to 16 μm, making it perfect for fine details, intricate designs, and small characters. | Built with high-quality components and a well-designed system to deliver clean and consistent engraving, even for complex and detailed designs. | Built with a pulse fiber laser source with no moving parts or mirrors, ensuring maintenance-free operation, high efficiency, and reliable performance without the need for an external chiller.
Model | CFM 20
Types | Marking / Engraver
Laser Source | Ytterbium Pulsed Fiber Laser
Laser Power | 20 W
Wavelength | 1064 nm
Beam Quality M² | < 1.6
Pulse Repetition Rate | 20 Hz-100kHz
Working Area | 100 × 100 mm
Working Life of Laser | >50,000 hours
Marking Speed | < 7000 mm/sec
Laser Frequency | 20 - 100 kHz
Axis | Auto Z Axis
Cooling System | Air Cooling
Minimum Width | 0.3 mm
Minimum Character | 0.3 mm
Repeatability | 0.003 mm
Power Consumption | 800 W
Power Supply | 220V AC / 60Hz / 1P /4A or 110V / 50Hz / 1P / 8A (110V on request)
Dimension (L × W × H) | 635 × 400 × 660 mm
Net Weight (Approx.) | 40 kg
Packing Dimension (L × W × H) | 965 × 610 × 914 mm
Weight | 45 kg

## Product Inspection Solutions
Check Weigher CW 200
A CW Series checkweigher is an automatic dynamic weighing system that accurately measures, classifies, and rejects underweight or overweight products, ensuring 100% online inspection, regulatory compliance, consistent quality, and customer protection across pharmaceutical, food, chemical, and cosmetic industries.
| Capacity | 5 ~ 200 g
| Readability | ± 0.2 g
| Max Speed (pcs / min) | 150
High-Precision Weighing | Fast & Efficient Operation | Robust Housing | Smart Display & Indicators | Flexible Reject Options | Data Management & Connectivity
CW Series ensures extremely accurate product weight measurement, with scale divisions from 0.01g to 10g depending on model, providing reliable quality control and reducing errors in industrial production lines. | Max speed ranges up to 600 pcs/min, allowing rapid weighing of products on production lines. Adjustable belt speed ensures efficient workflow without compromising accuracy, enhancing overall productivity. | Constructed from corrosion-resistant stainless steel (SS304) or mild steel, CW Series withstands industrial environments. The design enables easy cleaning, durability, and hygiene, supporting long-term use and safe product handling. | Equipped with a touch screen display, indicator lamps, and buzzer for high/low, acceptance, and rejection signals. Provides clear visual feedback to operators, streamlining production and preventing human error. | CW Series supports multiple reject systems—Push, Air Jet, Arm, or Drop Belt—allowing seamless adaptation to different products and production lines for precise, automated sorting of underweight or overweight items. | Capable of storing 100 batches with 2000 reports per batch. RS232, optional USB, and Ethernet connectivity enable direct data transfer to Windows systems, MS Access, or SQL Server for traceability and reporting.
| Specifications - CW 200 | Specifications - CW 500 | Specifications - CW 1K | Specifications - CW 3K
Maximum Capacity | 0.5 ~ 200g | 50~500g | 600~1200g | 1.2~3kg
Scale Division | 0.01g | 0.05g | 0.1g | 0.5g
Precision ³± | 0.2g | 0.3g | 0.5g | 1g
Protection class | IP54 | IP54 | IP54 | IP54
Max Speed (pcs / min) ± | 180 | 180 | 150 | 120
Belt Width (mm) | 100 | 150 | 230 | 300
Height (mm) | 750 x ± 50 | 750 x ± 50 | 750 x ± 50 | 750 x ± 50
Interface | Rs232C, Optional USB, Ethernet Windows Direct communication/MS access & SQL server | Rs232C, Optional USB, Ethernet | Rs232C, Optional USB, Ethernet | Rs232C, Optional USB, Ethernet
Power Supply | AC110 / 220V, ± 10% 50Hz | AC110 / 220V, ± 10% 50Hz | AC110 / 220V, ± 10% 50Hz | AC110 / 220V, ± 10% 50Hz
Working Condition Temp | 5-40°C, RH<95% | 5~40°C, RH<95% | 5~40°C, RH<95% | 5~40°C, RH<95%
Balance Dimension (W x D X H)
Packaging dimensions (W x D X H)
Net weight
Gross weight
Metal detector MDC 200
The MDC Series metal detector with integrated conveyor ensures reliable quality assurance for packaged food products. Its robust design enables safe detection of steel, stainless steel, and non-ferrous metal contaminants, including aluminum, with high accuracy and flexibility.
| Tunnel Width (mm) | 300
| Tunnel Height (mm) | 80 / 100 / 120 / 150 / 180
| Max Speed (pcs / min) | 60
Metal Detection | Reject System | Touch Interface | Stainless Construction | Adjustable Tunnel | Data Connectivity
MDC Series accurately detects ferrous, non-ferrous, and stainless steel contaminants in food, pharmaceutical, and industrial products. Adjustable sensitivity ensures high detection performance across diverse production lines. | Integrated rejector options including Push, Air Jet, Arm, and Drop Belt allow automated removal of contaminated products, maintaining product safety and hygiene while minimizing manual intervention on high-speed lines. | User-friendly touchscreen displays high/low, acceptance, and reject signals. Intuitive controls allow operators to monitor processes effectively, adjust sensitivity, and ensure reliable detection with minimal training. | Robust SS304 stainless steel frame with anti-vibration legs provides long-lasting durability and corrosion resistance. The hygienic design ensures easy cleaning, ideal for food, pharmaceutical, and industrial environments. | Height, width, and conveyor speed can be customized to fit various product sizes and production requirements, providing flexibility for different industries while maintaining accurate and efficient metal detection. | Equipped with RS232 interface and memory storage for up to 10 products. Enables integration with production line systems and traceability software for reporting and quality control compliance.
| Specifications - MDC 300 | Specifications - MDC 350 | Specifications - MDC 400 | Specifications - MDC 500 | Specifications - MDC 600
Tunnel Width (mm) | 300 | 350 | 400 | 500 | 600
Sensibility (dia. mm) Ferrous | 0.4 / 0.4 / 0.6 / 0.8 / 0.8 | 0.6 / 0.8 / 0.8 / 0.8 / 1.0 | 0.6 / 0.8 / 1.0 / 1.0 / 1.3 / 1.6 | 1.0 / 1.0 / 1.3 / 1.3 / 1.6 / 1.8 | 1.0 / 1.3 / 1.6 / 1.8 / 2.3
Sensibility (dia. mm) Non-Ferrous | 0.6 / 0.6 / 0.8 / 1.0 / 1.0 | 0.8 / 1.0 / 1.0 / 1.0 / 1.2 | 0.8 / 1.0 / 1.2 / 1.2 / 1.5 / 1.8 | 1.2 / 1.2 / 1.5 / 1.5 / 1.8 / 2.0 | 1.2 / 1.5 / 1.8 / 2.0 / 2.5
Sensibility (dia. mm) S.S 304 | 1.0 / 1.0 / 1.3 / 1.3 / 1.6 | 1.0 / 1.0 / 1.3 / 1.5 / 1.8 | 1.3 / 1.3 / 1.6 / 1.8 / 2.0 / 2.3 | 1.6 / 1.8 / 1.8 / 2.0 / 2.3 / 2.8 | 1.8 / 2.3 / 2.3 / 2.8 / 3.3
Belt width | Standard size :380 mm, Customized on request | Standard size :380 mm, Customized on request | Standard size :380 mm, Customized on request | Standard size :380 mm, Customized on request | Standard size :380 mm, Customized on request
Interface | RS232C | RS232C | RS232C | RS232C | RS232C
Power Supply
Balance Dimension (W x D X H)
Packaging dimensions (W x D X H)
Net weight
Gross weight
Check weigher with Metal detector combo CMD 3K
A single-platform CMD series system integrates metal detection and checkweighing for the food industry, identifying contamination, verifying product weight, and classifying packed products to ensure quality, safety, compliance, and efficient production control.
| Capacity | 1~ 3000 g
| Readability | ± 0.25g
| Max Speed (pcs / min) | 60
Balanced Coil | High Speed | Accurate Weighing | Smart Rejection | Touch Display | Robust Construction
Equipped with a balanced coil and digital weighing sensor, the CMD 3K ensures precise, stable, and repeatable weighing for conductive and non-conductive products with minimal interference, maximizing accuracy on production lines. | Conveyor belt operates at 60 m/min, enabling fast product throughput without compromising weighing accuracy. Ideal for industrial and food production applications, this high-speed design maintains efficiency and productivity consistently. | Offers ±0.5 g accuracy over a 10–3000 g range with 0.1 g minimum scale. Ensures reliable measurements for various product types, improving quality control and reducing product rejection or wastage. | Flexible rejection systems include pusher, air jet, or flapper options. Automatically removes underweight, overweight, or metal-contaminated products, enhancing efficiency, minimizing human error, and maintaining product integrity during production. | Graphical touch-screen interface provides intuitive control, easy parameter adjustments, and real-time monitoring. Simplifies operation, reduces training time, and allows operators to efficiently manage multiple product types with memory storage. | Stainless steel SS304 grade construction with anti-vibration adjustable legs ensures durability, hygiene compliance, and stability in industrial environments, enabling long-term, low-maintenance operation even under continuous production demands.
Specifications - CMD 3K
Maximum Capacity | 10 - 3000 g
Minimum Scale | 0.1 g
Detection Theory | Balanced Coil & Digital Weighing Sensor
Belt Speed | 60 m/min.
Accuracy | ±0.5 g
Reject Systems | Pusher /Air Jet / Flapper
Display | Graphical with Touch Screen
Alarm Method | Buzzer and lamp
Power Supply | AC110V/220V, 50-60Hz
Rated Output | Approx. 300W
Balance Dimension (W x D X H) | 1630 x 800 x 750 ± 20mm
Packaging dimensions (W x D X H)
Net weight
Gross weight
Heavy Pack Checkweigher CW 30K
A Shipper carton checkweigher CW-K series ensures accurate weight verification of packed cartons, detects under / overfills, improves shipment accuracy, reduces errors, and maintains consistent quality control in packaging and logistics operations.
| Capacity | 12~30kg
| Readability | 2g / 5g
| Max Speed (pcs / min) | 50/30
Flexible Design | Weighing sensor | Robust Housing | New Touchscreen as Smartphone | Smart Control
Product moves through heavy-duty conveyor at controlled speed, applying uniform force on load cells for accurate dynamic weighing, enabling real-time classification of large packaged loads in industrial environments. | High-quality load cells deliver excellent accuracy, repeatability, long-term stability, robust construction, and reliable performance under continuous dynamic and industrial operating conditions. | Check weigher systems feature a rigid stainless-steel frame with vibration-resistant construction, hygienic design, smooth surfaces, and modular layout, ensuring stability, accuracy, easy cleaning, and long-term reliability in industrial environments. | The touch screen display offers an intuitive user interface with clear visuals, real-time weight data, easy parameter setting, multilingual support, and fast operator interaction for efficient check weigher system control. | Simple setup boosts productivity with adjustable speed, 100-product memory, 32-bit high-speed CPU, DSP technology, and capacity to store and easily print up to 2000 logs.
| Specifications - CW 30K | Specifications - CW60K
Maximum Capacity | 12-30kg | 30-60kg
Scale Division | 2g/5g | 5g
Precision | ± 5g/10g | ± 10g
Max Speed (pcs / min) | 50/30 | 30
Belt Width (mm) | 500 | 600
Conveyor Belt Length (mm) | 1600 | 1600
Height (mm) | 750 x ± 50 | 750 x ± 50
Working Condition Temp | 5-40°C, RH<95% | 5-40°C, RH<95%
Protection class | IP54 | IP54
Interface | Rs232C, Optional USB, Ethernet Windows Direct communication/MS access & SQL server | Rs232C, Optional USB, Ethernet Windows Direct communication/MS access & SQL server
Power Supply | 230 VAC, 50 Hz, Max 15 W | 230 VAC, 50 Hz, Max 15 W
Balance Dimension (W x D X H)
Packaging dimensions (W x D X H)
Net weight
Gross weight
Baby Scales BS 20 LCD
Accurate baby scale with high-precision sensors, LCD display, tare function, and auto-off. Compact, safe design ensures precise weight measurement for infants. Easy-to-read, reliable, and suitable for home or clinical use.
| Capacity | 20 Kg
| Accuracy | 5g
| Display | LCD Backlite
High-Precision Sensors | Bright LCD Display | Convenient Measurement | Portable & Durable | Certified & Compliant
advanced averaging weight technology delivers extremely sensitive, accurate, and reliable measurements, ensuring infants’ weight is captured precisely every time. | Large, backlit LCD ensures easy readability; Save and Freeze features allow quick, convenient monitoring for caregivers and medical staff | Includes attachable self-adhesive baby length ruler; auto-zero, tare, and auto-hold functions streamline weighing and height assessment simultaneously. | Lightweight, battery-operated design with adjustable leveling feet provides stability, portability, and long-lasting use across hospital, clinic, or field environments. | Fully BIS-certified and ISO-approved, meeting IS:9281 standards, ensuring safety, accuracy, and compliance with legal metrology and healthcare regulations.
Specifications - BS 20
Maximum Capacity | 20 kg
Accuracy | 5 g
Function Keys | Music on Hold
Display | LCD Backlite
Working Condition Temp | 10ºC to 40ºC
Protection class | IP54
Power Supply | 6Vdc / 1amp
Balance Dimension (W x D X H) | 545 x 330 x 145, ±5mm
Packaging dimensions (W x D X H) | 600 x 350 x 150 mm
Net weight | 4 Kg
Gross weight | 5 Kg
Baby Scales BS 20
Lightweight and battery-operated, this infant scale features advanced sensors, backlit display, attachable length ruler, and adjustable leveling feet, delivering fast, precise, and reliable measurements in hospitals and clinics.
| Capacity | 20 Kg
| Accuracy | 5g
| Display | LED Display
High-Precision Sensors | LED Display | Convenient Measurement | Portable & Durable | Certified & Compliant
advanced averaging weight technology delivers extremely sensitive, accurate, and reliable measurements, ensuring infants’ weight is captured precisely every time. | Bright LED screen provides clear, easy-to-read weight measurements; Save and Freeze functions enable quick, convenient monitoring for caregivers and healthcare professionals. | Includes attachable self-adhesive baby length ruler; auto-zero, tare, and auto-hold functions streamline weighing and height assessment simultaneously. | Lightweight, battery-operated design with adjustable leveling feet provides stability, portability, and long-lasting use across hospital, clinic, or field environments. | Fully BIS-certified and ISO-approved, meeting IS:9281 standards, ensuring safety, accuracy, and compliance with legal metrology and healthcare regulations.
Specifications - BS 20
Maximum Capacity | 20 kg
Accuracy | 5 g
Function Keys | Music on Hold
Display | LED Display
Working Condition Temp | 10ºC to 40ºC
Protection class | IP54
Power Supply | 6Vdc / 1amp
Balance Dimension (W x D X H) | 545 x 330 x 145, ±5mm
Packaging dimensions (W x D X H) | 600 x 350 x 150 mm
Net weight | 4 Kg
Gross weight | 5 Kg
Mother And Child Scale HS 180
Designed for precise weighing of both mothers and infants, the HS Series combines safety, reliability, and ease-of-use for hospitals, clinics, and healthcare facilities.
| Capacity | 180 Kg
| Graduation | Graduation
| Display | LCD Backlite
High-Precision Sensors | Tempered-Glass Platform | Convenient Operation | Integrated Monitoring | User-Friendly Display | Data Storage & Connectivity
Advanced strain gauge technology ensures accurate and repeatable weight measurement for both mother and baby | 6 mm tempered-glass pan with stainless steel cover provides durability and easy cleaning. | Step-on technology calibration, auto ON/OFF, low-battery and overload indication for simple use. | Supports input of ID, gender, age, and height with in-built child growth monitoring system. | Large backlit LCD clearly shows mother, baby, and combined weights with alternate display options. | USB port and memory for at least 500 measurements; real-time clock for traceable documentation.
Specifications - HS 180
Maximum Capacity | 180 kg
Graduation | Graduation
Display | LCD Backlite
Working Condition Temp | 10ºC to 40ºC
Power Supply | 2 x 1.5V AAA / 2* 3V CR2032 Lithium Battery(included)
Balance Dimension (W x D X H) | 302 x 302 x 26
Packaging dimensions (W x D X H) | 325 x 325 x 50
Net weight | 2 Kg
Gross weight | 2.5 Kg
Health Scales HS 180G
Precision health scale with strain-gauge sensors, backlit LCD display, auto-zero/tare/hold functions, lightweight design, and BIS/ISO certification for clinical and pediatric use.
| Capacity | 180 Kg
| Accuracy | 100 g
| Display | LCD Backlite
High-Precision Sensors | Bright LCD Display | Convenient Operation | Portable & Durable | Certified & Compliant
Advanced strain-gauge technology ensures accurate and reliable weight measurements for clinical and pediatric applications. | Backlit, easy-to-read screen with auto-hold, save, and freeze features for effortless monitoring and data recording. | Auto-zero, tare, and multiple unit conversion simplify usage, making daily measurements fast and efficient. | Lightweight design with non-slip base, battery operation, and adjustable leveling feet for consistent performance anywhere. | BIS-approved, ISO 9001:2015 and ISO 13485:2016 certified; conforms to IS:9281 standards for safety and accuracy.
Specifications - HS 180
Maximum Capacity | 180 kg
Accuracy | 100 g
Pan Size | 280 x 280 x 23.5
Display | LCD Backlite
Power Supply | 2 x 1.5V AAA / 2* 3V CR2032 Lithium Battery
Balance Dimension (W x D X H) | 302 x 300 x 24
Packaging dimensions (W x D X H)
Net weight | 2 Kg
Gross weight | 2.5 Kg
Height and Weight Scales HS 200
High-precision digital scale for accurate weight and BMI measurement. Features large backlit display, auto-zero, stabilization, and energy-saving functions. Ideal for hospitals, clinics, gyms, and infant weighing.
| Capacity | 200 Kg
| Readability | 100 g
| Display | LCD Backlite/LED
Design & Build | Display | Capacity & Accuracy | Power & Operation | Applications & Certification
Compact, ergonomic scale crafted for hospitals and clinics, built to OIML III standards ensuring precise and reliable healthcare measurements. | Large 3.2” backlit LCD/LED with temperature reading, providing clear, accurate, and easy-to-read measurements for healthcare staff and caregivers. | Measures up to 200 kg / 440 lbs with 100 g precision, featuring auto-zero and stabilization for consistent, reliable results. | Dual power options: batteries or AC 110/220V, with auto power-off and energy-saving functions for efficient, uninterrupted operation in any environment. | Suitable for hospitals, clinics, gyms, telemedicine, and infant weighing; BIS ISI certified, ISO 9001:2015 & ISO 13485:2016 approved for reliability.
Specifications - HS 200
Maximum Capacity | 200 kg
Readability | 100 g
Pan Size | 280 x 280 x 23.5
Display | LCD Backlite/LED
Division | 100 g
Height rod | 210
Height range to be measured | 190
Working Condition Temp | 10ºC to 40ºC
Power Supply | 2 x 1.5V Battery
Balance Dimension (W x D X H)
Packaging dimensions (W x D X H) | 780 x 375 x 210
Net weight | 14 Kg
Gross weight | 15 Kg
Animal Weighing Scales CTG 150A
High-capacity veterinary scale designed for clinics, kennels, and grooming centers. Compact, portable, with MotionLock for accurate readings, multiple weighing modes, and durable stainless steel platform for reliable animal weight measurements.
| Capacity | 150 kg
| Readability | 50 g
| Display | LCD Backlite
Precision Weighing | Durable Platform | MotionLock Technology | User-Friendly Interface | Flexible Power
Accurately measures animal weight in kilograms, pounds, or pieces, ensuring reliable results for veterinary and commercial animal applications. | Removable stainless steel platform resists wear, simplifies cleaning, and supports large animals safely for consistent weight readings. | Maintains accurate measurements even with animal movement, minimizing errors and enhancing precision during clinical or grooming operations. | Large LCD backlit display, intuitive controls, and clear readouts simplify operation for veterinary staff and ensure efficient workflow. | Supports battery or AC power; includes auto sleep/power-down and low battery indication to maximize operational uptime in any location.
| Specifications - CTG 150A | Specifications - CTG 300A
Maximum Capacity | 150 kg | 300 kg
Readability | 50 g | 50 g
Pan Size | 600 x 500 mm | 900 x 600 mm
Display | LCD Backlite | LCD Backlite
Working Condition Temp
Power Supply
Balance Dimension (W x D X H)
Packaging dimensions (W x D X H)
Net weight
Gross weight
Animal Weighing Scales CTG 600A
Reliable and durable veterinary scale offering accurate, motion-stable measurements. Ideal for clinics, kennels, and grooming shops with portable design, removable stainless steel platform, and long-lasting battery operation.
| Capacity | 600 kg
| Accuracy | 100
| Pan Size | 800 x 800 mm
High Capacity | Motion Lock | Portable Design | Battery Operated | Easy Cleanup
Supports heavy animals with robust construction, ensuring precise weight measurements even under large loads and demanding veterinary applications. | Exclusive software ensures accurate weight readings even if the animal moves, maintaining reliability and stability for veterinary and grooming operations. | Compact, lightweight design allows easy relocation between clinics or grooming areas while maintaining accurate, professional-grade performance for animal weighing. | Runs on long-lasting battery power with auto power-off feature to conserve energy and provide consistent, uninterrupted operation. | Removable stainless steel platform simplifies cleaning and sanitation, ensuring hygienic operation and preventing cross-contamination between animals.
| Specifications - CTG 600A | Specifications - CTG1TA | Specifications - CTG 2TA1 | Specifications - CTG 2TA2
Maximum Capacity | 600 kg | 1 Ton | 2 Ton
Accuracy | 100 | 200 | 200 | 500
Readability
Pan Size | 800 x 800 mm | 1000 x 1000 mm | 1200 x 1500 mm | 1500 x 1500 mm
Display | LCD Backlite | LCD Backlite
Working Condition Temp
Power Supply
Balance Dimension (W x D X H)
Packaging dimensions (W x D X H)
Net weight
Gross weight
Double Pan Scales
A precision electronic balance engineered for safe and accurate weighing of centrifuge buckets and blood bags. Designed to be compact, durable, and user-friendly, it ensures reliable performance in laboratory and clinical environments.
| Capacity | 600 kg
| Accuracy | ±1gm
| Display | LCD backlit
Compact Design | High Accuracy | User-Friendly | Versatile Power | Approved Standards
Small footprint with sturdy construction allows effortless placement in labs while maintaining stability during sensitive weighing operations. | ±1 g precision ensures reliable measurement of blood bags or centrifuge samples for clinical and research purposes. | Jumbo LCD display and simple button layout allow quick, intuitive operation without specialized training. | Operates via 12 VDC adaptor or load cell interface, ensuring compatibility with multiple lab power setups. | Certified by Legal Metrology Dept., BIS ISI, and conforms to IS 9281 standards for accuracy and reliability.
Specifications - CTG 600A
Maximum Capacity | 6 kg
Accuracy | ±1gm
Readability
Pan Size
Display | LCD backlit
Working Condition Temp
Power Supply
Balance Dimension (W x D X H) | 300 x 375 x 12 mm
Packaging dimensions (W x D X H)
Net weight
Gross weight

## XRF
CUBE
A versatile XRF solution delivering elemental and coating analysis using micro-focus X-ray, proportional counter detection, precise CCD spot targeting, and an extensive calibration library with certified safety.
| Detector | Gas Filled Prop Counter
| Measuring Direction | Top to Bottom
| Measurement Time & Accuracy | 10 to 60 sec & ±0.3%
User Friendly Software | Safe | Material Analysis | Micro focus X-ray | Measuring Directions | Diverse Calibration Sample | Silicon Pin [Si-Pin] | Silicon drift Detector [FSDD]
Most comprehensive software for material analysis including Coating thickness measurement as well as bath analysis. Convenient calibration function High resolution CCD camera with optical zoom help positioning the spot for measurement. | The shutter is a safety device. It prevents primary radiation from entering the measuring chamber when measurement is in progress, thus safeguarding operator from risk of X-Ray exposures. X-Ray Auto cut off if accidentally door open. AERB approved | Versatile Material Analysis software x-Master based on both Empirical and Fundamental method help yield accurate Measurement as well as identify and measure any metal standard free. | The X-ray generates the primary X-ray radiation. Advanced micro-focus tubes are used to concentrate a large part of the primary radiation onto a very small measuring spot. | User-friendly precision: position the spot anywhere you see fit. Ideal for uneven surface jewellery and rings with inner & outer surface testing. | More than three decades of field experience have helped us build a wide calibration library, enabling diverse and application-specific calibration for every unit before it leaves our facility. | Gold Testing machine installed with original AMPTEK by USA . Suitable for simple measurement. Ideal for measuring complex shape samples and different measuring distance. | Gold Testing machine installed with original AMPTEK by USA . Suitable for simple measurement. Ideal for measuring complex shape samples and different measuring distance.
Sprcifications- | CUBE, prop counter | CUBE, High Resolution Si-PIN | CUBE, Fast Silicon Drift Detector(FSDD)
Detector | Gas Filled Proportional Counter | High Resolution Si-PIN | Fast Silicon Drift Detector(FSDD)
Measuring Direction | Top to Bottom | Top to Bottom | Top to Bottom
Applied Application | Gold, Silver, Platinium Analysis in Jewellery Alloys | Coating Thickness, Bath Analysis | Coating Thickness, Bath Analysis
Detector Range | Au, Ag, Cu, Zn, Ni, Pd, Cd, Pt, Fe, Co, Mn,Pb, Cr & more | Au, Ag, Cu, Zn, Ni, Pd, Cd, Rh, Ru, Pt, Fe, Co, In. Os, Ir, Mn,Pb,Ti, Hf, Cr, V, W, Ta & more | Au, Ag, Cu, Zn, Ni, Pd, Cd, Rh, Ru, Pt, Fe, Co, In. Os, Ir, Mn,Pb,Ti, Hf, Cr, V, W, Ta & more
Measurement Time | 10 to 60 sec. | 10 to 60 sec. | 10 to 60 sec.
Focus | Auto focus | Auto focus | Auto focus
X-ray Tube | Power 50 Watts Oil cooling X-ray Tube | Power 50 Watts Oil cooling X-ray Tube | Power 50 Watts Oil cooling X-ray Tube
High Voltage | 50kV (1 mA), Software Controlled Optimized | 50kV (1 mA), Software Controlled Optimized | 50kV (1 mA), Software Controlled Optimized
Collimator | 0.3 mm | 0.3 mm / 0.5 mm (Optional four positions collimator charger) | 0.3 mm / 0.5 mm (Optional four positions collimator charger)
Sample Stage | Manual scissors-type Z-stage | Manual scissors-type Z-stage | Manual scissors-type Z-stage
Power Supply | 230 VAC, 50/60Hz, 120W / 100W | 230 VAC, 50/60Hz, 120W / 100W | 230 VAC, 50/60Hz, 120W / 100W
Inside Chamber | 330 x 200 x 170 mm (W x D x H) | 330 x 200 x 170 mm (W x D x H) | 330 x 200 x 170 mm (W x D x H)
Dimension | 350 x 450 x 310 mm (W x D x H) | 350 x 450 x 310 mm (W x D x H) | 350 x 450 x 310 mm (W x D x H)
Operating System | Windows 10 | Windows 10 | Windows 10
Sample View | Crosshair-guided zoomable video & image capture for precise targeting | Crosshair-guided zoomable video & image capture for precise targeting | Crosshair-guided zoomable video & image capture for precise targeting
System | External PC set | External PC set | External PC set
Operating Conditions | 15°–25° with <90% humidity | 15°–25°C with <90% humidity | 15°–25°C with <90% humidity
Integrations | Compatible with electronic balances | Compatible with electronic balances | Compatible with electronic balances
Approvals | ISO Approved, AERB | ISO Approved, AERB | ISO Approved, AERB
Accuracy | ±0.3% | ±0.2% | ±0 1%
Dimension (W x D X H) | 350 × 450 × 310 mm | 350 × 450 × 310 mm | 350 × 450 × 310 mm
Packaging dimensions (W x D X H) | 560 x 560 x 690 mm | 560 x 560 x 690 mm | 560 x 560 x 690 mm
Net weight | 27 kg | 40 kg | 40 kg
Gross weight | 48 kg | 48 kg | 48 kg
Compact Eco
A versatile XRF solution delivering elemental and coating analysis using micro-focus X-ray, proportional counter detection, precise CCD spot targeting, and an extensive calibration library with certified safety.
| Detector | Gas Filled Prop Counter
| Measuring Direction | Top to Bottom
| Measurement Time & Accuracy | 10 to 60 sec & ±0.3%
User Friendly Software | Safe | Coating Thickness | Bath Analysis | Material Analysis | Micro focus X-ray | Diverse Calibration Sample | Digital Pulse Processor [DPP] | Silicon Drift Detector [FSDD] | Silicon PIN Technology
Most comprehensive software for material analysis including Coating thickness measurement as well as bath analysis. Convenient calibration function High resolution CCD camera with optical zoom help positioning the spot for measurement. | The shutter is a safety device. It prevents primary radiation from entering the measuring chamber when measurement is in progress, thus safeguarding operator from risk of X-Ray exposures. X-Ray Auto cut off if accidentally door open. AERB approved | From simple coating thickness measurements in the electroplating industry, such as zinc on iron, multi-coating applications, sophisticated precious metal coating | Bath analysis in the electroplating industry. Measurement of e.g. thin gold, palladium in jewellery industry. | Versatile Material Analysis software x-Master based on both Empirical and Fundamental method help yield accurate Measurement as well as identify and measure any metal standard free | The X-ray generates the primary X-ray radiation. Advanced micro-focus tubes are used to concentrate a large part of the primary radiation onto a very small measuring spot. | More than three decades of field experience have helped us build a wide calibration library, enabling diverse and application-specific calibration for every unit before it leaves our facility. | DPP generates very high pulse and count rates amplifying the signal detected by the detector. Working together with the detector, the DPP plays a crucial role in maintaining very high stability and energy resolution. | Gold Testing machine installed with AMPTEK by USA Suitable for simple measurement. Ideal for measuring complex shape samples and different measuring distance. | Gold Testing machine installed with original AMPTEK by USA. Suitable for simple measurement. Ideal for measuring complex shape samples and different measuring distance.
| Specifications - Compact Eco, prop counter | Specifications - Compact Eco, High Resolution Si-PIN | Specifications - Compact Eco, Fast Silicon Drift Detector(FSDD)
Detector | Gas Filled Proportional Counter | High Resolution Si-PIN | Fast Silicon Drift Detector(FSDD)
Measuring Direction | Top to Bottom | Top to Bottom | Top to Bottom
Applied Application | Gold, Silver, Platinium Analysis in Jewellery Alloys | Coating Thickness, Bath Analysis | Coating Thickness, Bath Analysis
Detector Range | Au, Ag, Cu, Zn, Ni, Pd, Cd, Pt, Fe, Co, Mn,Pb, Cr & more | Au, Ag, Cu, Zn, Ni, Pd, Cd, Rh, Ru, Pt, Fe, Co, In. Os, Ir, Mn,Pb,Ti, Hf, Cr, V, W, Ta & more | Au, Ag, Cu, Zn, Ni, Pd, Cd, Rh, Ru, Pt, Fe, Co, In. Os, Ir, Mn,Pb,Ti, Hf, Cr, V, W, Ta & more
Measurement Time | 10 to 60 sec. | 10 to 60 sec. | 10 to 60 sec.
Focus | Auto focus | Auto focus | Auto focus
X-ray Tube | Power 50 Watts Oil cooling X-ray Tube | Power 50 Watts Oil cooling X-ray Tube | Power 50 Watts Oil cooling X-ray Tube
High Voltage | 50kV (1 mA), Software Controlled Optimized | 50kV (1 mA), Software Controlled Optimized | 50kV (1 mA), Software Controlled Optimized
Collimator | 0.3 mm | 0.3 mm / 0.5 mm (Optional four positions collimator charger) | 0.3 mm / 0.5 mm (Optional four positions collimator charger)
Sample Stage | Manual scissors-type Z-stage | Manual scissors-type Z-stage | Manual scissors-type Z-stage
Power Supply | 230 VAC, 50/60Hz, 120W / 100W | 230 VAC, 50/60Hz, 120W / 100W | 230 VAC, 50/60Hz, 120W / 100W
Inside Chamber | 330 x 200 x 170 mm (W x D x H) | 330 x 200 x 170 mm (W x D x H) | 330 x 200 x 170 mm (W x D x H)
Dimension | 630 x 430 x 420 mm (W x D x H) | 630 x 430 x 420 mm (W x D x H) | 630 x 430 x 420 mm (W x D x H)
Operating System | Windows 10 | Windows 10 | Windows 10
Sample View | Crosshair-guided zoomable video & image capture for precise targeting | Crosshair-guided zoomable video & image capture for precise targeting | Crosshair-guided zoomable video & image capture for precise targeting
System | External PC set | Standard PC Set | Standard PC Set
Operating Conditions | 15°–25° with <90% humidity | 15°–25°C with <90% humidity | 15°–25°C with <90% humidity
Integrations | Compatible with electronic balances | Compatible with electronic balances | Compatible with electronic balances
Approvals | ISO Approved, AERB | ISO Approved, AERB | ISO Approved, AERB
Accuracy | ±0.3% | ±0.2% | ±0 1%
Dimension (W x D X H) | 630 × 430 × 420 mm | 630 × 430 × 420 mm | 630 × 430 × 420 mm
Packaging dimensions (W x D X H) | 680 x 790 x 800 mm | 680 x 790 x 800 mm | 680 x 790 x 800 mm
Net weight | 27 kg | 40 kg | 40 kg
Gross weight | 60 kg | 60 kg | 60 kg
StaRk
Fast, non-destructive material analysis with x-Master software (empirical + fundamental) and easy calibration. Bottom-to-top micro-focus X-ray supports quick sample positioning. Toshiba (Japan) proportional counter suits complex shapes and varying distances. Safety shutter with door-open auto cutoff; AERB approved. Extensive calibration library included.
| Detector | Gas Filled Prop Counter
| Measuring Direction | Bottom to Top
| Measurement Time & Accuracy | 10 to 60 sec & ±0.3%
User Friendly Software | Safe | Material Analysis | Micro focus X-ray | Measuring Directions | Diverse Calibration Sample | Silicon Drift Detector [FSDD] | Silicon PIN Technology
Most comprehensive software for material analysis including Coating thickness measurement as well as bath analysis. Convenient calibration function High resolution CCD camera with optical zoom help positioning the spot for measurement. | The shutter is a safety device. It prevents primary radiation from entering the measuring chamber when measurement is in progress, thus safeguarding operator from risk of X-Ray exposures. X-Ray Auto cut off if accidentally door open. AERB approved | Versatile Material Analysis software x-Master based on both Empirical and Fundamental method help yield accurate Measurement as well as identify and measure any metal standard free | The X-ray generates the primary X-ray radiation. Advanced micro-focus tubes are used to concentrate a large part of the primary radiation onto a very small measuring spot. | User-friendly precision: position the spot anywhere you see fit. Ideal for uneven surface jewellery and rings with inner & outer surface testing. | More than three decades of field experience have helped us build a wide calibration library, enabling diverse and application-specific calibration for every unit before it leaves our facility. | Gold Testing machine installed with AMPTEK by USA Suitable for simple measurement. Ideal for measuring complex shape samples and different measuring distance. | Gold Testing machine installed with original AMPTEK by USA. Suitable for simple measurement. Ideal for measuring complex shape samples and different measuring distance.
| Specifications - StaRk, prop counter | Specifications - StaRk, High Resolution Si-PIN | Specifications -StaRk, Fast Silicon Drift Detector(FSDD)
Detector | Gas Filled Proportional Counter | High Resolution Si-PIN | Fast Silicon Drift Detector(FSDD)
X-ray Tube | Mini-focus, high performance W-target, Spot 0.2mm - 0.8mm Long Life, High efficiency, tungsten target air cooled | Mini-focus, high performance W-target, Spot 0.2mm - 0.8mm Long Life, High efficiency, tungsten target air cooled | Mini-focus, high performance W-target, Spot 0.2mm - 0.8mm Long Life, High efficiency, tungsten target air cooled
Measuring Direction | Bottom to Top | Bottom to Top | Bottom to Top
System | Standard PC set | Standard PC set | Standard PC set
Collimator | 0.3mm0 or 0.5mm0, Optional four positions Collimator changer | 0.3mm0 or 0.5mm0, Optional four positions Collimator changer | 0.3mm0 or 0.5mm0, Optional four positions Collimator changer
Detector Range | Au, Ag, Cu, Cd, Zn, Ni ,Pd | Au, Ag, Cu, Zn, Ni, Pd, Cd, Rh, Ru, Pt and more. | Au, Ag, Cu, Zn, Ni, Pd, Cd, Rh, Ru, Pt and more.
Analysis Range | 0.01% - 99.99% | 0.2% - 99.99% | 0.01% - 99.99%
Measurement Time | 10 – 60 seconds | 10 – 60 seconds | 10 – 60 seconds
High Voltage | 50kV / 1.2mAJ Software Control Optimized | 50kV / 1.2mAJ Software Control Optimized | 50kV / 1.2mAJ Software Control Optimized
Sample Positioning | Auto focussing | Auto focussing | Auto focussing
Video Microscope | High resolution CCD Color Video camera with cross hair | High resolution CCD Color Video camera with cross hair | High resolution CCD Color Video camera with cross hair
PC Interface | External | External | External
Power Supply | 110V – 260V, 50/60 Hz | 110V – 260V, 50/60 Hz | 110V – 260V, 50/60 Hz
Temperature range | 15°C - 25°C | 15°C - 25°C | 15°C - 25°C
Operating System | Windows 10 | Windows 10 | Windows 10
Applied Application | Gold & Silver Analysis in jewellery and alloys, Coating Thickness Measurement on Metals | Gold & Silver Analysis in jewellery and alloys, Coating Thickness Measurement on Metals | Gold & Silver Analysis in jewellery and alloys, Coating Thickness Measurement on Metals
Chamber | Large Chamber to placewide Object | Large Chamber to place wide Object | Large Chamber to place wide Object
Sample Chamber | 300 x 290 x 100 mm (W x D x H) | 300 x 290 x 100 mm (W x D x H) | 300 x 290 x 100 mm (W x D x H)
Accuracy | ± 0.3% | ± 0.2% | ± 0.1%
Dimension | 350 x 450 x 310 mm (W x D x H) | 350 x 450 x 310 mm (W x D x H) | 350 x 450 x 310 mm (W x D x H)
Packaging dimensions (W x D X H) | 590 x 610 x 650 mm (W x D x H) | 590 x 610 x 650 mm (W x D x H) | 590 x 610 x 650 mm (W x D x H)
Net weight | 32 kg | 32 kg | 32 kg
Gross weight | 56 kg | 56 kg | 56 kg
Stark Inbuilt
A compact, high-performance XRF system with advanced FSDD technology for fast, accurate metal and alloy analysis. Detects multiple elements simultaneously, ideal for gold purity testing, jewellery verification, and measurement with reliable, high-resolution results.
| Detector | Fast Silicon Drift Detector (FSDD)
| Measuring Direction | Bottom to Top
| Measurement Time & Accuracy | 10 to 60 sec & ± 0.1%
User Friendly Software | Safe | Material Analysis | Micro focus X-ray | Measuring Directions | Proportional Counter | Silicon PIN (Si-PIN) Detector | Fast Silicon Drift Detector (FSDD)
Powerful software for gold analysis with easy calibration and a high-resolution CCD camera for precise sample positioning. | The shutter is a safety device. It prevents primary radiation from entering the measuring chamber when measurement is in progress, thus safeguarding operator from risk of X-Ray exposures. X-Ray Auto cut off if accidentally door open. AERB approved. | Versatile Material Analysis software X-Master based on both Empirical and Fundamental method help yield accurate Measurement as well as identify and measure any metal standard free. | The X-ray generates the primary X-ray radiation. Advanced micro-focus tubes are used to concentrate a large part of the primary radiation onto a very small measuring spot. | User-friendly precision: position the spot anywhere you see fit. Ideal for uneven surface jewellery and rings with outer surface testing. | Gold Testing machine installed with Toshiba - Made in Japan Detector. Suitable for simple measurement. Ideal for measuring complex shape samples and different measuring distance. | The Si-PIN detector delivers reliable and stable performance for everyday metal and jewellery testing. It is ideal for routine alloy analysis, offering consistent results with good resolution for standard applications. | Equipped with a high-resolution AMPTEK FSDD detector, the StaRk Inbuilt FSDD delivers excellent peak separation and high count rate performance for accurate results.
Sprcifications- | StaRk Inbuint, Prop Counter | StaRk Inbuint, High Resolution Si-PIN | StaRk Inbuilt, Fast Silicon Drift Detector(FSDD)
Detector | Gas Filled Proportional Counter | High Resolution Si-PIN | Fast Silicon Drift Detector(FSDD)
X-ray Tube | Power 50 Watts Oil cooling X-ray Tube | Power 50 Watts Oil cooling X-ray Tube | Power 50 Watts Oil cooling X-ray Tube
Measuring Direction | Bottom to Top | Bottom to Top | Bottom to Top
System | Inbulit PC | Inbulit PC | Inbulit PC
Collimator | 0.3 mm | 2.00 mm (1.00 mm optional) | 1.50 mm
Detector Range | Au, Ag, Cu, Zn, Ni, Pd, Cd, Pt, Fe, Co, Mn,Pb, Cr & more | Au, Ag, Cu, Zn, Ni, Pd, Cd, Rh, Ru, Pt, Fe, Co, In, Os, Ir, Mn, Pb, Ti, Hf, Cr, V, W, Ta & more | Au, Ag, Cu, Zn, Ni, Pd, Cd, Rh, Ru, Pt, Fe, Co, In, Os, Ir, Mn, Pb, Ti, Hf, Cr, V, W, Ta & more
Analysis Range | 0.01% - 99.99% | 0.2% - 99.99% | 0.1% - 99.99%
Measurement Time | 10 to 60 sec. | 10 to 60 sec. | 10 to 60 sec.
High Voltage | 50 kV (1 mA), Software Control Optimized | 50 kV (1 mA), Software Control Optimized | 50 kV (1 mA), Software Control Optimized
Sample Positioning | Auto focussing | Auto focussing | Auto focussing
Video Microscope | High resolution CCD Color Video camera with cross hair | High resolution CCD Color Video camera with cross hair | High resolution CCD Color Video camera with cross hair
PC Interface | Touch Screen Industrial Grade PC | Touch Screen Industrial Grade PC | Touch Screen Industrial Grade PC
Power Supply | 230 VAC, 50/60Hz, 120W / 100W | 230 VAC, 50/60Hz, 120W / 100W | 230 VAC, 50/60Hz, 120W / 100W
Temperature range | 15°C - 25°C | 15°C - 25°C | 15°C - 25°C
Operating System | Windows 10 | Windows 11 | Windows 11
Applied Application | Gold, Silver, Platinum Analysis in Jewellery Alloys | Gold, Silver, Platinum Analysis in Jewellery Alloys | Gold, Silver, Platinum Analysis in Jewellery Alloys
Chamber | Large Chamber to placewide Object | Large Chamber to placewide Object | Large Chamber to placewide Object
Sample Chamber | 300 x 290 x 100 mm (W x D x H) | 300 x 290 x 100 mm (W x D x H) | 300 x 290 x 100 mm (W x D x H)
Accuracy | ±0.3% | ±0.2% | ±0.1%
Dimension (W x D X H) | 395 × 390 × 320 mm | 395 × 390 × 320 mm | 395 × 390 × 320 mm
Packaging dimensions (W x D X H) | 550 × 510 × 680 mm | 550 × 510 × 680 mm | 550 × 510 × 680 mm
Net weight | 30 kg | 30 kg | 30 kg
Gross weight | 45 kg | 45 kg | 45 kg
Stark Pro
A compact, high-performance XRF system with advanced FSDD technology for fast, accurate metal and alloy analysis. Detects multiple elements simultaneously, ideal for gold purity testing, jewellery verification, and coating thickness measurement with reliable, high-resolution results.
| Detector | Fast Silicon Drift Detector (FSDD)
| Measuring Direction | Bottom to Top
| Measurement Time & Accuracy | 10 to 60 sec & ± 0.2%
User Friendly Software | Safe | Material Analysis | Micro focus X-ray | Measuring Directions | Diverse Calibration Sample | Silicon PIN (Si-PIN) Detector | Fast Silicon Drift Detector (FSDD)
Most comprehensive software for material analysis including measurement as well as bath analysis. Convenient calibration function High resolution CCD camera with optical zoom help positioning the spot for measurement. | The shutter is a safety device. It prevents primary radiation from entering the measuring chamber when measurement is in progress, thus safeguarding operator from risk of X-Ray exposures. X-Ray Auto cut off if accidentally door open. AERB approved. | Versatile Material Analysis software X-Master based on both Empirical and Fundamental method help yield accurate Measurement as well as identify and measure any metal standard free. | The X-ray generates the primary X-ray radiation. Advanced micro-focus tubes are used to concentrate a large part of the primary radiation onto a very small measuring spot. | User-friendly precision: position the spot anywhere you see fit. Ideal for uneven surface jewellery and rings with inner & outer surface testing. | More than three decades of field experience have helped us build a wide calibration library, enabling diverse and application-specific calibration for every unit before it leaves our facility. | The Si-PIN detector delivers reliable and stable performance for everyday metal and jewellery testing. It is ideal for routine alloy analysis, offering consistent results with good resolution for standard applications. | Equipped with a high-resolution AMPTEK SDD detector, the StaRk Pro SDD delivers excellent peak separation and high count rate performance for accurate results.
Sprcifications- | StaRk Pro , High Resolution Si-PIN | StaRk Pro , Fast Silicon Drift Detector(FSDD)
Detector | High Resolution Si-PIN | Fast Silicon Drift Detector(FSDD)
X-ray Tube | Power 50 Watts Oil Cooling X-ray Tube | Power 50 Watts Oil Cooling X-ray Tube
Measuring Direction | Bottom to Top | Bottom to Top
System | Standard PC set | Standard PC set
Collimator | 2.00 mm (1.00 mm optional) | 1.50 mm
Detector Range | Au, Ag, Cu, Zn, Ni, Pd, Cd, Rh, Ru, Pt, Fe, Co, In, Os, Ir, Mn, Pb, Ti, Hf, Cr, V, W, Ta & more | Au, Ag, Cu, Zn, Ni, Pd, Cd, Rh, Ru, Pt, Fe, Co, In, Os, Ir, Mn, Pb, Ti, Hf, Cr, V, W, Ta & more
Analysis Range | 0.2% - 99.99% | 0.1% - 99.99%
Measurement Time | 10 to 60 sec. | 10 to 60 sec.
High Voltage | 50 kV (1 mA), Software Control Optimized | 50 kV (1 mA), Software Control Optimized
Sample Positioning | Auto focussing | Auto focussing
Video Microscope | High resolution CCD Color Video camera with cross hair | High resolution CCD Color Video camera with cross hair
PC Interface | External | External
Power Supply | 230VAC, 50/60 Hz, 120W / 100W (110V avaliable on request) | 230VAC, 50/60 Hz, 120W / 100W (110V avaliable on request)
Temperature range | 15°C - 25°C | 15°C - 25°C
Operating System | Window 11 | Window 11
Applied Application | Gold, Silver, Platinum Analysis in Jewellery Alloys | Gold, Silver, Platinum Analysis in Jewellery Alloys
Chamber | Large Chamber to placewide Object | Large Chamber to placewide Object
Sample Chamber | 300 x 290 x 100 mm (W x D x H) | 300 x 290 x 100 mm (W x D x H)
Accuracy | 0.2% | 0.1%
Dimension (W x D X H) | 395 x 390 x 320 mm (W x D x H) | 395 x 390 x 320 mm (W x D x H)
Packaging dimensions (W x D X H) | 590 x 610 x 650 mm (W x D x H) | 590 x 610 x 650 mm (W x D x H) | 590 x 610 x 650 mm (W x D x H)
Net weight | 28 kg | 28 kg
Gross weight | 56 kg | 56 kg | 56 kg
Stark Pro Inbuilt
A compact, high-performance XRF system with advanced FSDD technology for fast, accurate metal and alloy analysis. Detects multiple elements simultaneously, ideal for gold purity testing, jewellery verification, and coating thickness measurement with reliable, high-resolution results.
| Detector | Fast Silicon Drift Detector (FSDD)
| Measuring Direction | Bottom to Top
| Measurement Time & Accuracy | 10 to 60 sec & ± 0.2%
User Friendly Software | Safe | Material Analysis | Micro focus X-ray | Measuring Directions | Diverse Calibration Sample | Silicon PIN (Si-PIN) Detector | Fast Silicon Drift Detector (FSDD)
Most comprehensive software for material analysis including & measurement as well as bath analysis. Convenient calibration function High resolution CCD camera with optical zoom help positioning the spot for measurement. | The shutter is a safety device. It prevents primary radiation from entering the measuring chamber when measurement is in progress, thus safeguarding operator from risk of X-Ray exposures. X-Ray Auto cut off if accidentally door open. AERB approved. | Versatile Material Analysis software X-Master based on both Empirical and Fundamental method help yield accurate Measurement as well as identify and measure any metal standard free. | The X-ray generates the primary X-ray radiation. Advanced micro-focus tubes are used to concentrate a large part of the primary radiation onto a very small measuring spot. | User-friendly precision: position the spot anywhere you see fit. Ideal for uneven surface jewellery and rings with inner & outer surface testing. | More than three decades of field experience have helped us build a wide calibration library, enabling diverse and application-specific calibration for every unit before it leaves our facility. | The Si-PIN detector delivers reliable and stable performance for everyday metal and jewellery testing. It is ideal for routine alloy analysis, offering consistent results with good resolution for standard applications. | Equipped with a high-resolution AMPTEK FSDD detector, the StaRk Pro Inbuilt FSDD delivers excellent peak separation and high count rate performance for accurate results.
Sprcifications- | StaRk Pro Inbuilt, High Resolution Si-PIN | StaRk Pro Inbuilt , Fast Silicon Drift Detector(FSDD)
Detector | High Resolution Si-PIN | Fast Silicon Drift Detector(FSDD)
X-ray Tube | Powerful 50 Watts Oil Cooling X-ray Tube | Powerful 50 Watts Oil Cooling X-ray Tube
Measuring Direction | Bottom to Top | Bottom to Top
System | Inbulit PC | Inbulit PC
Collimator | 2.0 mm (1.00 mm optional) | 1.50 mm
Detector Range | Au, Ag, Cu, Zn, Ni, Pd, Cd, Rh, Ru, Pt, Fe, Co, In, Os, Ir, Mn, Pb, Ti, Hf, Cr, V, W, Ta & more | Au, Ag, Cu, Zn, Ni, Pd, Cd, Rh, Ru, Pt, Fe, Co, In, Os, Ir, Mn, Pb, Ti, Hf, Cr, V, W, Ta & more
Analysis Range | 0.2% - 99.99% | 0.01% - 99.99%
Measurement Time | 10 - 60 sec. | 10 - 60 sec.
High Voltage | 50 kV (1 mA), Software Control Optimized | 50 kV (1 mA), Software Control Optimized
Sample Positioning | Auto focussing | Auto focussing
Video Microscope | High resolution CCD Color Video camera with cross hair | High resolution CCD Color Video camera with cross hair
PC Interface | Touch Screen Industrial Grade PC | Touch Screen Industrial Grade PC
Power Supply | 230VAC, 50/60 Hz, 120W / 100W (110V avaliable on request) | 230VAC, 50/60 Hz, 120W / 100W (110V avaliable on request) |  | `
Temperature range | 15°C - 25°C | 15°C - 25°C
Operating System | Windows 11 | Windows 11
Applied Application | Gold, Silver, Platinum Analysis in Jewellery Alloys | Gold, Silver, Platinum Analysis in Jewellery Alloys
Chamber | Large Chamber to placewide Object | Large Chamber to placewide Object
Sample Chamber | 300 x 290 x 100 mm (W x D x H) | 300 x 290 x 100 mm (W x D x H)
Accuracy | ±0.2% | ±0.1%
Dimension (W x D X H) | 395 x 390 x 320 mm (W x D x H) | 395 x 390 x 320 mm (W x D x H)
Packaging dimensions (W x D X H) | 550 × 510 × 680 mm | 550 × 510 × 680 mm | 550 × 510 × 680 mm
Net weight | 30 kg | 30 kg | 30 kg
Gross weight | 45 kg | 45 kg | 45 kg
Axiom X-ray Fluorescence (XRF)
Analysis for PCBs delivers high-precision, non-destructive elemental verification for quality control and compliance. It detects a broad range of elements and provides rapid on-site results in minutes. Integrated software supports coating thickness and material analysis. A slotted chamber fits large flat boards, with optional XY joystick control.
| Detector | Fast Silicon Drift Detector(FSDD)
| Measuring Direction | Bottom to Top
| Measurement Time & Accuracy | 10 to 60 sec & ±0.1
User Friendly Software | Safe | Material Analysis | Micro focus X-ray | Measuring Directions | Diverse Calibration Sample | Proportional Counter | Silicon PIN (Si-PIN) Detector | Fast Silicon Drift Detector (SDD)
Easy-to-use software helps users perform material analysis, coating thickness measurement, and bath analysis with confidence. The calibration function makes setup simple, while the high-resolution CCD camera with optical zoom helps position the measurement spot accurately. | The safety shutter prevents primary X-ray radiation from entering the measuring chamber during operation, helping protect the operator from exposure. The system also includes automatic X-ray cut-off if the door is opened accidentally and is AERB approved. | XMasteR software supports material analysis, coating thickness measurement, and bath analysis using both empirical and Fundamental Parameter methods. It helps identify and measure metals accurately, even without standard samples. | The micro-focus X-ray tube generates primary X-ray radiation and focuses it onto a very small measurement spot, enabling precise analysis of small areas and fine sample features. | User-friendly precision: position the spot anywhere you see fit. Ideal for uneven surface jewellery and rings with inner & outer surface testing. | More than three decades of field experience have helped us build a wide calibration library, enabling diverse and application-specific calibration for every unit before it leaves our facility. | The proportional counter features a large active detection area, providing excellent count rates and highly repeatable measurements. It is ideal for general coating thickness measurement, bath analysis, and complex-shaped samples. | The Si-PIN detector delivers stable and reliable performance for coating thickness and material analysis. It offers high energy resolution and accurate measurements, making it suitable for routine and demanding applications. | The SDD provides superior energy resolution, high count rates, and excellent sensitivity for precise elemental analysis. It is ideal for ultra-thin coatings, multilayer measurements, and trace element detection.
Sprcifications- | Axiom, Prop Counter | Axiom, High Resolution Si-PIN | Axiom, Fast Silicon Drift Detector(FSDD)
Detector | Gas Filled Proportional Counter | High Resolution Si-PIN | Fast Silicon Drift Detector(FSDD)
X-ray Tube | High performing, stable with long life tungsten tube. Spot size approx. 0.5 × 0.5 mm to assure minimal beam spread | High performing, stable with long life tungsten tube. Spot size approx. 0.5 × 0.5 mm to assure minimal beam spread | High performing, stable with long life tungsten tube. Spot size approx. 0.5 × 0.5 mm to assure minimal beam spread
Measuring Direction | Bottom to Top | Bottom to Top | Bottom to Top
System | External PC set | External PC set | External PC set
Collimator | Single Collimator 0.3 mm Ø (Optional sizes: 0.1 mm, 0.2 mm, 0.5 mm, 0.7 mm, 1.00 mm) | Single Collimator 0.3 mm Ø (Optional sizes: 0.1 mm, 0.2 mm, 0.5 mm, 0.7 mm) | Single Collimator 0.3 mm Ø (Optional sizes: 0.1 mm, 0.2 mm, 0.5 mm, 0.7 mm)
Detector Range | Au, Ag, Cu, Zn, Ni, Pd, Cd, Pt, Fe, Co, Mn,Pb, Cr & more | Au, Ag, Cu, Zn, Ni, Pd, Cd, Rh, Ru, Pt, Fe, Co, In. Os, Ir, Mn,Pb,Ti, Hf, Cr, V, W, Ta & more | Au, Ag, Cu, Zn, Ni, Pd, Cd, Rh, Ru, Pt, Fe, Co, In. Os, Ir, Mn,Pb,Ti, Hf, Cr, V, W, Ta & more
Analysis Range | Na (Sodium) to U (Uranium) | Na (Sodium) to U (Uranium) | Na (Sodium) to U (Uranium)
Measurement Time | 10 to 60 sec. | 10 to 60 sec. | 10 to 60 sec.
High Voltage | 50 kV, 1.2 mA (60 W) | 50 kV, 1.2 mA (60 W) | 50 kV, 1.2 mA (60 W)
Sample Positioning | Motorised Programmable X, Y | Motorised Programmable X, Y | Motorised Programmable X, Y
Video Microscope | High-resolution CCD color camera for optical monitoring of the measurement location with 20X zoom, Electronic cross hair with scale and beam size indication. | High-resolution CCD color camera for optical monitoring of the measurement location with 20X zoom, Electronic cross hair with scale and beam size indication. | High-resolution CCD color camera for optical monitoring of the measurement location with 20X zoom, Electronic cross hair with scale and beam size indication.
PC Interface | External | External | External
Power Supply | AC 110V or AC 230V, 50-60 Hz | AC 110V or AC 230V, 50-60 Hz | AC 110V or AC 230V, 50-60 Hz
Temperature range | 15°C - 25°C | 15°C - 25°C | 15°C - 25°C
Operating System | Window 10 | Window 10 | Window 10
Applied Application | Gold & Silver Analysis in jewellery and alloys, Coating Thickness Measurement on Metals | Gold & Silver Analysis in jewellery and alloys, Coating Thickness Measurement on Metals | Gold & Silver Analysis in jewellery and alloys, Coating Thickness Measurement on Metals
Chamber | Large Slotted Chamber | Large Slotted Chamber | Large Slotted Chamber
Sample Chamber | 460 × 500 × 190 mm | 460 × 500 × 190 mm | 460 × 500 × 190 mm
Accuracy | ±0.1% | ±0.1% | ±0.1%
Dimension (W x D X H) | 500 × 740 × 530 mm | 500 × 740 × 530 mm | 500 × 740 × 530 mm
Packaging dimensions (W x D X H)
Net weight | 108 kg | 108 kg | 108 kg
Gross weight
Hand Held XRF Nanoris
A portable, non-destructive analyzer for rapid elemental testing of metals, alloys, precious metals and coatings. It identifies alloy grades in seconds and checks plating thickness without damaging samples. Lightweight for real-time on-site PMI and scrap sorting. Imported detector ensures low-noise multi-element performance, with PC/mobile remote-control data sync.
| Detector | Gas Filled Prop Counter
| Measuring Direction | (RSD) ≤ 1%
| Measurement Element Range | Mg (12) - U (92)
NDT | Computer Software | Portability | High quality imported detectors | Automobile Industry | XRF Precious Metal Analyzer | Alloy Field waste recyling (Scrap)
Quickly and accurately analyze the plating thickness and sample element content of the plating sample without damaging the test object. | Remote control system, C # data synchronization control system based on Windows system, Both computer and mobile software can operate instruments and synchronize data through software. . | The instrument has a smaller volume and lighter weight, making it convenient for real-time on-site testing and can easily handle samples of various shapes | The instrument adopts imported detectors suitable for analyzing multi-element coatings, with low noise and ?exible response to coating structure. | In the automotive industry, the Tektronix spectrometer is used in the field of ternary catalysis to evaluate whether automobile exhaust emissions meet standards. It can detect harmful gases and particulate matter in car exhaust, as well as their concentration and composition, providing data support for formulating car emission standards | The detection of the content of dozens of elements such as gold, silver, and platinum group metals (ruthenium, rhodium, palladium, osmium, iridium, platinum) from K to U shows that the content can be both% and K gold | For the scrap metal recycling industry, an accurate and portable handheld spectrometer can help identify alloys and analyze material chemical composition, while the Aczet handheld spectrometer can provide fast and accurate elemental analysis for scrap metal recycling. Relying on Aczet digital and intelligent industry solutions, Aczet’s handheld spectral analyzer can measure element content as low as several parts per million.
| Specifications - Hand Held XRF Nanoris Si-PIN | Specifications -Hand Held XRF Nanoris, Fast Silicon Drift Detector(FSDD)
Detector | Si - Pin | Fast Silicon Drift Detector(FSDD)
Element Range | Ti (22) - U (92) | Mg (12) - U (92
Measurement Accuracy | (RSD) < 1% | (RSD) < 1%
Data Transmission | USB, E-mail, Wi-Fi, Bluetooth | USB, E-mail, Wi-Fi, Bluetooth
Network Connectivity | Wi-Fi, Bluetooth | Wi-Fi, Bluetooth
Control System | Customized Industrial - Grade Control System : CPU : 2.4 Ghz, RAM : 2GB, Internal Storage : 4GB (Expandable upto 32GB) | Customized Industrial - Grade Control System : CPU : 2.4 Ghz, RAM : 2GB, Internal Storage : 4GB (Expandable upto 32GB)
Application mode | Precious Alloy, RoHs, Alloys | Precious Alloys, Light Alloys, Mining, Minerals
Camera | High Definition Camera | High Definition Camera
Detector Window Protection | Shield Window | Shield Window
Analysis Methods | Empirical coefficient method, Intelligent FP Method | Empirical coefficient method, Intelligent FP Method
Excitation Source | 50kV / 200mA, X-ray fluorescence / 500mA (Optional) | 50kV / 200mA, X-ray fluorescence / 500mA (Optional
Detector Cooling Method | Two-stage thermoelectric cooling | Two-stage thermoelectric cooling
Safety | Proximity Sensor for Radiation Safety / Radiation Protection (Optional) | Proximity Sensor for Radiation Safety / Radiation Protection (Optional)
Battery Life | 6 - 7 Hours. | 6 - 7 Hours.
Dimensions | 254 mm (L) x 91 mm (W) x 319 mm (H) | 254 mm (L) x 91 mm (W) x 319 mm (H)
Material | Alloy, polycarbonate | Alloy, polycarbonate
Dimension (W x D X H) | 91 × 254 × 319 mm | 91 × 254 × 319 mm
Packaging dimensions (W x D X H) | 550 x 440 x 500 mm | 550 x 440 x 500 mm
Net weight | 1.75 kg (Including battery) | 1.75 kg (Including battery)
Gross weight
Melting Furnaces
IM Series Melting Furnaces use medium-frequency induction heating to generate heat directly within the material, ensuring fast, uniform, and energy-efficient melting with minimal heat loss and stable, safe operation.
| Capacity | 1 kg Gold
| Input | 3500 W
| Frequency | 50/60 Hz
Induction Heating Technology | Fast & Uniform Melting | Energy Efficient Operation | Safe & Stable Performance
Uses medium-frequency induction to generate heat directly within the material for efficient melting | Ensures rapid melting with consistent temperature distribution across the entire material | Minimizes heat loss by focusing heat within the crucible, improving overall efficiency | Provides controlled heating with minimal external heat, ensuring safe and reliable operation
Model | IM 1 | IM 2 | IM 5
Melting Capacity | 1 kg Gold | 2 kg Gold | 5 kg Gold
Power | 3500 W | 5000 W | 8000 W
Voltage | 220V 50/60 Hz | 220V 50/60 Hz | 220V 50/60 Hz
Current | 16 Amp Single Phase | 16 Amp Single Phase | 16 Amp Single Phase
Can Melt Metal | Gold / Silver / Copper / Iron / Steel | Gold / Silver / Copper / Iron / Steel | Gold / Silver / Copper / Iron / Steel
Max Temperature | 1600°C | 1600°C | 1600°C
Cooling System | Water Cooling Inbuilt Chiller | Water Cooling Inbuilt Chiller | Water Cooling Inbuilt Chiller
Melting Time | 3–5 mins | 3–5 mins | 3–5 mins
Chiller | In-built | In-built | In-built
Stabilization | Outside body remains cool, continuous 24 hrs operation | Outside body remains cool, continuous 24 hrs operation | Outside body remains cool, continuous 24 hrs operation
Velocity | 8–10 mins to melt 1 kg gold | 8–10 mins to melt 1 kg gold | 8–10 mins to melt 1 kg gold
Operation | Easy to operate, energy saving & time saving | Easy to operate, energy saving & time saving | Easy to operate, energy saving & time saving
Dimension | 550 × 305 × 400 mm | 550 × 305 × 400 mm | 560 × 370 × 750 mm
Net Weight | 28 kg | 28 kg | 51 kg
Ultrasonic Cleaner CUB 2.0
Designed for deep, efficient cleaning of jewellery, electronics, medical tools, and industrial parts, these ultrasonic cleaners feature 40 KHz power, 304 stainless steel tanks, digital timer/heater, degassing mode, and built-in safety protection.
| Capacity | 1.8 Ltr
| Input | 110/220 AC, 50/60 Hz
| Frequency | 40 Khz
Stainless Steel Tank | Built-in Transducer with Degassing | Enhanced Cleaning Action | Efficient Cooling System
Durable 304 stainless steel tank ensures corrosion resistance, safety, and long-lasting performance | Integrated industrial transducer with degassing function enhances ultrasonic efficiency for deeper and more effective cleaning | Improved cavitation process allows ultrasonic waves to penetrate deeply, ensuring thorough and uniform cleaning | Advanced cooling system maintains stable operation, improving performance and extending machine life
Specifications | CUB 2.0 | CUB 5 | CUB 10 | CUB 15 | CUB 22
Capacity (L) | 1.8 Ltr. | 4.5 Ltr. | 9.8 Ltr. | 13.8 Ltr. | 20 Ltr.
Frequency | 40 KHz | 40 KHz | 40 KHz | 40 KHz | 40 KHz
Adjustable Time | 1 to 99 minutes | 1 to 99 minutes | 1 to 99 minutes | 1 to 99 minutes | 1 to 99 minutes
Adjustable Temperature | 0 to 80°C | 0 to 80°C | 0 to 80°C | 0 to 80°C | 0 to 80°C
Ultrasonic Power | 60 W | 120 W | 240 W | 360 W | 480 W
Heating Power | 100 W | 300 W | 300 W | 400 W | 500 W
Tank Size (L×W×H) mm | 150 × 140 × 100 | 300 × 155 × 100 | 300 × 240 × 150 | 330 × 300 × 150 | 500 × 300 × 150
Overall Size (L×W×H) mm | 190 × 170 × 220 | 312 × 162 × 240 | 330 × 270 × 310 | 360 × 330 × 330 | 550 × 340 × 330
Power Requirements | 230 Vac, ±20%, 50–60 Hz | 230 Vac, ±20%, 50–60 Hz | 230 Vac, ±20%, 50–60 Hz | 230 Vac, ±20%, 50–60 Hz | 230 Vac, ±20%, 50–60 Hz
Weight Approx. | 4.2 kg | 6 kg | 11 kg | 21 kg | 21 kg

## Jewellery Accesories
Accessories for Jewellery Machines:
Dot Matrix Printer CPR 03
Compact Dot Matrix Printer CPR-03 with RS-232 interface provides reliable laboratory data recording, accurate result printouts, enhanced traceability, and regulatory compliance. Universal compatibility enables seamless connection with most balance and instrument brands.
| Model | CPR 03
| Consumable | Paper Roll / Black ink printer ribbon cartridge
| Printers | Dot Matrix Printer
| No of Characters per line | 24 Character
| Printing Speed | Faster Than 1 line per second
| Paper Type | Normal Paper Roll, Wood Free Paper
| Paper size | Width : 57.5 ±0.5mm. Core Diameter : 8mm Min.
| Power Supply | AC Adopter, 100-270V AC, 50Hz - 60Hz
| Interface | RS232
| Bound Fate | 1200, 2400, 4800, 9600, 19200
Dot Matrix Printer CPR 02
Compact Dot Matrix Printer CPR-02 with RS-232 interface provides reliable laboratory data recording, accurate result printouts, enhanced traceability, and regulatory compliance for weighing and measurement applications.
| Model | CPR 02
| Consumable | Paper Roll / Black ink printer ribbon cartridge
| Printers | Dot Matrix Printer
| No of Characters per line | 24 Character
| Printing Speed | Faster Than 1 line per second
| Paper Type | Normal Paper Roll, Wood Free Paper
| Paper size | Width : 57.5 ±0.5mm. Core Diameter : 8mm Min.
| Power Supply | AC Adopter, 100-270V AC, 50Hz - 60Hz
| Interface | RS232
| Bound Fate | 1200, 2400, 4800, 9600, 19200
Label printer CLP 04
CLP-04 label printer supports barcode and matrix code printing with specialized volumetric flask labels for easy scanning and removal.
| Model | CLP 04
Barcode Reader CBR01
CBR-01 Barcode Reader enables fast, accurate barcode data capture.
| Model | CBR01
Antistatic Ionizer CI 01
The CI 01 Antistatic Ionizer ensures safe and effective static elimination, operating reliably between 0°C and 40°C for stable laboratory and weighing applications.
| Model | CI 01
Anti Vibration Weighing Table VRD - AVT 600
The VRD-AVT 600 anti-vibration weighing table provides a stable, durable platform for precision balances, minimizing vibrations and ensuring accurate results in laboratory environments.
| Model | VRD-AVT 600
Vibrosorb-Anti Vibration Pad VRD 03
The VRD 03 Vibrosorb Anti-Vibration Pad, large size, offers strong vibration absorption with 400 × 240 × 25 mm dimensions and 8.5 kg mass for stable precision weighing.
| Model | VRD 03
Remote Display
The SRD 01 Remote Display provides clear external weight indication with 19 mm digit size. It operates on 6V DC power from the measuring indicator scale and includes a 1.5 m connection cable for convenient placement.
| Model | SRD 01
Multicollimator Changer
The Multicollimator Changer allows easy selection and switching of collimators, helping optimize XRF measurement accuracy, improve analysis flexibility, and support different sample testing requirements.
| Model | XMC
BE Window Xray tube
The BE Window X-ray Tube provides efficient low-energy X-ray transmission for accurate XRF analysis, improving sensitivity for light elements and ensuring reliable elemental testing results.
| Model | XBW
Red Light Pointer
The Red Light Pointer helps accurately identify the measurement spot before XRF analysis, ensuring precise sample positioning, easy operation, and reliable testing results.
| Model | XRP
XY motrorised Stage with Joystick
The XY Motorised Stage with Joystick enables smooth and precise sample movement during XRF analysis, allowing accurate positioning, easy control, and reliable testing of different sample areas.
| Model | XMS
Calibration Standard
The Calibration Standard ensures accurate XRF measurement verification by providing a reliable reference material for calibration, performance checking, and consistent elemental analysis results.
| Model | XCS
Lab Jack Stand Table
The Lab Jack Stand Table provides stable height adjustment and firm sample support during XRF analysis, enabling accurate positioning, easy operation, and reliable measurement results.
| Model | XLJ
Liquid sample container
The Liquid Sample Container is designed to safely hold liquid samples during XRF analysis, ensuring secure placement, accurate measurement, and reliable elemental testing results.
| Model | XLC
Holding Stand
The Holding Stand provides secure and stable sample positioning during XRF analysis, helping ensure accurate alignment, easy handling, and reliable measurement results.
| Model | XHS
Standard Rotary
Rotary attachment for marking and cutting rings, bracelets, and curved surfaces, supporting both inner and outer marking.
| Model | LSR
Professional Rotary
Rotary attachment for marking and cutting rings, bracelets, and curved surfaces, supporting inner and outer marking.
| Model | LPR
Cutting Zig
Automatic plate feeder designed for serial marking and cutting applications.
| Model | LCZ
Eye Protection Cover
Protective cover that helps provide eye protection, dust prevention, and noise reduction.
| Model | LEP
Dust Collector
Extraction system for dust recovery and smoke filtering during laser operation.
| Model | LDC
External Water Chiller for Forza 160 / 180
Provides efficient cooling to maintain stable temperature, improve laser performance, and support continuous operation.
| Model | LWC
Internal Wrist Arm
Provides ergonomic arm support for comfortable operation and reduced operator fatigue.
| Model | LWA
Motorised Table
Height-adjustable table for improved working comfort and precise positioning.
| Model | LMT

## Laser Welding
Laser Spot Welding Machine Stout 250
Stout is a high-performance laser spot welding system designed for precision joining of fine jewellery and micro components. It combines high pulse energy, stable beam control, and real-time visual monitoring to deliver consistent, high-quality welds with minimal thermal impact.
| Laser Source | Nd:Ce:YAG
| Wavelength | 1064 nm
| Max Pulse Energy | 250 J
Spacious Chamber Access | Innovative Smoother Spot | Double Cooling System | Precision Targeting Microscope | Capacitive Touch Display | 100 Program Memory
Stout features a large welding chamber with a wide 180° opening for easy handling of larger components. It includes joystick control for dynamic parameter adjustment, integrated LED lighting for clear visibility, and gas/argon provision, ensuring flexible, precise, and comfortable operation for intricate jewellery work. | The advanced optical design ensures an even and stable energy spot, resulting in cleaner weld surfaces and better distribution of filler material. This helps achieve smoother finishing and stronger, more refined weld joints. | Advanced dual cooling design efficiently manages heat during welding, ensuring stable performance and continuous operation across varying working conditions. | Leica 10X stereo microscope provides advanced optics with clear and sharp visualization, enabling accurate micro-welding with reduced eye strain and improved productivity. | Equipped with a vibrant 8-inch colour capacitive touch screen, the system offers a smooth and intuitive user experience. The highly responsive display makes navigation easy and adjustments quick, ensuring comfortable and seamless operation. | Allows storage of up to 100 welding programs with customizable parameters, eliminating the need to manually remember or adjust settings for different metals.
Specifications | Stout 180 | Stout 225 | Stout 250
Laser Source | Nd:Ce:YAG | Nd:Ce:YAG | Moto Nd:Ce:YAG
Wavelength | 1064 nm | 1064 nm | 1064 nm
Maximum Laser Energy | 180 J | 225 J | 250 J
Peak Power | 9 kW | 9 kW | 10 kW
Average Power | 90 W | 100 W | 100 W
Pulse Duration | 0.1 – 20 ms | 0.1 – 25 ms | 0.1 – 25 ms
Frequency of Repetition | 0.5 to 30 Hz | 0.5 to 30 Hz | 0.5 to 30 Hz
Spot Setting | Motorized | Motorized | Motorized
Spot Diameter | 0.2 to 2 mm (0.1 mm optional) | 0.2 to 2 mm (0.1 mm optional) | 0.2 to 2 mm (0.1 mm optional)
Microscope | 10× Leica Stereo Microscope | 10× Leica Stereo Microscope | 10× Leica Stereo Microscope
Display Theme | Light / Dark | Light / Dark | Light / Dark
Wave Program | 6, Preset Pulse Shapes | 6, Preset Pulse Shapes | 6, Preset Pulse Shapes
Program Memory | 100 | 100 | 100
Internal Control | Joystick | Joystick | Joystick
External Control / Display | 8" Colour TFT Touch Screen | 8" Colour TFT Touch Screen | 8" Colour TFT Touch Screen
Argon / Air Nozzle | Yes, Double Nozzle | Yes, Double Nozzle | Yes, Double Nozzle
Compressed Air Nozzle | Yes | Yes | Yes
Pre/Post Mode for Valve | Yes | Yes | Yes
Cooling System | Double Cooling System 4L | Double Cooling System 4L | Double Cooling System 4L
Provision for External Chiller | Yes | Yes | Yes
Camera | In-screen Camera | In-screen Camera | In-screen Camera
HDMI (Output) | Yes | Yes | Yes
Input Power | 230V AC, 50/60 Hz, Single Phase (110V on request) | 230V AC, 50/60 Hz, Single Phase (110V on request) | 230V AC, 50/60 Hz, Single Phase (110V on request)
Dimension (W x D X H) | 850 × 500 × 590 mm | 850 × 500 × 590 mm | 850 × 500 × 590 mm
Packaging dimensions (W x D X H) | 650 x 1040 x 840 mm | 650 x 1040 x 840 mm | 650 x 1040 x 840 mm
Net weight | 52 kg | 52 kg | 52 kg
Gross weight | 84 Kg | 84 Kg | 84 Kg
Laser Spot Welding Machine Stout 250-3D
Stout 3D is a high-performance laser spot welding system equipped with an advanced 3D vision system for precise depth perception and accurate targeting. It combines high pulse energy, stable beam control, and enhanced real-time visual monitoring to deliver precise, consistent, and high-quality welds with minimal thermal impact.
| Laser Source | Moto Nd:Ce:YAG
| Wavelength | 1064 nm
| Max Pulse Energy | 250 J
Spacious Chamber Access | Innovative Smoother Spot | Double Cooling System | 3D Vision Sytem | Capacitive Touch Display | 100 Program Memory | Sanitization Safety
Stout-3D features a large welding chamber with a wide 180° opening for easy handling of larger components. It includes joystick control for dynamic parameter adjustment, integrated LED lighting for clear visibility, and gas/argon provision, ensuring flexible, precise, and comfortable operation for intricate jewellery work. | The advanced optical design ensures an even and stable energy spot, resulting in cleaner weld surfaces and better distribution of filler material. This helps achieve smoother finishing and stronger, more refined weld joints. | Advanced dual cooling design efficiently manages heat during welding, ensuring stable performance and continuous operation across varying working conditions. | Advanced 3D stereoscopic viewing ensures clear, strain-free visibility for precise welding on complex surfaces, improving productivity and allowing comfortable long-duration operation, even while wearing spectacles. | Equipped with a vibrant 8-inch colour capacitive touch screen, the system offers a smooth and intuitive user experience. The highly responsive display makes navigation easy and adjustments quick, ensuring comfortable and seamless operation. | Allows storage of up to 100 welding programs with customizable parameters, eliminating the need to manually remember or adjust settings for different metals. | Ensures a clean and safe working environment with an integrated dust filter that removes fumes, dust, and odors during operation.
Specifications | Stout 180-3D | Stout 225-3D | Stout 250-3D
Laser Source | Nd:Ce:YAG | Nd:Ce:YAG | Moto Nd:Ce:YAG
Wavelength | 1064 nm | 1064 nm | 1064 nm
Maximum Laser Energy | 180 J | 225 J | 250 J
Peak Power | 9 kW | 9 kW | 10 kW
Average Power | 90 W | 100 W | 100 W
Pulse Duration | 0.1 – 20 ms | 0.1 – 25 ms | 0.1 – 25 ms
Frequency of Repetition | 0.5 to 30 Hz | 0.5 to 30 Hz | 0.5 to 30 Hz
Spot Setting | Motorized | Motorized | Motorized
Spot Diameter | 0.2 to 2 mm (0.1 mm optional) | 0.2 to 2 mm (0.1 mm optional) | 0.2 to 2 mm (0.1 mm optional)
Microscope | Lynx Stereo 3D Vision | Lynx Stereo 3D Vision | Lynx Stereo 3D Vision
Display Theme | Light / Dark | Light / Dark | Light / Dark
Wave Program | 6, Preset Pulse Shapes | 6, Preset Pulse Shapes | 6, Preset Pulse Shapes
Program Memory | 100 | 100 | 100
Internal Control | Joystick | Joystick | Joystick
External Control / Display | 8" Colour TFT Touch Screen | 8" Colour TFT Touch Screen | 8" Colour TFT Touch Screen
Argon / Air Nozzle | Yes, Double Nozzle | Yes, Double Nozzle | Yes, Double Nozzle
Compressed Air Nozzle | Yes | Yes | Yes
Pre/Post Mode for Valve | Yes | Yes | Yes
Cooling System | Double Cooling System 4L | Double Cooling System 4L | Double Cooling System 4L
Provision for External Chiller | Yes | Yes | Yes
Camera | In-screen Camera | In-screen Camera | In-screen Camera
HDMI (Output) | Yes | Yes | Yes
Input Power | 230V AC, 50/60 Hz, Single Phase (110V on request) | 230V AC, 50/60 Hz, Single Phase (110V on request) | 230V AC, 50/60 Hz, Single Phase (110V on request)
Dimension (W x D X H) | 850 × 500 × 590 mm | 850 × 500 × 590 mm | 850 × 500 × 590 mm
Packaging dimensions (W x D X H) | 650 x 1040 x 840 mm | 650 x 1040 x 840 mm | 650 x 1040 x 840 mm
Net weight | 55 kg | 55 kg | 55 kg
Gross weight | 84 Kg | 84 Kg | 84 Kg
Economical Laser Spot Welding Machine Forza 160
Compact benchtop design for easy mobility and space-saving operation. Delivers high productivity with fast, noiseless welding and reduced processing time. The integrated microscope and camera system with ergonomic binocular optics ensures precise, fatigue-free operation. An efficient air-cooled system provides effective heat dissipation, low maintenance, and consistent high-quality welds without the need for secondary processing.
| Laser Source | Nd:Ce:YAG
| Wavelength | 1064 nm
| Max Pulse Energy | 160 J
Spacious Chamber Access | Innovative Smoother Spot | Double Cooling System | Precision Targeting Microscope | Capacitive Touch Display | 100 Program Memory
Forza features a large welding chamber with a wide 180° opening for easy handling of larger components. It includes joystick control for dynamic parameter adjustment, integrated LED lighting for clear visibility, and gas/argon provision, ensuring flexible, precise, and comfortable operation for intricate jewellery work. | The advanced optical design ensures an even and stable energy spot, resulting in cleaner weld surfaces and better distribution of filler material. This helps achieve smoother finishing and stronger, more refined weld joints. | Advanced dual cooling design efficiently manages heat during welding, ensuring stable performance and continuous operation across varying working conditions. | Leica 10x stereo microscope provides advanced optics with clear and sharp visualization, enabling accurate micro-welding with reduced eye strain and improved productivity. | Equipped with a vibrant 8-inch colour capacitive touch screen, the system offers a smooth and intuitive user experience. The highly responsive display makes navigation easy and adjustments quick, ensuring comfortable and seamless operation. | Allows storage of up to 100 welding programs with customizable parameters, eliminating the need to manually remember or adjust settings for different metals.
Specifications | Forza 60 | Forza 100 | Forza 160
Laser Source | Nd:Ce:YAG | Nd:Ce:YAG | Nd:Ce:YAG
Wavelength | 1064 nm | 1064 nm | 1064 nm
Maximum Laser Energy | 60 J | 100 J | 160 J
Peak Power | 3 kW | 4 kW | 6.5 kW
Average Power | 30 W | 45 W | 75 W
Pulse Duration | 20 ms | 25 ms | 25 ms
Frequency of Repetition | 0.5 to 10 Hz | 0.5 to 10 Hz | 0.5 to 20 Hz
Spot Setting | Motorized | Motorized | Motorized
Spot Diameter | 0.2 to 2 mm | 0.2 to 2 mm | 0.2 to 2 mm
Microscope | 10x Aczet approved Stereo Microscope | 10x Aczet approved Stereo Microscope | 10x Aczet approved Stereo Microscope
Display Theme | Dark / Light | Dark / Light | Dark / Light
Wave Program | 6 Preset Pulse Shape | 6 Preset Pulse Shape | 6 Preset Pulse Shape
Program Memory | 100 | 100 | 100
Internal Control | Not Available | Not Available | Joystick
External Control / Display | 5" Colour TFT Touch Screen | 5" Colour TFT Touch Screen | 8" Colour TFT Touch Screen
Argon / Air Nozzle | Yes, Single Nozzle | Yes, Single Nozzle | Yes, Double Nozzle
Compressed Air Nozzle | No | No | Yes
Pre/Post Mode for Valve | Not Available | Not Available | Yes
Cooling System | Single Cooling System 3L | Single Cooling System 3L | Single Cooling System 3L
Provision for External Chiller | Yes | Yes | Yes
Camera | No Camera | No Camera | In-screen Camera
Input Power | 230V AC, 50/60 Hz, Single Phase (110V on request) | 230V AC, 50/60 Hz, Single Phase (110V on request) | 230V AC, 50/60 Hz, Single Phase (110V on request)
Dimension (W × D × H) | 516 × 671 × 582 mm | 516 × 671 × 582 mm | 516 × 671 × 582 mm
Packaging dimensions (W x D X H) | 900 × 670 × 860 mm | 1100 × 700 × 850 mm | 1040 × 650 × 850 mm
Net weight | 52 kg | 52 kg | 54 kg
Gross weight | 79 kg | 78 kg | 78 kg
Laser Spot Welding Machine CMLW Series
Laser Source
| Wavelength | 1064 nm
| Max Pulse Energy | 100 J
Spacious Chamber Access | Innovative Smoother Spot | Double Cooling System | Precision Targeting Microscope | Capacitive Touch Display | 100 Program Memory
Model | CMLW 200
Wavelength | 1064 nm
Maximum Laser Energy | 100 J
Peak Power | ≤ 5 kW
Average Power | 200 W
Pulse Duration
Frequency of Repetition | 1 – 20 Hz (Adjustable)
Spot Setting
Spot Diameter | 0.2 – 1.5 mm (Adjustable)
Microscope
Display Theme
Wave Program
Program Memory
Internal Control
External Control / Display
Argon / Air Nozzle
Compressed Air Nozzle
Pre/Post Mode for Valve
Cooling System
Provision for External Chiller | External chiller
Camera
Input Power | Single Phase AC 220V ±10%, 50/60 Hz
Dimension (W × D × H) | 790 × 390 × 1090 mm
Packaging dimensions (W x D X H)
Net weight | 85 kg
Gross weight

## Laser Marking
Laser Marking Machine - LM 60
Laser marking machines deliver high-quality cutting and engraving for precise marking of logos, text, and graphics on metal surfaces in minimal time. With versatile power configurations, they support both marking and foil cutting,ensuring precision, flexibility, and higher production efficiency.
| Laser Source | Pulsed Powerful Fiber Laser
| Laser Power | 60 W
| Wavelength | 1064 nm
Simple and Versatile Software | Motorized Z-Axis | Precision and Accuracy | Quality Components | Air-Cooled Fiber Laser
Intuitive and easy-to-use software allows smooth operation while supporting engraving of logos, text, images, and vector files on various metals, with compatibility for formats like SVG, DXF, BMP, JPG, and more. | The motorized Z-axis helps you quickly and accurately adjust focus, making the overall operation easier and more efficient. | Delivers high precision and accuracy with a marking area of up to 160 × 160 mm and resolution up to 16 μm, making it perfect for fine details, intricate designs, and small characters. | Built with high-quality components and a well-designed system to deliver clean and consistent engraving, even for complex and detailed designs. | Built with a pulse fiber laser source with no moving parts or mirrors, ensuring maintenance-free operation, high efficiency, and reliable performance without the need for an external chiller.
Model | LM 20 | LM 60
Types | Marking / Engraver | Marking / Engraver / Cutting
Laser Source | Pulsed Powerful Fiber Laser | Pulsed Powerful Fiber Laser
Laser Power | 20 W | 60 W
Wavelength | 1064 nm | 1064 nm
Beam Quality M² | Excellent, M < 1.4 | Excellent, M < 1.4
Pulse Repetition Rate | 20 Hz-100 kHz | 20 Hz-100 kHz
Working Area | 100 × 100 mm | 100 × 100 mm
Working Life of Laser | >50,000 hours | >50,000 hours
Marking Speed | < 7000 mm/sec | < 7000 mm/sec
Laser Frequency | 20 - 100 kHz | 20 - 100 kHz
Axis | Auto Z Axis | Auto Z Axis
Cooling System | Air Cooling | Air Cooling
Minimum Width | 0.3 mm | 0.3 mm
Minimum Character | 0.3 mm | 0.3 mm
Repeatability | 0.003 mm | 0.003 mm
Power Consumption | 800 W | 800 W
Power Supply | 220V AC / 50Hz Single Phase (110V on request) | 220V AC / 50Hz Single Phase (110V on request)
Dimension (L × W × H) | 420 × 365 × 575 mm | 420 × 365 × 575 mm
Packing Dimension (L × W × H) | 580 x 470 x 550 mm | 580 x 470 x 550 mm
Net Weight (Approx.) | 19 kg | 19 kg
Gross Weight | 25 kg | 25 kg
Laser Marking Machine - CFM 100
Laser marking machines deliver high-quality cutting and engraving for precise marking of logos, text, and graphics on metal surfaces in minimal time. With versatile power configurations, they support both marking and foil cutting, ensuring precision, flexibility, and higher production efficiency.
| Laser Source | Ytterbium Pulsed Fiber Laser
| Laser Power | 100 W
| Wavelength | 1064 nm
Simple and Versatile Software | Motorized Z-Axis | Precision and Accuracy | Quality Components | Air-Cooled Fiber Laser
Intuitive and easy-to-use software allows smooth operation while supporting engraving of logos, text, images, and vector files on various metals, with compatibility for formats like SVG, DXF, BMP, JPG, and more. | The motorized Z-axis helps you quickly and accurately adjust focus, making the overall operation easier and more efficient. | Delivers high precision and accuracy with a marking area of up to 160 × 160 mm and resolution up to 16 μm, making it perfect for fine details, intricate designs, and small characters. | Built with high-quality components and a well-designed system to deliver clean and consistent engraving, even for complex and detailed designs. | Built with a pulse fiber laser source with no moving parts or mirrors, ensuring maintenance-free operation, high efficiency, and reliable performance without the need for an external chiller.
Model | CFM 20 | CFM 50 | CFM 60 | CFM 100
Types | Marking / Engraver | Marking / Engraver / Cutting | Marking / Engraver / Cutting | Marking / Engraver / Cutting
Laser Source | Ytterbium Pulsed Fiber Laser | Ytterbium Pulsed Fiber Laser | Ytterbium Pulsed Fiber Laser | Ytterbium Pulsed Fiber Laser
Laser Power | 20 W | 50 W | 60 W | 100 W
Wavelength | 1064 nm | 1064 nm | 1064 nm | 1064 nm
Beam Quality M² | < 1.6 | < 1.6 | < 1.6 | < 1.6
Pulse Repetition Rate | 20 Hz-100kHz | 20 Hz-100kHz | 20 Hz-100kHz | 20 Hz-100kHz
Working Area | 100 × 100 mm | 100 × 100 mm | 100 × 100 mm | 100 × 100 mm
Working Life of Laser | >50,000 hours | >50,000 hours | >50,000 hours | >50,000 hours
Marking Speed | < 7000 mm/sec | < 7000 mm/sec | < 7000 mm/sec | < 7000 mm/sec
Laser Frequency | 20 - 100 kHz | 20 - 100 kHz | 20 - 100 kHz | 20 - 100 kHz
Axis | Auto Z Axis | Auto Z Axis | Auto Z Axis | Auto Z Axis
Cooling System | Air Cooling | Air Cooling | Air Cooling | Air Cooling
Minimum Width | 0.3 mm | 0.3 mm | 0.3 mm | 0.3 mm
Minimum Character | 0.3 mm | 0.3 mm | 0.3 mm | 0.3 mm
Repeatability | 0.003 mm | 0.003 mm | 0.003 mm | 0.003 mm
Power Consumption | 800 W | 800 W | 800 W | 800 W
Power Supply | 220V AC / 60Hz / 1P /4A or 110V / 50Hz / 1P / 8A (110V on request) | 220V AC / 60Hz / 1P /4A or 110V / 50Hz / 1P / 8A (110V on request) | 220V AC / 60Hz / 1P /4A or 110V / 50Hz / 1P / 8A (110V on request) | 220V AC / 60Hz / 1P /4A or 110V / 50Hz / 1P / 8A (110V on request)
Dimension (L × W × H) | 635 × 400 × 660 mm | 635 × 400 × 660 mm | 635 × 400 × 660 mm | 735 × 400 × 660 mm
Packing Dimension (L × W × H) | 965 × 610 × 914 mm | 965 × 610 × 914 mm | 965 × 610 × 914 mm | 1016 × 610 × 914 mm
Net Weight (Approx.) | 40 kg | 40 kg | 40 kg | 48 kg
Weight | 45 kg | 45 kg | 45 kg | 45 kg
Laser Marking Machine - CFEP 10
High-speed laser marking and engraving for jewellery and precious metals—create sharp logos, hallmarks, text, and graphics with faster production. Easy software supports images/vectors and common formats (SVG, DXF, BMP, PLT, JPG, DWG). Motorized Z-axis focusing, up to 160×160 mm area, 16 μm resolution. Air-cooled pulse fiber laser: low maintenance, no external chiller.
| Laser Source | Diode End Pump
| Laser Power | 10 W
| Wavelength | 1064 ±1 nm
Simple and Versatile Software | Motorized Z-Axis | Precision and Accuracy | Quality Components | Diode-Pumped Laser System
Intuitive and easy-to-use software allows smooth operation while supporting engraving of logos, text, images, and vector files on various metals, with compatibility for formats like SVG, DXF, BMP, JPG, and more. | The motorized Z-axis helps you quickly and accurately adjust focus, making the overall operation easier and more efficient. | Delivers high precision and accuracy with a marking area of up to 160 × 160 mm and resolution up to 16 μm, making it perfect for fine details, intricate designs, and small characters. | Built with high-quality components and a well-designed system to deliver clean and consistent engraving, even for complex and detailed designs. | Built with a pulse diode-pumped laser source that has no moving parts or mirrors, ensuring long-lasting, maintenance-free operation, high efficiency, and consistent performance, all without requiring an external chiller.
Model | CLEP 10
Types | Marking / Engraver
Laser Source | Diode End Pump
Laser Power | 10 W
Wavelength | 1064 nm ± 1nm
Beam Quality M² | < 1.2
Pulse Repetition Rate | 25 Hz
Working Area | 100 × 100 mm
Working Life of Laser | <20,000 hours
Marking Speed | < 7000 mm/sec
Laser Frequency | 20 - 100 kHz
Axis | Auto Z Axis
Cooling System | Air Cooling
Minimum Width | 0.3 mm
Minimum Character | 0.3 mm
Repeatability | 0.003 mm
Power Consumption | 800 W
Power Supply | 220V AC / 60Hz / 1P /4A or 110V / 50Hz / 1P / 8A (110V on request)
Dimension (L × W × H) | 580 × 460 × 640 mm
Packing Dimension (L × W × H) | 933 × 588 × 840 mm
Net Weight (Approx.) | 55 kg
Weight | 45 kg
Multicolour Laser Engraver
Multicolour Laser Engraveing Machine is designed for precision marking, engraving, and color processing on gold jewellery. It enables high-quality detailing with vibrant color effects, ensuring improved craftsmanship, flexibility, and efficient production.
Laser Source
Laser Power
Wavelength
User-Friendly Operation | Compact & Powerful Design | Multi-Function Capability | Vivid Color Processing | Enhanced Craftsmanship
Designed for easy use, enabling smooth operation for all skill levels while ensuring consistent and reliable results on every piece. | Lightweight and compact design enables easy portability without compromising performance, making it ideal for both small workshops and large-scale manufacturing environments. | Supports marking, cutting, and vibrant color processing on various materials, including gold (14K, 18K) as well as stainless steel and brass. | Delivers high-precision color marking to create vibrant and detailed designs, ideal for enhancing the visual appeal of gold jewellery. | Advanced laser technology improves design quality and aesthetics, supporting intricate patterns and high-value jewellery creation.
Model
Types | Multicolour Marking
Laser Source
Laser Power
Wavelength
Beam Quality M²
Pulse Repetition Rate
Working Area (mm)
Working Life of Laser
Marking Speed
Laser Frequency
Axis
Cooling System
Minimum Width
Minimum Character
Repeatability
Power Consumption
Power Supply
Dimension (L × W × H)
Net Weight (Approx.)
Packing Dimension (L × W × H)
Weight
Laser Marking Machine - CFM 3D
CFM 3D laser marking machines deliver high-quality cutting and engraving for precise marking of logos, text, and graphics on metal surfaces in minimal time. With versatile power configurations, they support both marking and foil cutting, ensuring precision, flexibility, and higher production efficiency.
| Laser Source | Ytterbium Pulsed Fiber Laser
| Laser Power | 100 W
| Wavelength | 1064 nm
Simple and Versatile Software | Motorized Z-Axis | Precision and Accuracy | Quality Components | Air-Cooled Fiber Laser
Intuitive and easy-to-use software allows smooth operation while supporting engraving of logos, text, images, and vector files on various metals, with compatibility for formats like SVG, DXF, BMP, JPG, and more. | The motorized Z-axis helps you quickly and accurately adjust focus, making the overall operation easier and more efficient. | Delivers high precision and accuracy with a marking area of up to 160 × 160 mm and resolution up to 16 μm, making it perfect for fine details, intricate designs, and small characters. | Built with high-quality components and a well-designed system to deliver clean and consistent engraving, even for complex and detailed designs. | Built with a pulse fiber laser source with no moving parts or mirrors, ensuring maintenance-free operation, high efficiency, and reliable performance without the need for an external chiller.
Model | CFM 30-3D | CFM 50-3D | CFM 100-3D
Types | 3D Marking / Marking / Engraver /Cutting | 3D Marking / Marking / Engraver /Cutting | 3D Marking / Marking / Engraver /Cutting
Laser Source | Ytterbium Pulsed Fiber Laser | Ytterbium Pulsed Fiber Laser | Ytterbium Pulsed Fiber Laser
Laser Power | 30 W | 50 W | 100 W
Wavelength | 1064 nm | 1064 nm | 1064 nm
Beam Quality M² | < 1.6 | < 1.6 | < 1.6
Pulse Repetition Rate | 20 Hz - 100 kHz | 20 Hz - 100 kHz | 20 Hz - 100 kHz
Working Area | 100 × 100 mm | 100 × 100 mm | 100 × 100 mm
Working Life of Laser | >50,000 hours | >50,000 hours | >50,000 hours
Marking Speed | < 7000 mm/sec | < 7000 mm/sec | < 7000 mm/sec
Laser Frequency | 20 - 100 kHz | 20 - 100 kHz | 20 - 100 kHz
Axis | Auto Z Axis | Auto Z Axis | Auto Z Axis
Cooling System | Air Cooling | Air Cooling | Air Cooling
Minimum Width | 0.3 mm | 0.3 mm | 0.3 mm
Minimum Character | 0.3 mm | 0.3 mm | 0.3 mm
Repeatability | 0.003 mm | 0.003 mm | 0.003 mm
Power Consumption | 800 W | 800 W | 800 W
Power Supply | 220V AC / 60Hz / 1P /4A or 110V / 50Hz / 1P / 8A (110V on request) | 220V AC / 60Hz / 1P /4A or 110V / 50Hz / 1P / 8A (110V on request) | 220V AC / 60Hz / 1P /4A or 110V / 50Hz / 1P / 8A (110V on request)
Dimension (L × W × H) | 635 × 400 × 660 mm | 635 × 400 × 660 mm | 635 × 400 × 660 mm
Packing Dimension (L × W × H) | 1040 x 600 x 920 mm | 1040 x 600 x 920 mm | 1040 x 600 x 920 mm
Net Weight (Approx.) | 70 kg | 70 kg | 70 kg
Weight | 90 kg | 90 kg | 90 kg

## G&J Balances
Hallmarking balances |  | Hallmarking Microbalance AXR 2MA
Delivers precise weighing for the smallest jewelry samples. Engineered for uncompromising accuracy in hallmarking environments, this next-generation weighing technology is designed exclusively for hallmarking applications, offering durable construction and long-term reliability. High-precision performance with a 2.1 g capacity and 1 µg readability ensures accurate micro-weighing of precious metals. An intuitive color TFT touchscreen with internal calibration and leveling alerts enables fast, error-free operation. 21 CFR Part 11–compliant user management ensures secure, regulated workflows, while multiple connectivity options (USB, RS232, Ethernet) support seamless data recording and traceability. Compact, overload-protected, and power-efficient—ideal for modern gold and precious metal hallmarking laboratories.
|  | Weighing Capacity | 2.1 g
|  | Readability | 1 µg
|  | Minimum Weight (USP, 0.1%, typical) | 1 mg
| Excellent Weighing Performance | Exceptional Accuracy for small samples | New Touchscreen as Smartphone | 21 CFR Part 11 Compliance | Minimum Sample Weight | Diversity | Digital Level | Audit Safe Result | 10 Easy-to-use Built-in-Applications | Data Free Transfer
| The Microbalance, with its two-terminal concept, offers excellent weighing performance by separating the electronics from the precise mechanical measuring cell. This design helps to eliminate the effects of temperature and its influence on the balance's performance. | With one microgram readability, low minimum weight, and an improved repeatability by 20 percent, you can accurately weigh small samples with the utmost accuracy. This is one of the biggest merits of this machine. | It’s 5.7" large colour TFT display and the capacitive touchscreen allows you to effortless swipe through operations, configurations, Tasks and Users. It has an Intuitive interface which is operable through cotton, silicon or rubber gloves. | Our machine fully meets the requirements of 21 CFR Part 11/ EU GMP Annex 11. It also provides features like Full data security, Electronic signatures, Validations, Electronic records and Audit trails. | The balance displays whether the value of the weight is below/above the minimum weight threshold. 3 MSW method can be configure e.g. USP, GMP and User defined with and Without Tare Weight. | The tasks can be custom-made for each small unique configuration of parameter and make it quick, errorless operation. This saves you time and provides higher productivity. It also has customised print format. | Instead of manually keep track of your level your scale. It automatically alerts before performance can be compromised. Digital level saves your time and gives you maximum precision in your measurements. | Innovative audit safe with alibi and audit trial. As you carry out your operation, your results and changed parameters are automatically recorded on Alibi and Audi trial makes it operation complete audit safe. | The ten easy to use build in applications include density determination, check weighing, totaling, and formulation, simplify your routine weighing tasks. These applications make your analytical operations much easier and accurate. | Simply transfer task parameters and results to a PC via many interface ports like RS232, USB A, USB B, Ethernet and wifi. Transcription errors are completely eliminated and reliability of your data is preserved.
Specifications - Microbalance AXR 2MA
| Maximum Capacity | 2.1 g
| Minimum load (OIML) | 0.1 mg
| Readability | 1 µg
| Verification unit [e] | 1 mg
| Repeatability | 1 µg
| Linearity | 3 µg
| Tare Range (Subtractive) | -2.1 g
| Standard repeatability [5% Max] | 0.5 µg
| Minimum Weight USP | 1 mg
| Minimum weight ((U=1%, k=2) | 0.1 mg
| Weighing Pan Diameter | Ø 26 mm
| Setting Time | 8-10 Sec
| Display | 5.7'' large colour TFT display
| Calibration | Motorised Internal Calibration (automatic/PSAC)
| User Management | Log History [21 CFR Part 11 Compliant ] 7 User management Password protection AuditTrai Minimum Sample Weight (MSW)
| Features | Touch-free operation Infrared Proximity Sensor [Automatic Door] 4 built-in languages English, German, French, Spanish Levelling guidance
| Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
| Operating temperature | 15º to 35ºC
| Protection class | IP 43
| Interfaces | RS232/USB
| Power Supply | DC Adapter, input 100 ~ 240V 0.5A output 13V / 0.9A
Dimensions for Electronic Box (W x D x H)
Dimensions for Mechanical Box (W x D x H)
Net weight
Gross weight
Hallmarking balances |  | Hallmarking Microbalance CM 2
Delivers precise weighing for the smallest jewelry samples, featuring 2.1 g capacity, 1 µg readability, bright LCD display, internal calibration, USB/RS232 connectivity, overload protection, ideal for hallmarking laboratories worldwide use.
|  | Weighing Capacity | 2.1 g
|  | Readability | 1 µg
|  | Minimum Weight (USP, 0.1%, typical) | 1 mg
| Excellent Weighing Performance | Brilliant Proven Technology | Automatic Internal Calibration (PSAC) | Touch Free Draftshield Operational | Bright, High-Contrast Display | 10 Easy-to-Use Built-In Applications | ISO/GLP Documentation | Seamless Connectivity | Quick & Simple Leveling | Easy Maintenance & Cleaning
| The Microbalance, with its two-terminal concept, offers excellent weighing performance by separating the electronics from the precise mechanical measuring cell. This design helps to eliminate the effects of temperature and its influence on the balance's performance. | The sensor truly stands out with its unparalleled design based on the electromagnetic force compensation (EMFC). It provides an exceptional performance that meets the highest of the set international standards. | Motorised, built-in weight internal calibration ensures precise weighing under all conditions. The Perfect Self Automatic Calibration (PSAC) function automatically maintains accuracy by adjusting for time and temperature changes. | The draft shield has wide opening, clear visibility and Infrared sensors to open and close the draft-shield doors with the wave of a hand. It simplifies dosing, minimizes spills, and reduces cross-contamination. | Excellent readable display with bright LED-backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | The built-in real-time clock ensures accurate, time-stamped data, providing traceable documentation that reflects true test results and supports compliance with SOPs and GLP regulations. | The USB and RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data. | An easy-to-view circular bubble level and large adjustable feet make balance leveling quick, simple, and precise for stable, accurate weighing. | Washable parts dismantle in seconds without tools. Dishwasher-safe components and a smooth housing design make routine cleaning quick, simple, and hygienic for everyday laboratory use.
Specifications - Hallmarking Microbalance CM 2
| Maximum Capacity | 2.1 g
| Minimum load (OIML) | 0.1 mg
| Readability | 1 µg
| Verification unit [e] | 1 mg
| Repeatability | 1 µg
| Linearity | 3 µg
| Tare Range (Subtractive) | -2.1 g
| Standard repeatability [5% Max] | 0.5 µg
| Minimum Weight USP | 1 mg
| Minimum weight ((U=1%, k=2) | 0.1 mg
| Weighing Pan Diameter | Ø 30 mm
| Setting Time | 8-10 Sec
| Display | LCD Backlit
| Calibration | Motorised Internal Calibration (automatic/PSAC)
| Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
| Operating temperature | 15º to 35ºC
| Interfaces | RS232/USB
| Power Supply | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A
| Weighing chamber height | Ø75 x 86.5 mm
| Dimensions for Electronic Box (W x D x H) | 240 x 210 x 65 mm
| Dimensions for Mechanical Box (W x D x H) | 360 x 140 x 146 mm
Packaging dimensions (W x D X H)
Net weight
Gross weight
Hallmarking balances |  | Hallmarking Analytical CY 285C
Delivers precise weighing for the smallest jewelry samples. Intuitive operation with solid construction, 80/220 g capacity, 0.01/0.1 mg readability, bright LCD, internal calibration, USB/RS232, overload protection for hallmarking labs.
|  | Weighing Capacity | 80/220 g
|  | Readability | 0.01/0.1 mg
|  | Minimum Weight (USP, 0.1%, typical) | 30 mg
| Brilliant Proven Technology | Automatic Internal Calibration (PSAC) | Revolution in Mass Measurement Quantity | Bright, High-Contrast Display | Automatic Internal Calibration (PSAC) | 10 Easy-to-Use Built-In Applications | ISO/GLP Documentation | Seamless Connectivity | Easy Maintenance & Cleaning | Quick & Simple Leveling
| The sensor really stands out because of its unique design based on electromagnetic force compensation (EMFC). It works very well and meets the highest international standards. | Motorised, built-in weight internal calibration ensures precise weighing under all conditions. The Perfect Self Automatic Calibration (PSAC) function automatically maintains accuracy by adjusting for time and temperature changes. | Our weighing system sets a new standard for the quality of mass measurement by meeting strict USP requirements (chapters 41 and 1251) for the best weighing accuracy and sd ≤ 1d repeatability. | Excellent readable display with bright LED-backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | Motorised, built-in weight internal calibration ensures precise weighing under all conditions. The Perfect Self Automatic Calibration (PSAC) function automatically maintains accuracy by adjusting for time and temperature changes. | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | The built-in real-time clock ensures accurate, time-stamped data, providing traceable documentation that reflects true test results and supports compliance with SOPs and GLP regulations. | The USB and RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data. | Washable parts dismantle in seconds without tools. Dishwasher-safe components and a smooth housing design make routine cleaning quick, simple, and hygienic for everyday laboratory use. | An easy-to-view circular bubble level and large adjustable feet make balance leveling quick, simple, and precise for stable, accurate weighing.
Specifications - Analytical Balance CY 285C
| Maximum Capacity | 82 / 220 g
| Readability | 0.01 / 0.1 mg
| Repeatability | 0.02 / 0.1 mg
| Weighing Pan Diameter | Ø 80 mm Grid pan
| Setting Time | 5- 6 Sec
| Display | LCD Backlit
| Calibration | Motorised Internal Calibration (automatic/PSAC)
| Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
| Interfaces | RS232/USB
| Power Supply | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A
| Balance Dimension (W x D X H) | 342.5 x 212 x 350 mm
| Packaging dimensions (W x D X H) | 490 × 400 × 520 mm
| Net weight | 8 kg
| Gross weight | 12 kg
Hallmarking balances |  | Hallmarking Analytical CY 224C
Delivers precise weighing for the smallest jewelry samples. Intuitive solid construction with 220 g capacity, 0.1 mg readability, bright LCD, internal calibration, USB/RS232 connectivity, overload protection—ideal for hallmarking laboratories.
|  | Weighing Capacity | 220 g
|  | Readability | 0.1 mg
|  | Minimum Weight (USP, 0.1%, typical) | 160 mg
| Brilliant Proven Technology | Solid & Reliable Design | Bright, High-Contrast Display | Seamless Connectivity | 10 Easy-to-Use Built-In Applications | ISO/GLP Documentation | Quick & Simple Leveling | Easy Maintenance & Cleaning
| The sensor truly stands out with its unparalleled design based on the electromagnetic force compensation (EMFC). It provides an exceptional performance that meets the highest of the set international standards. | Built with a solid design and high-quality materials for long-term durability, ensuring reliable and consistent operation with accurate results. Supplied with a standard protective cover for added safety. | Excellent readable display with bright LED-backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | The USB and RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data. | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | The built-in real-time clock ensures accurate, time-stamped data, providing traceable documentation that reflects true test results and supports compliance with SOPs and GLP regulations. | An easy-to-view circular bubble level and large adjustable feet make balance leveling quick, simple, and precise for stable, accurate weighing. | Washable parts dismantle in seconds without tools. Dishwasher-safe components and a smooth housing design make routine cleaning quick, simple, and hygienic for everyday laboratory use.
Specifications - Hallmarking CY 224C
| Maximum Capacity | 220 g
| Readability | 0.1 mg
| Repeatability | 0.1 mg
| Weighing Pan Diameter | Ø 90 mm
| Setting Time | 2- 3 Sec
| Display | LCD Backlit
| Calibration | Motorised Internal Calibration (automatic/PSAC)
| Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
| Interfaces | RS232/USB
| Power Supply | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A
| Balance Dimension (W x D X H) | 342.5 x 212 x 350 mm
| Packaging dimensions (W x D X H) | 490 × 400 × 520 mm
| Net weight | 8 kg
| Gross weight | 12 kg
Hallmarking balances |  | Hallmarking Precision Balance CY 3102
Delivers precise weighing for the smallest jewelry samples. Intuitive solid construction with 3100 g capacity, 0.01 g readability, bright LCD display, external calibration, USB/RS232 connectivity, overload protection—ideal for modern hallmarking laboratories.
|  | Weighing Capacity | 3100 g
|  | Readability | 0.01 g
|  | Minimum Weight (USP, 0.1%, typical) | 10 g
| Brilliant Proven Technology | Solid & Reliable Design | Bright, High-Contrast Display | 10 Easy-to-Use Built-In Applications | ISO/GLP Documentation | Quick & Simple Leveling | Easy Maintenance & Cleaning
| The Precision Balances use a MonoTech single-block weighing sensor ensuring uniform load transmission, thermal stability, rapid stabilization, accurate corner-load performance, high repeatability, minimal drift, and long-term reliability in continuous operation. | Built with a solid design and high-quality materials for long-term durability, ensuring reliable and consistent operation with accurate results. Supplied with a standard protective cover for added safety. | Excellent readable display with bright LED-backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | The built-in real-time clock ensures accurate, time-stamped data, providing traceable documentation that reflects true test results and supports compliance with SOPs and GLP regulations. | An easy-to-view circular bubble level and large adjustable feet make balance leveling quick, simple, and precise for stable, accurate weighing. | Washable parts dismantle in seconds without tools. Dishwasher-safe components and a smooth housing design make routine cleaning quick, simple, and hygienic for everyday laboratory use.
Specifications - Hallmarking Precision Balance CY 3102
| Maximum Capacity | 3100 g
| Readability | 0.01 g
| Repeatability | 0.01 g
| Weighing Pan | 180 x 195 mm
| Setting Time | 2- 3 Sec
| Display | LCD Backlit
| Calibration | External Calibration
| Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
| Operating temperature | 15º to 35ºC
| Relative humidity | 40% ÷ 80%
| Sensitivity Drift | ± 2ppm (2 x 10-6 / ºC)
| Protection class | IP 43
| Leveling system | manual
| Interfaces | RS-232/USB
| Power Supply | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A
| Compliance | ISO, GLP, GMP, USP
| Balance Dimension (W x D X H) | 342.5 x 212 x 89.5 mm
| Packaging dimensions (W x D X H) | 490 × 400 × 520 mm
| Net weight | 7.5 kg
| Gross weight | 12 kg
Diamond Scales |  | Diamond Scales CY 603KC
Trusted precision for accurate jewelry weighing. Offers 600 ct capacity with 0.001 ct readability. Features a large, bright LCD backlit display, internal calibration, easy leveling, USB and RS232 connectivity, overload protection, and a compact footprint.
|  | Weighing Capacity | 600 ct
|  | Readability | 0.001 ct
|  | Calibration | Motorised Internal Calibration (automatic/PSAC)
| Brilliant Proven Technology | Solid & Reliable Design | Bright, High-Contrast Display | Seamless Connectivity | 10 Easy-to-Use Built-In Applications | ISO/GLP Documentation
| The sensor truly stands out with its unparalleled design based on the electromagnetic force compensation (EMFC). It provides an exceptional performance that meets the highest of the set international standards. | Built with a solid design and high-quality materials for long-term durability, ensuring reliable and consistent operation with accurate results. Supplied with a standard protective cover for added safety. | Excellent readable display with bright LED-backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | The USB and RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data. | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | The built-in real-time clock ensures accurate, time-stamped data, providing traceable documentation that reflects true test results and supports compliance with SOPs and GLP regulations.
|  | Specifications - Diamond Scales CY 603KC | Specifications - Diamond Scales CY 1603KC
| Maximum Capacity | 600 ct | 1600 ct
| Readability | 0.001 ct | 0.001 ct
| Repeatability | 0.001 ct | 0.001 ct
| Weighing Pan Diameter | Ø 90 mm | Ø 90 mm
| Setting Time | 2- 3 Sec | 2- 3 Sec
| Display | LCD Backlit | LCD Backlit
| Calibration | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC)
| Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
| Interfaces | RS232/USB | RS232/USB
| Power Supply | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A
| Balance Dimension (W x D X H) | 342.5 × 212 × 279 mm | 342.5 × 212 × 279 mm
| Packaging dimensions (W x D X H) | 540 x 430 x 530 mm | 540 x 430 x 530 mm
Net weight
Gross weight
Diamond Scales CY 603K
Trusted precision for accurate jewelry weighing. Offers 600 ct capacity with 0.001 ct readability. Features a large, bright LCD backlit display, external calibration, easy leveling, USB and RS232 connectivity, overload protection, and a compact footprint.
|  | Weighing Capacity | 600 ct
|  | Readability | 0.001 ct
|  | Calibration | External Calibration
| Brilliant Proven Technology | Solid & Reliable Design | Bright, High-Contrast Display | Seamless Connectivity | 10 Easy-to-Use Built-In Applications | ISO/GLP Documentation
| The sensor truly stands out with its unparalleled design based on the electromagnetic force compensation (EMFC). It provides an exceptional performance that meets the highest of the set international standards. | Built with a solid design and high-quality materials for long-term durability, ensuring reliable and consistent operation with accurate results. Supplied with a standard protective cover for added safety. | Excellent readable display with bright LED-backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | The USB and RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data. | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | The built-in real-time clock ensures accurate, time-stamped data, providing traceable documentation that reflects true test results and supports compliance with SOPs and GLP regulations.
|  | Specifications - Diamond Scales CY 253K | Specifications - Diamond Scales CY 603K
| Maximum Capacity | 250 ct | 600 ct
| Readability | 0.001 ct | 0.001 ct
| Repeatability | 0.001 ct | 0.001 ct
| Weighing Pan Diameter | Ø 90 mm | Ø 90 mm
| Setting Time | 2- 3 Sec | 2- 3 Sec
| Display | LCD Backlit | LCD Backlit
| Calibration | External Calibration | External Calibration
| Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
| Interfaces | RS232/USB | RS232/USB
| Power Supply | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A
| Balance Dimension (W x D X H) | 342.5 × 212 × 279 mm | 342.5 × 212 × 279 mm
| Packaging dimensions (W x D X H) | 540 x 430 x 530 mm | 540 x 430 x 530 mm
Net weight
Gross weight
Diamond Scales CG 103K
Trusted precision for accurate jewelry weighing. Offers a 100 ct capacity with 0.001 ct readability, ensuring reliable and cost-effective measurements. Its compact design, inbuilt rechargeable battery, and bright LCD backlit display enable easy use anywhere—even during power outages.
|  | Weighing Capacity | 100 ct
|  | Readability | 0.001 ct
|  | Calibration | External Calibration
| Precision Load Cell Technology | Solid & Reliable Design | Bright, High-Contrast Display | Power Efficiency
| Equipped with a precision aluminum load cell (IP65) and advanced 8-bit microprocessor design for accurate and stable performance. The level indicator with adjustable feet ensures proper setup, while 150% overload protection enhances safety and reliability. | Designed with durable, high-quality housing to provide a solid and reliable structure, ensuring long-term stability and repeatable performance. | Excellent readable display with bright backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | Built-in rechargeable backup battery and automatic shutoff features extend battery life, ensuring reliable performance and convenient everyday operation anywhere anytime.
Specifications - Diamond Scales CG 103K
| Maximum Capacity | 100 ct
| Readability | 0.001 ct
| Repeatability | 0.001 ct
| Weighing Pan Diameter | Ø 35 mm
| Setting Time | 2-3 sec
| Display | LCD Backlit
| Calibration | External
| Units of measure | g, ct, ozt, tl, tola, gn, dwt
| Operating temperature | 10ºC to 40ºC
| Interfaces | RS232
| Built-in Application | Counting, Percent Weighing, Check weighing
| Battery | 6V 1.2Ah rechargeable battery
| Power Supply | AC 220V / 50-60 Hz
| Balance Dimension (W x D X H) | 200 x 285 x 170 mm
Packaging dimensions (W x D X H)
| Net weight | 3 kg
| Gross weight | 3 Kg
Gold -Diamond Scales |  | Gold -Diamond Scales CY 513C
Designed to weigh gold and diamond jewelry samples with accuracy. Intuitive, solid-built construction with 510 g / 2550 ct capacity, 0.001 g / 0.005 ct readability, bright LCD backlit display, internal calibration, easy leveling, and USB/RS232 connectivity.
|  | Weighing Capacity | 510 g / 2550 ct
|  | Readability | 0.001 g / 0.005 ct
|  | Calibration | Motorised Internal Calibration (automatic/PSAC)
| Brilliant Proven Technology | Solid & Reliable Design | Bright, High-Contrast Display | Seamless Connectivity | 10 Easy-to-Use Built-In Applications | ISO/GLP Documentation | Quick & Simple Leveling | Easy Maintenance & Cleaning
| The sensor truly stands out with its unparalleled design based on the electromagnetic force compensation (EMFC). It provides an exceptional performance that meets the highest of the set international standards. | Built with a solid design and high-quality materials for long-term durability, ensuring reliable and consistent operation with accurate results. Supplied with a standard protective cover for added safety. | Excellent readable display with bright LED-backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | The USB and RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data. | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | The built-in real-time clock ensures accurate, time-stamped data, providing traceable documentation that reflects true test results and supports compliance with SOPs and GLP regulations. | An easy-to-view circular bubble level and large adjustable feet make balance leveling quick, simple, and precise for stable, accurate weighing. | Washable parts dismantle in seconds without tools. Dishwasher-safe components and a smooth housing design make routine cleaning quick, simple, and hygienic for everyday laboratory use.
|  | Specifications - Gold -Diamond Scales CY 513C | Specifications - Gold -Diamond Scales CY 1003C
| Maximum Capacity | 510 g / 2550 ct | 1000 g / 5000 ct
| Readability | 0.001 g / 0.005 ct | 0.001 g / 0.005 ct
| Repeatability | 0.001 g / 0.005 ct | 0.001 g / 0.005 ct
| Weighing Pan | 128 X 128 mm | 128 X 128 mm
| Setting Time | 2- 3 Sec | 2- 3 Sec
| Display | LCD Backlit | LCD Backlit
| Calibration | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC)
| Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
| Interfaces | RS232/USB | RS232/USB
| Power Supply | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A
| Balance Dimension (W x D X H) | 342.5 x 212 x 165 mm | 342.5 x 212 x 165 mm
| Packaging dimensions (W x D X H) | 490 × 400 × 520 mm | 490 × 400 × 520 mm
| Net weight | 7.8 kg | 7.8 kg
| Gross weight | 12 kg | 12 kg
Gold -Diamond Scales |  | Gold -Diamond Scales CY 223
Designed to weigh gold and diamond jewelry samples with accuracy. Intuitive, solid-built construction with 220 g / 1100 ct capacity, 0.001 g / 0.005 ct readability, bright LCD backlit display, external calibration, easy leveling, and USB/RS232 connectivity.
|  | Weighing Capacity | 220 g /1100 ct
|  | Readability | 0.001 g / 0.005 ct
|  | Calibration | External Calibration
| Brilliant Proven Technology | Solid & Reliable Design | Bright, High-Contrast Display | 10 Easy-to-Use Built-In Applications | ISO/GLP Documentation | Quick & Simple Leveling | Easy Maintenance & Cleaning
| The sensor truly stands out with its unparalleled design based on the electromagnetic force compensation (EMFC). It provides an exceptional performance that meets the highest of the set international standards. | Built with a solid design and high-quality materials for long-term durability, ensuring reliable and consistent operation with accurate results. Supplied with a standard protective cover for added safety. | Excellent readable display with bright LED-backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | The built-in real-time clock ensures accurate, time-stamped data, providing traceable documentation that reflects true test results and supports compliance with SOPs and GLP regulations. | An easy-to-view circular bubble level and large adjustable feet make balance leveling quick, simple, and precise for stable, accurate weighing. | Washable parts dismantle in seconds without tools. Dishwasher-safe components and a smooth housing design make routine cleaning quick, simple, and hygienic for everyday laboratory use.
|  | Specifications - Gold -Diamond Scales CY 223 | Specifications - Gold -Diamond Scales CY 323 | Specifications - Precision Balance CY 513 | Specifications - Gold -Diamond Scales CY 1003
| Maximum Capacity | 220 g / 1100 ct | 320 g / 1600 ct | 510 g / 2550 ct | 1000 g / 5000 ct
| Readability | 0.001 g / 0.005 ct | 0.001 g / 0.005 ct | 0.001 g / 0.005 ct | 0.001 g / 0.005 ct
| Repeatability | 0.001 g / 0.005 ct | 0.001 g / 0.005 ct | 0.001 g / 0.005 ct | 0.001 g / 0.005 ct
| Weighing Pan | 128 X 128 mm | 128 X 128 mm | 128 X 128 mm | 128 X 128 mm
| Setting Time | 2- 3 Sec | 2- 3 Sec | 2- 3 Sec | 2- 3 Sec
| Display | LCD Backlit | LCD Backlit | LCD Backlit | LCD Backlit
| Calibration | Eternal Calibration | Eternal Calibration | Eternal Calibration | Eternal Calibration
| Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
| Interfaces | RS232/USB | RS232/USB | RS232/USB | RS232/USB
| Power Supply | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A
| Balance Dimension (W x D X H) | 342.5 x 212 x 165 mm | 342.5 x 212 x 165 mm | 342.5 x 212 x 165 mm | 342.5 x 212 x 165 mm
| Packaging dimensions (W x D X H) | 490 × 400 × 520 mm | 490 × 400 × 520 mm | 490 × 400 × 520 mm | 490 × 400 × 520 mm
| Net weight | 7.8 kg | 7.8 kg | 7.8 kg | 7.8 kg
| Gross weight | 12 kg | 12 kg | 12 kg | 12 kg
Gold Scales CG 203L
Gold Scales |  | Designed for accurate gold weighing, featuring 200 g capacity, 0.001 g readability, bright LCD backlit display, and inbuilt rechargeable battery for reliable daily use in the jewelry business operations anywhere.
|  | Weighing Capacity | 200 g
|  | Readability | 0.001 g
|  | Calibration | External Calibration
| Precision Load Cell Technology | Solid & Reliable Design | Bright, High-Contrast Display
| Equipped with a precision aluminum load cell (IP65) and advanced 8-bit microprocessor design for accurate and stable performance. The level indicator with adjustable feet ensures proper setup, while 150% overload protection enhances safety and reliability. | Designed with durable, high-quality housing to provide a solid and reliable structure, ensuring long-term stability and repeatable performance. | Excellent readable display with bright backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation.
|  | Specifications - Gold Scales CG 203L | Specifications - Gold Scales CG 303L | Specifications -Gold Scales CG 503L
| Maximum Capacity | 200 g | 300 g | 500 g
| Readability | 0.001 g | 0.001 g | 0.001 g
| Repeatability | 0.001 g | 0.001 g | 0.001 g
| Weighing Pan Diameter | Ø 90 mm | Ø 90 mm | Ø 90 mm
| Setting Time | 2-3 sec | 2-3 sec | 2-3 sec
| Display | LCD Backlit | LCD Backlit | LCD Backlit
| Calibration | External | External | External
| Units of measure | g, ct, ozt, tl, tola, gn, dwt | g, ct, ozt, tl, tola, gn, dwt | g, ct, ozt, tl, tola, gn, dwt
| Operating temperature | 10ºC to 40ºC | 10ºC to 40ºC | 10ºC to 40ºC
| Interfaces | RS232 | RS232 | RS232
| Battery | 6V 1.2Ah rechargeable battery | 6V 1.2Ah rechargeable battery | 6V 1.2Ah rechargeable battery
| Power Supply | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz
| Balance Dimension (W x D X H) | 200 x 285 x 170 mm | 200 x 285 x 170 mm | 200 x 285 x 170 mm
Packaging dimensions (W x D X H)
| Net weight | 3 kg | 3 kg | 3 kg
| Gross weight | 3 Kg | 3 Kg | 3 Kg
Gold Scales |  | Gold Scales ALT 3102C
Designed for accurate gold jewellery weighing, the new generation technology with Solid built Construction. 3100 g capacity; 0.01 g readability; Colour TFT touchscreen display; internal Calibration; optimal feature levelling alert, easy levelling; and user management simplify and optimise your daily weighing processes. USB and RS232 connection; overload protection; small footprint; Power saving mode.
|  | Weighing Capacity | 3100 g
|  | Readability | 0.01 g
|  | Calibration | Motorised Internal Calibration (automatic/PSAC)
| Advanced EMFC Technology for Trusted Results | Intuitive Touchscreen | Automatic Internal Adjustment with PSAC | Infrared Proximity Sensor | ALIBI Memory secure storage of measurements | 10 Easy-to-Use Built-In Applications | Easy Leveling for Accurate Results | Optional On-Screen Leveling Assistance | Audit Trial Report | Seamless Connectivity
| Advanced EMFC sensor technology delivers international-standard performance. Engineered in India, ALT Series balances provide long-term accuracy, fast stabilization, excellent repeatability, overload protection, self-test reliability, and consistent results. | The large 4.3” color TFT touchscreen ensures clear visibility and smooth operation, even while wearing gloves. Its intuitive interface and wipe-clean surfaces allow quick, hygienic cleaning after every weighing task. | Automatic motorized, built-in weight internal adjustment ensures continuous accuracy using internal test weights. The balance performs PSAC at scheduled times or automatically whenever ambient temperature changes are detected. | Equipped with two programmable hands-free sensors, the balance allows convenient contactless operation. Each sensor can be assigned functions such as Zero, Tare, Print, Home Screen, or user-defined tasks, improving workflow efficiency and maintaining hygiene. | ALIBI memory securely stores up to 1000 reports. ensuring reliable data storage, full traceability, and long-term safety. It guarantees continuity and protection of vital measurement data over extended periods. USB-A (Optional for mass data storage in a pendrive ) | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | ALT Series balances ensure maximum stability with adjustable feet and a front bubble level, delivering perfectly leveled, accurate, and reliable weighing results. | Optional real-time digital level guidance uses an integrated sensor to continuously monitor the balance position and provides on-screen instructions, enabling quick and accurate leveling within seconds. | AuditTrail adds value by providing complete, gapless documentation with automatic recording of calibration and leveling processes, including secure date and time stamps for full traceability. | The RS232, USB-C, and Ethernet interfaces allow easy connection to a printer, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data.
|  | Specifications - Gold Scales ALT 3102C | Specifications - Gold Scales ALT 6102C
| Maximum Capacity | 3100 g | 6100 g
| Readability | 0.01 g | 0.01 g
| Repeatability | 0.01 g | 0.01 g
| Weighing Pan | 190 x 205 mm | 190 x 205 mm
| Setting Time | 2- 3 Sec | 2- 3 Sec
| Display | 4.3" color TFT touchscreen | 4.3" color TFT touchscreen
| Calibration | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC)
| Features | 7 User management Password protection AuditTrai ALIBI Memory ( 1000 reports are stored. ) Minimum Sample Weight (MSW) Touch-free operation Infrared Proximity Sensor 4 English, German, French, Sp built-in languagesanish Levelling guidance ( Optional ) USB-A (Optional for mass data storage in a pendrive ) | 7 User management Password protection AuditTrai ALIBI Memory ( 1000 reports are stored. ) Minimum Sample Weight (MSW) Touch-free operation Infrared Proximity Sensor 4 built-in languages English, German, French, Spanish Levelling guidance ( Optional ) USB-A (Optional for mass data storage in a pendrive )
| Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
| Interfaces | RS232/USB | RS232/USB
| Power Supply | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz. | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz.
Balance Dimension (W x D X H)
Packaging dimensions (W x D X H)
Net weight
Gross weight
Gold Scales |  | Gold Scales CY 10002HC
Intuitive to operate with solid, durable construction. Designed for accurate gold jewelry weighing with 10 kg capacity and 0.01 g readability. Large, bright LCD backlit display for easy viewing. Internal calibration and simple leveling for accurate results. USB and RS232 connectivity for seamless data transfer. Large pan size for high-capacity weighing applications.
|  | Weighing Capacity | 10 kg
|  | Readability | 0.01 g
|  | Calibration | Motorised Internal Calibration (automatic/PSAC)
| Brilliant Proven Technology | Solid & Reliable Design | Bright, High-Contrast Display | Seamless Connectivity | 10 Easy-to-Use Built-In Applications | ISO/GLP Documentation | Quick & Simple Leveling | Easy Maintenance & Cleaning
| The Precision Balances use a MonoTech single-block weighing sensor ensuring uniform load transmission, thermal stability, rapid stabilization, accurate corner-load performance, high repeatability, minimal drift, and long-term reliability in continuous operation. | Built with a solid design and high-quality materials for long-term durability, ensuring reliable and consistent operation with accurate results. Supplied with a standard protective cover for added safety. | Excellent readable display with bright LED-backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | The USB and RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data. | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | The built-in real-time clock ensures accurate, time-stamped data, providing traceable documentation that reflects true test results and supports compliance with SOPs and GLP regulations. | An easy-to-view circular bubble level and large adjustable feet make balance leveling quick, simple, and precise for stable, accurate weighing. | Washable parts dismantle in seconds without tools. Dishwasher-safe components and a smooth housing design make routine cleaning quick, simple, and hygienic for everyday laboratory use.
|  | Specifications - Gold Scales CY 10002HC | Specifications - Gold Scales CY 12002HC
| Maximum Capacity | 10 kg | 12 kg
| Readability | 0.01 g | 0.01 g
| Repeatability | 0.01 g | 0.01 g
| Weighing Pan | 190 x 205 mm | 190 x 205 mm
| Setting Time | 2- 3 Sec | 2- 3 Sec
| Display | Backlit LCD Display | Backlit LCD Display
| Calibration | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC)
| Units of measure | gm, mg, ct, GN, mo, oz, dwt, t1T, t1H, t1S, mom, Bat, MS | gm, mg, ct, GN, mo, oz, dwt, t1T, t1H, t1S, mom, Bat, MS
| Interfaces | RS-232/USB | RS-232/USB
| Power Supply | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz. | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz.
| Balance Dimension (W x D X H) | 517 x 302 x 130 mm | 518 x 302 x 130 mm
Packaging dimensions (W x D X H)
| Net weight | 16.3 kg | 16.3 kg
Gross weight
Gold Scales |  | Gold Scales ALT 3102
Designed for accurate gold jewelry weighing with New generation technology with Solid built Construction. 3100 g capacity; 0.01 g readability; Colour TFT touchscreen display; external Calibration; optimal feature leveling alert, easy leveling; and user management simplify and optimize your daily weighing processes. USB and RS232 connection; overload protection; small footprint; Power saving mode.
|  | Weighing Capacity | 3100 g
|  | Readability | 0.01 g
|  | Calibration | Eternal Calibration
| Advanced EMFC Technology for Trusted Results | Intuitive Touchscreen | Advanced User Management & Access Control | ALIBI Memory secure storage of measurements | 10 Easy-to-Use Built-In Applications | Easy Leveling for Accurate Results | Optional On-Screen Leveling Assistance | Audit Trial Report | Seamless Connectivity
| Advanced EMFC sensor technology delivers international-standard performance. Engineered in India, ALT Series balances provide long-term accuracy, fast stabilization, excellent repeatability, overload protection, self-test reliability, and consistent results. | The large 4.3” color TFT touchscreen ensures clear visibility and smooth operation, even while wearing gloves. Its intuitive interface and wipe-clean surfaces allow quick, hygienic cleaning after every weighing task. | User Management ensures secure operation by allowing administrators to create up to 7 users and assign specific access rights, ensuring controlled usage, data integrity, and protection against unauthorized actions. | ALIBI memory securely stores up to 1000 reports. ensuring reliable data storage, full traceability, and long-term safety. It guarantees continuity and protection of vital measurement data over extended periods. USB-A (Optional for mass data storage in a pendrive ) | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | ALT Series balances ensure maximum stability with adjustable feet and a front bubble level, delivering perfectly leveled, accurate, and reliable weighing results. | Optional real-time digital level guidance uses an integrated sensor to continuously monitor the balance position and provides on-screen instructions, enabling quick and accurate leveling within seconds. | AuditTrail adds value by providing complete, gapless documentation with automatic recording of calibration and leveling processes, including secure date and time stamps for full traceability. | The RS232, USB-C, and Ethernet interfaces allow easy connection to a printer, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data.
|  | Specifications - Gold Scales ALT 3102 | Specifications - Gold Scales ALT 6102
| Maximum Capacity | 3100 g | 6100 g
| Readability | 0.01 g | 0.01 g
| Repeatability | 0.01 g | 0.01 g
| Weighing Pan | 190 x 205 mm | 190 x 205 mm
| Setting Time | 2- 3 Sec | 2- 3 Sec
| Display | 4.3" color TFT touchscreen | 4.3" color TFT touchscreen
| Calibration | Eternal Calibration | Eternal Calibration
| Features | 7 User management Password protection AuditTrai ALIBI Memory ( 1000 reports are stored. ) Minimum Sample Weight (MSW) Touch-free operation Infrared Proximity Sensor 4 English, German, French, Spanish built-in languages A Levelling guidance ( Optional ) USB-A (Optional for mass data storage in a pendrive ) | 7 User management Password protection AuditTrai ALIBI Memory ( 1000 reports are stored. ) Minimum Sample Weight (MSW) Touch-free operation Infrared Proximity Sensor 4 English, German, French, Spanish built-in languages A Levelling guidance ( Optional ) USB-A (Optional for mass data storage in a pendrive )
| Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
| Interfaces | RS232/USB | RS232/USB
| Power Supply | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz. | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz.
Balance Dimension (W x D X H)
Packaging dimensions (W x D X H)
Net weight
Gross weight
Gold Scales |  | Gold Scales CY 3102C
Designed for accurate gold jewelry weighing "Intuitive to operate with solid built construction. 3100 g capacity; 0.01 g readability; large, bright LCD backlit display; internal Calibration; easy leveling; USB and RS232 connection.
|  | Weighing Capacity | 3100 g
|  | Readability | 0.01 g
|  | Calibration | Motorised Internal Calibration (automatic/PSAC)
| Brilliant Proven Technology | Solid & Reliable Design | Bright, High-Contrast Display | Seamless Connectivity | 10 Easy-to-Use Built-In Applications | ISO/GLP Documentation | Quick & Simple Leveling | Easy Maintenance & Cleaning
| The Precision Balances use a MonoTech single-block weighing sensor ensuring uniform load transmission, thermal stability, rapid stabilization, accurate corner-load performance, high repeatability, minimal drift, and long-term reliability in continuous operation. | Built with a solid design and high-quality materials for long-term durability, ensuring reliable and consistent operation with accurate results. Supplied with a standard protective cover for added safety. | Excellent readable display with bright LED-backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | The USB and RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data. | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | The built-in real-time clock ensures accurate, time-stamped data, providing traceable documentation that reflects true test results and supports compliance with SOPs and GLP regulations. | An easy-to-view circular bubble level and large adjustable feet make balance leveling quick, simple, and precise for stable, accurate weighing. | Washable parts dismantle in seconds without tools. Dishwasher-safe components and a smooth housing design make routine cleaning quick, simple, and hygienic for everyday laboratory use.
|  | Specifications - Gold Scales CY 3102C | Specifications - Gold Scales CY 6102C
| Maximum Capacity | 3100 g | 6100 g
| Readability | 0.01 g | 0.01 g
| Repeatability | 0.01 g | 0.01 g
| Weighing Pan | 180 x 195 mm | 190 x 205 mm
| Setting Time | 2- 3 Sec | 2- 3 Sec
| Display | LCD Backlit | LCD Backlit
| Calibration | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC)
| Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
| Interfaces | RS-232/USB | RS-232/USB
| Power Supply | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A
| Balance Dimension (W x D X H) | 342.5 x 212 x 89.5 mm | 342.5 x 212 x 89.5 mm
| Packaging dimensions (W x D X H) | 490 × 400 × 520 mm | 490 × 400 × 520 mm
| Net weight | 7.5 kg | 7.5 kg
| Gross weight | 12 kg | 12 kg
Gold Scales |  | Gold Scales CY 3102
Designed for accurate gold jewelry weighing Intuitive to operate with Solid built construction. 3100 g capacity; 0.01 g readability; large, bright LCD backlit display; external calibration; easy leveling; USB and RS232 connection; overload protection; small footprint.
|  | Weighing Capacity | 3100 g
|  | Readability | 0.01 g
|  | Calibration | Eternal Calibration
| Brilliant Proven Technology | Solid & Reliable Design | Bright, High-Contrast Display | 10 Easy-to-Use Built-In Applications | ISO/GLP Documentation | Quick & Simple Leveling | Easy Maintenance & Cleaning
| The Precision Balances use a MonoTech single-block weighing sensor ensuring uniform load transmission, thermal stability, rapid stabilization, accurate corner-load performance, high repeatability, minimal drift, and long-term reliability in continuous operation. | Built with a solid design and high-quality materials for long-term durability, ensuring reliable and consistent operation with accurate results. Supplied with a standard protective cover for added safety. | Excellent readable display with bright LED-backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | The built-in real-time clock ensures accurate, time-stamped data, providing traceable documentation that reflects true test results and supports compliance with SOPs and GLP regulations. | An easy-to-view circular bubble level and large adjustable feet make balance leveling quick, simple, and precise for stable, accurate weighing. | Washable parts dismantle in seconds without tools. Dishwasher-safe components and a smooth housing design make routine cleaning quick, simple, and hygienic for everyday laboratory use.
|  | Specifications - Gold Scales CY 3102 | Specifications - Gold Scales CY 6102
| Maximum Capacity | 3100 g | 6100 g
| Readability | 0.01 g | 0.01 g
| Repeatability | 0.01 g | 0.01 g
| Weighing Pan | 180 x 195 mm | 180 x 195 mm
| Setting Time | 2- 3 Sec | 2- 3 Sec
| Display | LCD Backlit | LCD Backlit
| Calibration | Eternal Calibration | Eternal Calibration
| Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
| Interfaces | RS-232/USB | RS-232/USB
| Power Supply | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A
| Balance Dimension (W x D X H) | 342.5 x 212 x 89.5 mm | 342.5 x 212 x 89.5 mm
| Packaging dimensions (W x D X H) | 490 × 400 × 520 mm | 490 × 400 × 520 mm
| Net weight | 7.5 kg | 7.5 kg
| Gross weight | 12 kg | 12 kg
Gold Scales |  | Gold Scales CG 302
Compact, durable, and easy to carry for accurate jewelry weighing anywhere. Features 600 g capacity, 0.01 g readability, bright LED display, inbuilt rechargeable battery, and overload protection for reliable performance in daily jewelry operations.
|  | Weighing Capacity | 600 g
|  | Readability | 0.01 g
|  | Calibration | Eternal Calibration
| Precision Load Cell Technology | Solid & Reliable Design | Bright LED Red Display
| Equipped with a precision aluminum load cell (IP65) and advanced 8-bit microprocessor design for accurate and stable performance. The level indicator with adjustable feet ensures proper setup, while 150% overload protection enhances safety and reliability. | Designed with durable, high-quality housing to provide a solid and reliable structure, ensuring long-term stability and repeatable performance. | High-visibility display with bright LED red digits ensures clear readability. The high-contrast screen offers wide viewing angles, enabling quick and accurate readings in all lighting conditions.
|  | Specifications - Gold Scales CG 302 | Specifications - Gold Scales CG 602
| Maximum Capacity | 300 g | 600 g
| Readability | 0.01 g | 0.01 g
| Repeatability | 0.01 g | 0.01 g
| Weighing Pan Diameter | Ø 118 mm | Ø 118 mm
| Setting Time | 2-3 sec | 2-3 sec
| Display | Bright LED Red | Bright LED Red
| Calibration | Eternal Calibration | Eternal Calibration
| Units of measure | g, ct, ozt, tl, tola, gn, dwt | g, ct, ozt, tl, tola, gn, dwt
| Interfaces | RS232 | RS232
| Battery | 6V 1.2Ah rechargeable battery | 6V 1.2Ah rechargeable battery
| Power Supply | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz
| Balance Dimension (W x D X H) | 190 x 230 x 155 mm | 190 x 230 x 155 mm
Packaging dimensions (W x D X H)
| Net weight | 2 kg | 2 kg
| Gross weight | 3 kg | 3kg
Gold Scales |  | Gold Scales CG 602L
Compact, durable, and easy to carry for accurate jewelry weighing anywhere. Features 600 g capacity, 0.01 g readability, Backlit LCD display, inbuilt rechargeable battery, and overload protection for reliable performance in daily jewelry operations.
|  | Weighing Capacity | 600 g
|  | Readability | 0.01 g
|  | Calibration | Eternal Calibration
| Precision Load Cell Technology | Solid & Reliable Design | Bright, High-Contrast Display
| Equipped with a precision aluminum load cell (IP65) and advanced 8-bit microprocessor design for accurate and stable performance. The level indicator with adjustable feet ensures proper setup, while 150% overload protection enhances safety and reliability. | Designed with durable, high-quality housing to provide a solid and reliable structure, ensuring long-term stability and repeatable performance. | Excellent readable display with bright backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation.
|  | Specifications - Gold Scales CG 302L | Specifications -Gold Scales CG 602L | Specifications - Gold Scales CG 2002L | 1202L
| Maximum Capacity | 300 g | 600 g | 2000 gm
| Readability | 0.01 g | 0.01 g | 0.01 g
| Repeatability | 0.01 g | 0.01 g | 0.01 g
| Linearity | 0.02 g | 0.02 g | 0.02 g
| Tare Range (Subtractive) | -300 g | -600 g | -2000 gm
| Weighing Pan | 120 x 120 mm | 120 x 120 mm | 120 x 120 mm
| Setting Time | 2 - 3 sec | 2 - 3 sec | 2 - 3 sec
| Display | Backlite LCD Display | Backlite LCD Display | Backlite LCD Display
| Calibration | External | External | External
| Units of measure | g, ct, ozt, tl, tola, gn, dwt | g, ct, ozt, tl, tola, gn, dwt | g, ct, ozt, tl, tola, gn, dwt
| Interfaces | RS232 | RS232 | RS232
| Battery | 6V 1.2Ah rechargeable battery | 6V 1.2Ah rechargeable battery | 6V 1.2Ah rechargeable battery
| Power Supply | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz
| Balance Dimension (W x D X H) | 200 x 285 x 170 mm | 200 x 285 x 170 mm | 200 x 285 x 170 mm
Packaging dimensions (W x D X H)
| Net weight | 2 kg | 2 kg | 2 kg
| Gross weight | 3 Kg | 3 Kg | 3 Kg
Silver Scales |  | Silver Scales CY 3101
Designed for accurate silver jewelry weighing Intuitive to operate with Solid built construction. 3100 g capacity; 0.1 g readability; large, bright LCD backlit display; external calibration; easy leveling; USB and RS232 connection; overload protection; small footprint.
|  | Weighing Capacity | 3100 g
|  | Readability | 0.1 g
|  | Calibration | Eternal Calibration
| Brilliant Proven Technology | Solid & Reliable Design | Bright, High-Contrast Display | 10 Easy-to-Use Built-In Applications | ISO/GLP Documentation | Quick & Simple Leveling | Easy Maintenance & Cleaning
| The Precision Balances use a MonoTech single-block weighing sensor ensuring uniform load transmission, thermal stability, rapid stabilization, accurate corner-load performance, high repeatability, minimal drift, and long-term reliability in continuous operation. | Built with a solid design and high-quality materials for long-term durability, ensuring reliable and consistent operation with accurate results. Supplied with a standard protective cover for added safety. | Excellent readable display with bright LED-backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | The built-in real-time clock ensures accurate, time-stamped data, providing traceable documentation that reflects true test results and supports compliance with SOPs and GLP regulations. | An easy-to-view circular bubble level and large adjustable feet make balance leveling quick, simple, and precise for stable, accurate weighing. | Washable parts dismantle in seconds without tools. Dishwasher-safe components and a smooth housing design make routine cleaning quick, simple, and hygienic for everyday laboratory use.
|  | Specifications - Silver Scales CY 3101 | Specifications - Silver Scales CY 6101
| Maximum Capacity | 3100 g | 6100 g
| Readability | 0.1 g | 0.1 g
| Repeatability | 0.1 g | 0.1 g
| Weighing Pan | 180 x 195 mm | 180 x 195 mm
| Setting Time | 2- 3 Sec | 2- 3 Sec
| Display | LCD Backlit | LCD Backlit
| Calibration | External Calibration | External Calibration
| Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
| Interfaces | RS-232/USB | RS-232/USB
| Power Supply | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A | DC Adapter, input 100 ~ 240 0.5A output 12V / 1A
| Balance Dimension (W x D X H) | 342.5 x 212 x 89.5 mm | 342.5 x 212 x 89.5 mm
| Packaging dimensions (W x D X H) | 490 × 400 × 520 mm | 490 × 400 × 520 mm
| Net weight | 7.5 kg | 7.5 kg
| Gross weight | 12 kg | 12 kg
Silver Scales |  | High CapacitySilver Scales AXR 15K1HC
Designed for accurate silver jewelry weighing Next-generation precision weighing with robust, durable construction for long-term reliability. High-capacity performance with 15 kg capacity and 0.1 g readability for precise micro-weighing. Vibrant color TFT touchscreen with internal calibration and easy leveling alerts for intuitive operation. 21 CFR Part 11 compliant user management ensures secure, traceable laboratory workflows. Multiple connectivity options — USB, RS232, and Ethernet — for seamless data transfer. large size stainless steel pan , safe, and efficient design with overload protection and power-saving mode, perfect for modern showroom.
|  | Weighing Capacity | 15 kg
|  | Readability | 0.1 g
|  | Calibration | Motorised Internal Calibration (automatic/PSAC)
| Advanced Single-Block Technology weighing Sensor | Intuitive Touchscreen | Automatic Internal Adjustment with PSAC | 10 Easy-to-use Built-in-Applications | 21 CFR Part 11 Compliance | Data Free Transfer | Compact & Safe Design
| The Precision Balances use a MonoTech single-block weighing sensor ensuring uniform load transmission, thermal stability, rapid stabilization, accurate corner-load performance, high repeatability, minimal drift, and long-term reliability in continuous operation. | The large 4.3” color TFT touchscreen ensures clear visibility and smooth operation, even while wearing gloves. Its intuitive interface and wipe-clean surfaces allow quick, hygienic cleaning after every weighing task. | Automatic motorized, built-in weight internal adjustment ensures continuous accuracy using internal test weights. The balance performs PSAC at scheduled times or automatically whenever ambient temperature changes are detected. | The ten easy to use build in applications include density determination, check weighing, totaling, and formulation, simplify your routine weighing tasks. These applications make your analytical operations much easier and accurate. | Our machine fully meets the requirements of 21 CFR Part 11/ EU GMP Annex 11. It also provides features like Full data security, Electronic signatures, Validations, Electronic records and Audit trails. | Simply transfer task parameters and results to a PC via many interface ports like RS232, USB A, USB B, Ethernet and wifi. Transcription errors are completely eliminated and reliability of your data is preserved. | Space-efficient construction ensures easy integration in any workspace. Features overload protection and power-saving functionality for reliable and efficient operation.
|  | Specifications -High Capacity Silver Scales AXR 15K1HC | Specifications -High Capacity Silver Scales AXR 35K1HC
| Maximum Capacity | 15 kg | 35 kg
| Readability | 0.1 g | 0.1 g
| Repeatability | 0.1 g | 0.1 g
| Weighing Pan | 370 x 360 mm | 370 x 360 mm
| Setting Time | 2- 3 Sec | 2 - 3 Sec.
| Display | 4.3" color TFT touchscreen | 4.3" color TFT touchscreen
| Calibration | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC)
| User Management | Log History [21 CFR Part 11 Compliant ] 7 User management Password protection AuditTrai Minimum Sample Weight (MSW) | Log History [21 CFR Part 11 Compliant ] 7 User management Password protection AuditTrai Minimum Sample Weight (MSW)
| Features | Touch-free operation Infrared Proximity Sensor [Automatic Door] 4 built-in languages English, German, French, Spanish Levelling guidance | Touch-free operation Infrared Proximity Sensor [Automatic Door] 4 built-in languages English, German, French, Spanish Levelling guidance
| Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
| Interfaces | RS232/USB | RS232/USB
| Power Supply | DC Adapter, input 100 ~ 240V 0.5A output 13V / 0.9A | DC Adapter, input 100 ~ 240V 0.5A output 13V / 0.9A
Balance Dimension (W x D X H)
Packaging dimensions (W x D X H)
Net weight
Gross weight
Gold Scales |  | High Capacity Precision Gold Scales AXR 10K2HC
Designed for accurate gold jewelry weighing Next-generation precision weighing with robust, durable construction for long-term reliability. High-capacity performance with 10 kg capacity and 0.01 g readability for precise micro-weighing. Vibrant color TFT touchscreen with internal calibration and easy leveling alerts for intuitive operation. 21 CFR Part 11 compliant user management ensures secure, traceable laboratory workflows. Multiple connectivity options — USB, RS232, and Ethernet — for seamless data transfer. Compact, safe, and efficient design with overload protection and power-saving mode, perfect for modern showroom.
|  | Weighing Capacity | 10 kg
|  | Readability | 0.01 g
|  | Minimum Weight (USP, 0.1%, typical) | 10 g
| Advanced Single-Block Technology weighing Sensor | Intuitive Touchscreen | Automatic Internal Adjustment with PSAC | 10 Easy-to-use Built-in-Applications | 21 CFR Part 11 Compliance | Data Free Transfer | Compact & Safe Design
| The Precision Balances use a MonoTech single-block weighing sensor ensuring uniform load transmission, thermal stability, rapid stabilization, accurate corner-load performance, high repeatability, minimal drift, and long-term reliability in continuous operation. | The large 4.3” color TFT touchscreen ensures clear visibility and smooth operation, even while wearing gloves. Its intuitive interface and wipe-clean surfaces allow quick, hygienic cleaning after every weighing task. | Automatic motorized, built-in weight internal adjustment ensures continuous accuracy using internal test weights. The balance performs PSAC at scheduled times or automatically whenever ambient temperature changes are detected. | The ten easy to use build in applications include density determination, check weighing, totaling, and formulation, simplify your routine weighing tasks. These applications make your analytical operations much easier and accurate. | Our machine fully meets the requirements of 21 CFR Part 11/ EU GMP Annex 11. It also provides features like Full data security, Electronic signatures, Validations, Electronic records and Audit trails. | Simply transfer task parameters and results to a PC via many interface ports like RS232, USB A, USB B, Ethernet and wifi. Transcription errors are completely eliminated and reliability of your data is preserved. | Space-efficient construction ensures easy integration in any workspace. Features overload protection and power-saving functionality for reliable and efficient operation.
|  | Specifications - High Capacity Precision Gold Scales AXR 10K2HC | Specifications - High Capacity Precision Gold Scales AXR 12K2HC
| Maximum Capacity | 10 kg | 12 kg
| Readability | 0.01 g | 0.01 g
| Repeatability | 0.01 g | 0.01 g
| Weighing Pan | 190 x 205 mm | 190 x 205 mm
| Setting Time | 2- 3 Sec | 2- 3 Sec
| Display | 4.3" color TFT touchscreen | 4.3" color TFT touchscreen
| Calibration | Motorised Internal Calibration (automatic/PSAC) | Motorised Internal Calibration (automatic/PSAC)
| User Management | Log History [21 CFR Part 11 Compliant ] 7 User management Password protection AuditTrai Minimum Sample Weight (MSW) | Log History [21 CFR Part 11 Compliant ] 7 User management Password protection AuditTrai Minimum Sample Weight (MSW)
| Features | Touch-free operation Infrared Proximity Sensor [Automatic Door] 4 built-in languages English, German, French, Spanish Levelling guidance | Touch-free operation Infrared Proximity Sensor [Automatic Door] 4 built-in languages English, German, French, Spanish Levelling guidance
| Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
| Interfaces | RS232/USB | RS232/USB
| Power Supply | DC Adapter, input 100 ~ 240V 0.5A output 13V / 0.9A | DC Adapter, input 100 ~ 240V 0.5A output 13V / 0.9A
Balance Dimension (W x D X H)
Packaging dimensions (W x D X H)
Net weight
Gross weight
Silver Scales |  | High Capacity Silver Scales CY 15001H
Designed for accurate silver jewelry weighing Intuitive to operate with solid, durable construction. High-capacity precision balance with 15 kg capacity and 0.1 g readability. Large, bright LCD backlit display for easy viewing. external calibration and simple leveling for accurate results. USB and RS232 connectivity for seamless data transfer. Large pan size for high-capacity weighing applications.
|  | Weighing Capacity | 15 kg
|  | Readability | 0.1 g
|  | Calibration | External Calibration
| Brilliant Proven Technology | Solid & Reliable Design | Bright, High-Contrast Display | Seamless Connectivity | 10 Easy-to-Use Built-In Applications | ISO/GLP Documentation | Quick & Simple Leveling | Easy Maintenance & Cleaning
| The Precision Balances use a MonoTech single-block weighing sensor ensuring uniform load transmission, thermal stability, rapid stabilization, accurate corner-load performance, high repeatability, minimal drift, and long-term reliability in continuous operation. | Built with a solid design and high-quality materials for long-term durability, ensuring reliable and consistent operation with accurate results. Supplied with a standard protective cover for added safety. | Excellent readable display with bright LED-backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | The USB and RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data. | Includes density determination, check weighing, totalling, and formulation functions to simplify routine weighing tasks and improve efficiency. | The built-in real-time clock ensures accurate, time-stamped data, providing traceable documentation that reflects true test results and supports compliance with SOPs and GLP regulations. | An easy-to-view circular bubble level and large adjustable feet make balance leveling quick, simple, and precise for stable, accurate weighing. | Washable parts dismantle in seconds without tools. Dishwasher-safe components and a smooth housing design make routine cleaning quick, simple, and hygienic for everyday laboratory use.
|  | Specifications - High Capacity Silver Scales CY 15001H | Specifications -High Capacity Silver Scales CY 25001H | Specifications - High Capacity Silver Scales CY 35001H
| Maximum Capacity | 15 kg | 25kg | 35 kg
| Readability | 0.1 g | 0.1 g | 0.1 g
| Repeatability | 0.1 g | 0.1 g | 0.1 g
| Weighing Pan | 370 x 360 mm | 370 x 360 mm | 370 x 360 mm
| Setting Time | 2 - 3 Sec. | 2 - 3 Sec. | 2 - 3 Sec.
| Display | Backlit LCD Display | Backlit LCD Display | Backlit LCD Display
| Calibration | External Calibration | External Calibration | External Calibration
| Units of measure | gm, mg, ct, GN, mo, oz, dwt, t1T, t1H, t1S, mom, Bat, MS | gm, mg, ct, GN, mo, oz, dwt, t1T, t1H, t1S, mom, Bat, MS | gm, mg, ct, GN, mo, oz, dwt, t1T, t1H, t1S, mom, Bat, MS
| Interfaces | RS232/USB | RS232/USB | RS232/USB
| Power Supply | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz. | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz. | AC/DC Adapter, 100 ~ 240 VAC, 50/60 Hz.
| Balance Dimension (W x D X H) | 517 x 302 x 130 mm | 518 x 302 x 130 mm | 519 x 302 x 130 mm
Packaging dimensions (W x D X H)
| Net weight | 16.3 kg | 16.3 kg | 16.3 kg
Gross weight
Silver Scales |  | Silver Scales CG 3001L
Compact, durable, and easy to carry for accurate jewelry weighing anywhere. Features 3000 g capacity, 0.1 g readability, Backlit LCD display, inbuilt rechargeable battery, and overload protection for reliable performance in daily jewelry operations.
|  | Weighing Capacity | 3000 g
|  | Readability | 0.1 g
|  | Calibration | External Calibration "
| Precision Load Cell Technology | Solid & Reliable Design | Bright, High-Contrast Display | Power Efficiency
| Equipped with a precision aluminum load cell (IP65) and advanced 8-bit microprocessor design for accurate and stable performance. The level indicator with adjustable feet ensures proper setup, while 150% overload protection enhances safety and reliability. | Designed with durable, high-quality housing to provide a solid and reliable structure, ensuring long-term stability and repeatable performance. | Excellent readable display with bright backlit LCD and large digits. High-contrast glass ensures wide viewing angles, while the intuitive interface allows easy, quick, and comfortable operation. | Built-in rechargeable backup battery and automatic shutoff features extend battery life, ensuring reliable performance and convenient everyday operation anywhere anytime.
|  | Specifications - Silver Scales CG 3001L | Specifications - Silver Scales CG 6001L | Specifications - Silver Scales CG 10001L | Specifications - Silver Scales CG 15001L | Specifications - Silver Scales CG 30001L
| Maximum Capacity | 3000 g | 6000 g | 10000 g | 15000 g | -30000 g
| Readability | 0.1 g | 0.1 g | 0.1 g | 0.1 g | 0.1 g
| Repeatability | 0.1 g | 0.1 g | 0.1 g | 0.1 g | 0.1 g
| Weighing Pan | 175 x 155 mm | 175 x 155 mm | 175 x 155 mm | 275 x 220 mm | 275 x 220 mm
| Setting Time | 2-3 sec | 2-3 sec | 2-3 sec | 2-3 sec | 2-3 sec
| Display | LCD Backlit | LCD Backlit | LCD Backlit | LCD Backlit | LCD Backlit
| Calibration | External | External | External | External | External
| Units of measure | g, ct, ozt, tl, tola, gn, dwt | g, ct, ozt, tl, tola, gn, dwt | g, ct, ozt, tl, tola, gn, dwt | kg, g, lb, oz | kg, g, lb, oz
| Interfaces | RS232 | RS232 | RS232 | RS232 | RS232
| Battery | 6V 1.2Ah rechargeable battery | 6V 1.2Ah rechargeable battery | 6V 1.2Ah rechargeable battery | 6V 1.2Ah rechargeable battery | 6V 1.2Ah rechargeable battery
| Power Supply | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz | AC 220V / 50-60 Hz
| Balance Dimension (W x D X H) | 190 x 290 x 90 mm | 190 x 290 x 90 mm | 190 x 290 x 90 mm | 300 x 290 x 90 mm | 300 x 290 x 90 mm
Packaging dimensions (W x D X H)
| Net weight | 2 kg | 2 kg | 2 kg | 5 kg | 5 kg
| Gross weight | 3 Kg | 3 Kg | 3 Kg | 6 kg | 6 kg
Karat Tester |  | Karat Tester CY 613 GT
Designed for precious metal purity testing, featuring 610 g capacity and 0.001 g/cm³ resolution. Accurately measures gold density and purity using Archimedes’ principle. Ideal for banks, pawn shops, scrap gold buyers, refineries, and jewelry businesses.
|  | Weighing Capacity | 610 g
|  | Readability | 0.001 g
|  | Resolution | 0.001 g/cm³
| Quick and Easy Operation | Brilliant Proven Technology | Changing Liquid Medium | Easy to read LCD Screen | Water Temperature Compensation | Seamless Connectivity | ISO/GLP Documentation
| The CY-GT Series Electronic Karat Tester is engineered for simplified and accurate density analysis. Featuring intuitive on-screen guidance, it enables reliable results through a streamlined two-step measurement process: • Step 1 – Measurement of weight in air • Step 2 – Measurement of weight in water | The sensor is designed based on electromagnetic force compensation (EMFC), ensuring high precision and stable performance. It delivers reliable results in accordance with established international standards. | The densimeter supports the use of alternative liquid media to accommodate challenging samples, such as those with air bubbles or complex geometries, ensuring accurate and reliable density measurements. H₂O / Ethanol / Other | The balance features a large backlit LCD display that ensures clear visibility under all lighting conditions. The integrated backlight enhances readability, enabling accurate viewing even in low-light environments. | Designed to ensure high accuracy by compensating for temperature variations, delivering precise density values at the defined reference temperature. | The USB and RS232 interfaces allow easy connection to a printer, secondary display, PC, or other digital systems, enabling fast, reliable, and error-free transfer of weighing data. | The built-in real-time clock ensures accurate, time-stamped data, providing traceable documentation that reflects true test results and supports compliance with SOPs and GLP regulations.
|  | Specifications - Karat Tester CY 613 GT | Specifications - Karat Tester CY 612 GT | Specifications - Karat Testers CY 3102 GT
| Maximum Capacity | 610 g | 610 g | 3100 g
| Readability | 0.001 g | 0.01 g | 0.01 g
| Density Resolution | 0.001 g/m³ | 0.001 g/m³ | 0.01 g/m³
| Repeatability | 0.001 g | 0.001 g | 0.01 g
| Weight in Air Pan Size | 85 x 106 mm | 85 x 106 mm | 124 x 160 mm
| Weight in Water Pan Size | 92 x 94 mm | 92 x 94 mm | 128 x 141 mm
| Setting Time | 2- 3 Sec | 2- 3 Sec | 2- 3 Sec
| Display | LCD Backlit | LCD Backlit | LCD Backlit
| Calibration | External Calibration | External Calibration | External Calibration
| Units of measure | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2 | gm, kg, lb, oz, ozt, GN, dwt, ct, mg, mom, m, tlH1, tlH2, tlS, tlt1, tlt2, b, mg1, mg2
| Operating temperature | 15° to 35°C | 15° to 35°C | 15° to 35°C
| Leveling system | manual | manual | manual
| Interface | RS232/USB | RS232/USB | RS232/USB
| Power Supply | DC Adaptor, input 100-240 0.8A output 13V/15.A, 50-60Hz | DC Adaptor, input 100-240 0.8A output 13V/15.A, 50-60Hz | DC Adaptor, input 100-240 0.8A output 13V/15.A, 50-60Hz
| Balance Dimension (W x D X H) | 342.5 x 212 x 270 mm | 342.5 x 212 x 270 mm | 342.5 x 212 x 270 mm
Packaging dimensions (W x D X H)
| Net weight | 10 kg | 10 kg | 10 kg
| Gross weight | 14 Kg | 14 Kg | 14 Kg

## Weighing Accessories for the La
Density Kit (Analytical Balance) CDK 01
Quickly convert analytical balance into a density measurement system for solids and floating samples, ensuring fast setup and accurate results.
Model | CDK O1
Compatible Balance | Analytical Balance CY/AXR Series
Dot Matrix Printer CPR 02
Compact Dot Matrix Printer CPR-02 with RS-232 interface provides reliable laboratory data recording, accurate result printouts, enhanced traceability, and regulatory compliance for weighing and measurement applications.
Model | CPR 02
Consumable | Paper Roll / Black ink printer ribbon cartridge
Printers | Dot Matrix Printer
No of Characters per line | 24 Character
Printing Speed | Faster Than 1 line per second
Paper Type | Normal Paper Roll, Wood Free Paper
Paper size | Width : 57.5 ±0.5mm. Core Diameter : 8mm Min.
Power Supply | AC Adopter, 100-270V AC, 50Hz - 60Hz
Interface | RS232
Bound Fate | 1200, 2400, 4800, 9600, 19200
Dot Matrix Printer CPR 03
Compact Dot Matrix Printer CPR-03 with RS-232 interface provides reliable laboratory data recording, accurate result printouts, enhanced traceability, and regulatory compliance. Universal compatibility enables seamless connection with most balance and instrument brands.
Model | CPR 03
Consumable | Paper Roll / Black ink printer ribbon cartridge
Printers | Dot Matrix Printer
No of Characters per line | 24 Character
Printing Speed | Faster Than 1 line per second
Paper Type | Normal Paper Roll, Wood Free Paper
Paper size | Width : 57.5 ±0.5mm. Core Diameter : 8mm Min.
Power Supply | AC Adopter, 100-270V AC, 50Hz - 60Hz
Interface | RS232
Bound Fate | 1200, 2400, 4800, 9600, 19200
Label printer CLP 04
CLP-04 label printer supports barcode and matrix code printing with specialized volumetric flask labels for easy scanning and removal.
Model | CLP 04
Barcode Reader CBR01
CBR-01 Barcode Reader enables fast, accurate barcode data capture.
Model | CBR01
Anti Vibration Weighing Table VRD - AVT 600
The CRD-AVT 600 anti-vibration weighing table provides a stable, durable platform for precision balances, minimizing vibrations and ensuring accurate results in laboratory environments.
Model | VRD - AVT 600
Antistatic Ionizer CI 01
The CI 01 Antistatic Ionizer ensures safe and effective static elimination, operating reliably between 0°C and 40°C for stable laboratory and weighing applications.
Model | CI 01
Vibrosorb-Anti Vibration Pad VRD 02
The VRD 02 Vibrosorb Anti-Vibration Pad, medium size, provides stable weighing support with 315 × 215 × 25 mm dimensions and 6.5 kg mass for vibration-free precision.
Model | VRD 02
Vibrosorb-Anti Vibration Pad VRD 03
The VRD 03 Vibrosorb Anti-Vibration Pad, large size, offers strong vibration absorption with 400 × 240 × 25 mm dimensions and 8.5 kg mass for stable precision weighing.
Model | VRD 03
Tweezers 1mg - 50g
The TWT-50 Tweezers are specially designed for handling weights from 1 mg to 50 g safely and precisely, preventing contamination during calibration and weighing.
Model | TWT-50
Tweezers 1g - 200g
The TWT-200 Tweezers are designed for safe and precise handling of weights from 1 g to 200 g, helping prevent contamination during calibration and weighing.
Model | TWT-200
Weight fork 1 kg
The TWF-1 Weight Fork is designed for safe and easy handling of 1 kg calibration weights, helping prevent direct contact, contamination, and handling errors.
Model | TWF-1
Weight handle TWH-2
Designed for safe and easy handling of 2 kg calibration weights, helping prevent direct contact, contamination, and handling errors.
Model | TWH-2
Weight handle TWH-5
Designed for safe and convenient handling of 5 kg calibration weights, ensuring clean, secure, and accurate weight handling during calibration.
Model | TWH-5
Weight handle TWH-10
Designed for reliable handling of 10 kg calibration weights, reducing manual contact and supporting safe movement during weighing and calibration tasks.
Model | TWH-10
Weight handle TWH-20
Designed for heavy-duty handling of 20 kg calibration weights, providing better grip, safety, and protection against contamination during use.
Model | TWH-20
Test Weight Trolley
The Steel Weight Trolley is designed for safe and convenient movement of heavy calibration weights, providing strong support, easy handling, and secure transportation in laboratory and industrial areas
Model | LBA-X-22
Temperature Kit
The Temperature Kit for moisture balance A helps verify and monitor heating temperature accuracy, ensuring reliable moisture analysis, stable drying performance, and consistent test results.
Model | CTK-X-01
Aluminum Dishes ( Single-use aluminum sample pans. Pack of 40. )
This design is suitable for handling a moisture sample. Package with 40 pcs
Model | CMD 40
Glass Fiber Filters GFF 100
Glass fiber filters for measuring samples with a high moisture content. Ensures even sample spread and enlarges the sample surface. Package with 100 pcs
Model | GFF 100
Remote Display
The SRD 01 Remote Display provides clear external weight indication with 19 mm digit size. It operates on 6V DC power from the measuring indicator scale and includes a 1.5 m connection cable for convenient placement.
Model | SRD 01
Specific Gravity Stand
The Specific Gravity Stand is designed for accurate density and specific gravity measurement of solid samples, providing stable support for reliable weighing and analysis.
Model | LBA-X 13
Interface Cable- USB
The Interface Cable enables reliable connection between weighing instruments and external devices, ensuring smooth data transfer, accurate communication, and efficient laboratory operation.
Model | LBA-X 15
Interface Cable- RS232
The RS232 Interface Cable provides a stable connection between weighing instruments and computers, printers, or indicators, enabling reliable serial data transfer and smooth communication.
Model | LBA-X 15
Anti-Thief Lock
The Anti-Theft Lock helps secure weighing instruments from unauthorized movement or theft, providing added safety and protection in laboratories, classrooms, and industrial areas.
Model | LBA-X 16
Evaporation Trap Kit
Traps moisture in the atmosphere surrounding the weighing beaker to keep the environment stable.
Model | CET 01

## Weighing Indicator
Weighbridge Indicator
The Weighbridge Indicator features a bright red LED display for clear weight visibility. It ensures accurate vehicle weighing, easy operation, and reliable data handling for industrial applications.
Model | WB IND 01
Intelligent Weigh Terminal
The Intelligent Weigh Terminal offers advanced weighing control with clear data display, smart operation, and reliable connectivity, making it ideal for industrial weighing and process applications.
Model | WB IT 01
Built-in Printer Indicator
The Built-in Printer Indicator provides accurate weighing display with integrated printing support, allowing instant record generation, easy operation, and reliable documentation for industrial weighing applications.
Model | WB
4-20mA Weighing Indicator
The 4–20mA Weighing Indicator provides accurate weight measurement with analog output support for PLC and automation systems, enabling smooth integration in industrial weighing and process control applications.
Model | CTG Series
Stainless Steel Indicator
The Stainless Steel Indicator is built for tough industrial environments, offering durable protection, accurate weight display, easy operation, and reliable performance in demanding weighing applications.
Model | CTG Series
Ethernet with WiFi
Ethernet with WiFi connectivity enables fast and reliable data transfer from weighing instruments to PCs, networks, or automation systems, supporting flexible communication and efficient data management.
Model | CTG Series
AIRPORT Baggage Scale Indicator
The CTG 300 4LABS Airport Baggage Scale Indicator is designed for accurate baggage weighing up to 300 kg, with 100 g readability for fast, clear, and reliable airport check-in operations.
Model | CTG 300 4LABS
Jumbo Display (Weighbridge)
The Jumbo Display for weighbridges provides large, bright, and clear weight readings from a distance, ensuring easy visibility, accurate monitoring, and smooth vehicle weighing operations.
Model | AJD
Flameproof Indicator
The Flameproof Indicator is designed for safe and accurate weighing in hazardous areas, offering reliable performance, clear display, and protection against explosive atmospheres.
Model | CTG-F Series
Ramp & Roller (Mild Steel)
The Mild Steel Ramp & Roller, sized 600 × 600 mm, is designed for smooth loading, unloading, and movement of heavy materials, ensuring durability and reliable industrial handling.
Model | CTG-RR
Size | 600 x 600 mm
Size | 750 x 750 mm
Ramp & Roller (Stainless Steel)
The CTG-RRS Stainless Steel Ramp & Roller, sized 600 × 600 mm, is designed for smooth loading, unloading, and movement of heavy materials, offering corrosion resistance, durability, and reliable industrial handling.
Model | CTG-RR
Size | 600 x 600 mm
4 Pack Caster Wheels
The 4 Pack Caster Wheels provide smooth mobility and stable support for equipment, making movement, positioning, and handling easy in laboratory and industrial applications.
Model | CTG-TP
Wireless Handheld Weighing Indicator
The Wireless Handheld Weighing Indicator provides convenient remote weight monitoring with clear display, easy operation, and reliable wireless communication for flexible industrial weighing applications.
Model | HHI Series
crane scale anti heating shield
The Crane Scale Anti-Heating Shield protects the scale from high-temperature exposure during weighing, ensuring safer operation, improved durability, and reliable performance in demanding industrial environments.
Model

## Test Weights
Weight Set 1 mg - 500 mg E1
E1 OIML class weight set (1 mg–500 mg) with storage box, handling accessories, cleaning tools, and NABL calibration certificate.
OIML Class
| Nominal Value | 1 mg-500 mg
| Calibration Certificate | NABL – ISO 17025 Accredited Certificate
OIML E1 Class Accuracy | Complete Weight Range | NABL Calibration Certificate
High-precision E1 class weight set designed for accurate calibration and verification of analytical and micro balances. | Includes nominal values from 1 mg to 500 mg, suitable for fine calibration and laboratory weighing applications. | Supplied with NABL–ISO 17025 accredited calibration certificate, storage box, handling accessories, and cleaning tools for reliable use.
Weight Set 1 mg - 200 g E1
E1 OIML class weight set (1 mg–200 g) weight in a aluminum box, handling accessories, cleaning tools, and NABL calibration certificate.
| OIML Class | E1
| Nominal Value | 1 mg-200 g
| Calibration Certificate | NABL – ISO 17025 Accredited Certificate
OIML E1 Class Accuracy | Wide Calibration Range | Complete Certified Kit
High-precision E1 class weight set designed for accurate calibration and verification of analytical and micro balances. | Covers nominal values from 1 mg to 200 g, making it suitable for multiple laboratory and industrial calibration requirements. | Supplied in an aluminum box with handling accessories, cleaning tools, and a NABL–ISO 17025 accredited calibration certificate for reliable traceability.
Weight Set 1 g - 2 kg E1
E1 OIML class weight set (1 g–2 kg) weight in a aluminum box, handling accessories, cleaning tools, and NABL calibration certificate.
| OIML Class | E1
| Nominal Value | 1 g - 2 kg
| Calibration Certificate | NABL – ISO 17025 Accredited Certificate
OIML E1 Class Precision | 1 g to 2 kg Weight Range | Complete Certified Kit
High-accuracy E1 class weight set designed for reliable calibration and verification of precision, analytical, and laboratory balances. | Includes nominal values from 1 g to 2 kg, suitable for routine calibration across a wide range of weighing instruments. | Supplied in an aluminum box with handling accessories, cleaning tools, and a NABL–ISO 17025 accredited calibration certificate for full traceability.
Weight Set 1 g - 5 kg E1
E1 OIML class weight set (1 g–5kg) weight in a aluminum box, handling accessories, cleaning tools, and NABL calibration certificate.
| OIML Class | E1
| Nominal Value | 1 g - 5 Kg
| Calibration Certificate | NABL – ISO 17025 Accredited Certificate
OIML E1 Class Precision | 1 g to 5 kg Weight Range | Complete Certified Kit
High-accuracy E1 class weight set designed for reliable calibration and verification of precision, analytical, and laboratory balances. | Includes nominal values from 1 g to 5 kg, suitable for calibration of medium to high-capacity weighing instruments. | Supplied in an aluminum box with handling accessories, cleaning tools, and a NABL–ISO 17025 accredited calibration certificate for full traceability.
Single Weight 500 g E1
Single weight 500 g E1 OIML class knob weight in a plastic box including a calibration certificate.
| OIML Class | E1
| Nominal Value | 500 g
| Calibration Certificate | NABL – ISO 17025 Accredited Certificate
OIML E1 Class Precision | Single 500 g Nominal Weight | Certified & Protected
High-precision 500 g E1 class knob weight designed for accurate calibration and verification of laboratory and precision balances. | Ideal for routine balance checking, calibration support, and performance verification where a 500 g reference weight is required. | Supplied in a plastic box with a NABL–ISO 17025 accredited calibration certificate for reliable traceability and safe storage.
Single Weight 1 kg E1
Single weight 1 kg E1 OIML class knob weight in a plastic box including a calibration certificate.
| OIML Class | E1
| Nominal Value | 1 kg
| Calibration Certificate | NABL – ISO 17025 Accredited Certificate
OIML E1 Class Precision | Single 1 kg Nominal Weight | Certified & Protected
High-precision 1 kg E1 class knob weight designed for accurate calibration and verification of laboratory and precision balances. | Suitable for routine balance calibration, performance checking, and verification where a 1 kg reference weight is required. | Supplied in a plastic box with a NABL–ISO 17025 accredited calibration certificate for reliable traceability and safe storage.
Single Weight 2 kg E1
Single weight 2 kg E1 OIML class knob weight in a plastic box including a calibration certificate.
| OIML Class | E1
| Nominal Value | 2 kg
| Calibration Certificate | NABL – ISO 17025 Accredited Certificate
OIML E1 Class Accuracy | Single 2 kg Nominal Weight | Certified & Protected
High-precision 2 kg E1 class knob weight designed for accurate calibration and verification of laboratory and precision balances. | Suitable for routine calibration, balance performance checking, and verification where a 2 kg reference weight is required. | Supplied in a plastic box with a NABL–ISO 17025 accredited calibration certificate for reliable traceability and safe storage.
Single Weight 5 kg E1
Single weight 5 kg E1 OIML class knob weight in a aluminum box including a calibration certificate.
| OIML Class | E1
| Nominal Value | 5 kg
| Calibration Certificate | NABL – ISO 17025 Accredited Certificate
OIML E1 Class Accuracy | Single 5 kg Nominal Weight | Certified & Protected
High-precision 5 kg E1 class knob weight designed for accurate calibration and verification of laboratory and precision balances. | Suitable for routine calibration, balance performance checking, and verification where a 5 kg reference weight is required. | Supplied in a plastic box with a NABL–ISO 17025 accredited calibration certificate for reliable traceability and safe storage.
Single Weight 10 kg E1
Single weight 10 kg E1 OIML class knob weight in a aluminum box including a calibration certificate.
| OIML Class | E1
| Nominal Value | 10 kg
| Calibration Certificate | NABL – ISO 17025 Accredited Certificate
OIML E1 Class Accuracy | Single 10 kg Nominal Weight | Certified & Protected
High-precision 10 kg E1 class knob weight designed for accurate calibration and verification of laboratory and precision balances. | Suitable for routine calibration, balance performance checking, and verification where a 10 kg reference weight is required. | Supplied in a plastic box with a NABL–ISO 17025 accredited calibration certificate for reliable traceability and safe storage.
Single Weight 50 kg E2
Single weight 50 kg E2 OIML class Ring Type Handle, including a calibration certificate no storage cases are provided along with the test weights.
| OIML Class | E2
| Nominal Value | 50 kg
| Calibration Certificate | NABL – ISO 17025 Accredited Certificate
OIML E2 Class Accuracy | Single 50 kg Nominal Weight | Certified & Protected
High-precision 50 kg E2 class ring-type handle weight designed for calibration and verification of high-capacity balances and weighing systems. | Suitable for routine calibration, performance checking, and verification where a 50 kg reference weight is required. | Supplied in a plastic box with a NABL–ISO 17025 accredited calibration certificate for reliable traceability and safe storage.
Single Weight 100 kg F1
Single weight 100 kg E2 OIML class Ring Type Handle, including a calibration certificate no storage cases are provided along with the test weights.
| OIML Class | F1
| Nominal Value | 100 kg
| Calibration Certificate | NABL – ISO 17025 Accredited Certificate
OIML F1 Class Accuracy | Single 100 kg Nominal Weight | Certified & Protected
High-precision 100 kg F1 class ring-type handle weight designed for calibration and verification of high-capacity balances and weighing systems. | Suitable for routine calibration, performance checking, and verification where a 100 kg reference weight is required. | Supplied in a plastic box with a NABL–ISO 17025 accredited calibration certificate for reliable traceability and safe storage.